## README
SERS Sepsis Project — Bulk Powder Raman Spectral Library (NAM, LTA, LPS)
Bulk (neat powder) samples measured directly on glass slides; B&W Tek i-Raman Plus BWS465-785H (BTC665N-785H-SYS), 785 nm; processed range 400–1800 cm-1.
Laser power is the instrument power level (%) from the file header; NOT converted to mW (sample power not measured). Objective: not recorded in the raw file headers (probe_name field empty).
Original raw files were never modified; 01_Raw_Data_Renamed contains byte-identical renamed copies (MD5-checked).
Naming: <Analyte>_Bulk_<YYYYMMDD>_P<power>pct_T<time>s_A<acc>_S<rep>.csv ; blanks: Glass_Blank_<YYYYMMDD>_P..pct_T..s_A.._S...csv ; replicate numbers are per date/condition in acquisition order.
Two NAM files were saved by the instrument in TXTR (semicolon) format; their renamed copies keep the .txtr extension.
Version 2.1 (23 Sep 2026): NAM 60 s re-audit, 70 % supporting group, final peak hierarchy (primary / supporting / glass-background / uncertain), LTA labelled PRELIMINARY (n = 4), LPS pending.
Sheets: 1_File_Inventory ¦ 2_QC ¦ 3_Reference_Conditions ¦ 4_Final_Marker_Table ¦ 5_Excluded_Data ¦ 6_Literature ¦ 7_Literature_Bands ¦ 8_All_Reproducible_Peaks ¦ 9_Reproducibility ¦ 10_Bulk_Intensity ¦ 11_Future_Acquisition ¦ 12_NAM_60s_ReAudit ¦ 13_NAM_70pct_vs_Reference ¦ 14_NAM_CrossDay_Detection
Key QC rules: Saturated (raw ≥ 65 000 counts) → excluded; SNR < 10 → Low SNR (kept, not used for reference); polystyrene-like pattern → contamination (excluded); correlation to class mean (glass zone 1308–1455 cm-1 excluded) < 0.85 → review.
SNR = height of strongest marker band (baseline-corrected) / SD of high-frequency noise (despiked minus SG-smoothed). Pearson correlation measures spectral reproducibility, not molecular identity.
Experimental peaks and literature peaks are reported in separate columns. Assignments are tentative unless stated; wording: "consistent with", not "proves".
VERSION 3 (23 Sep 2026): pipeline corrections (spike width guard, hot-pixel fix, robust noise) — see sheet 24. New sheets: Traceability_Table (central), 15–25.
Decisions: NAM = GO for identity/markers + top-up re-acquisition (≥10 independent spots at 90 %/30 s/5 acc); LTA = RE-ACQUIRE (80 %/45 s/2 acc, 10–15 spots); LPS = RE-ACQUIRE (70/80/90 % screening, 30 s, 5 acc, matched glass blanks).
UPDATE 27 Sep 2026 — NAM session 24 Sep 2026 added (sheets 26–32). The raw folder "bulk-lps-24-09-26" contains NAM (operator-confirmed); files renamed NAM_Bulk_20260924_* / Glass_Blank_20260924_*.
Result: 21/41 NAM spectra pass QC (failures = dark-reference ripple). New independent spots SUPPORT the NAM reference (positions ≤1.5 cm-1, relative intensities within ±7 % for primary markers). Reference condition unchanged (90/30/5); reference still INTERIM pending ≥4 more clean independent spots at 90 %.
LPS: no LPS data yet — screening still required.
UPDATE 28 Sep 2026 — NAM reference declared FINAL (no further NAM acquisition). Reference = 90 %/30 s/5 acc, 30 May 2026, n = 15 (unchanged files). Limitation: 6 independent spots at the reference condition; supported by 21 independent spots at 60–90 % (24 Sep) and 10 at 70 %/10 s. See sheet 37.
UPDATE 28 Sep 2026 — LPS reference ESTABLISHED: E. coli O111:B4 (Sigma L2630, lot 0000442391), 60 %/40 s/5 acc, n = 10; primary markers 425 and 1080 cm-1. Sheets 38–42.

## Traceability_Table
CENTRAL TABLE — marker traceability: valid spectrum → reproducible → peak reproducible → not background → chemically plausible → literature
Every final marker (primary and supporting) is traced from QC-passed spectra to its literature comparison. Assignments are TENTATIVE; no pure-compound normal-Raman reference exists for NAM, LTA or LPS.
LPS: no marker (reference pending controlled glass-slide measurement).
Biomarker | Category | Experimental_peak | SD | Intensity_counts | CV_pct | SNR | Replicate_frequency | Valid_spectrum | Reproducible | Not_background | Literature_peak | Assignment | Evidence_type | Citation | Confidence
NAM | Primary marker | 541 | 0 | 715 ± 33 | 4.6 | 39.9 | 15/15 ref; min day 100 %; indep. spots 100 % | Yes (QC pass) | Yes | Yes | 400–600 (region) | Proposed: pyranose-ring skeletal deformation δ(C-C-O)/δ(C-C-C) | Review of pure related compounds (carbohydrates) | L6 | Low
NAM | Primary marker | 829 | 0 | 625 ± 30 | 4.8 | 34.8 | 15/15 ref; min day 100 %; indep. spots 100 % | Yes (QC pass) | Yes | Yes | 800–950 (region) | Proposed: ν(C-O-C)/ring mode in the carbohydrate anomeric region | Review of pure related compounds (carbohydrates) | L6 | Low
NAM | Primary marker | 870.5 | 0.5 | 872 ± 37 | 4.3 | 48.6 | 15/15 ref; min day 100 %; indep. spots 100 % | Yes (QC pass) | Yes | Yes | 800–950 (region) | Proposed: ν(C-O-C)/ν(C-C) ring mode (anomeric region) | Review of pure related compounds (carbohydrates) | L6 | Low
NAM | Primary marker | 929 | 0 | 917 ± 37 | 4 | 51.1 | 15/15 ref; min day 100 %; indep. spots 100 % | Yes (QC pass) | Yes | Yes | 800–950 (region) | Proposed: ν(C-O-C)/ν(C-C); possible lactyl C-COOH contribution | Review of pure related compounds (carbohydrates); Band compilation from biological tissues | L6, L5 | Low
NAM | Primary marker | 955 | 0 | 862 ± 44 | 5.1 | 48.1 | 15/15 ref; min day 100 %; indep. spots 100 % | Yes (QC pass) | Yes | Yes | 800–950 (upper edge) | Proposed: ν(C-C) / CH3 rocking (N-acetyl or lactyl methyl) – tentative | Band compilation from biological tissues; Review of pure related compounds (carbohydrates) | L5, L6 | Low
NAM | Primary marker | 1023 | 0.5 | 749 ± 47 | 6.2 | 41.8 | 15/15 ref; min day 100 %; indep. spots 100 % | Yes (QC pass) | Yes | Yes | 1000–1200 (region) | Proposed: ν(C-O)/ν(C-C) (C-OH stretching) of sugar ring | Review of pure related compounds (carbohydrates) | L6 | Low
NAM | Primary marker | 1127 | 0 | 581 ± 30 | 5.2 | 32.4 | 15/15 ref; min day 98 %; indep. spots 100 % | Yes (QC pass) | Yes | Yes | 1000–1200 (region) | Proposed: ν(C-O-C) (3-O-lactyl ether / ring) or ν(C-C) | Review of pure related compounds (carbohydrates) | L6 | Low
NAM | Supporting peak | 457 | 0.4 | 547 ± 32 | 5.8 | 30.5 | 15/15 ref; min day 97 %; indep. spots 100 % | Yes (QC pass) | Yes | Partly (edge/glass) | 400–600 (region) | Proposed: pyranose-ring skeletal deformation | Review of pure related compounds (carbohydrates) | L6 | Low
NAM | Supporting peak | 487 | 0 | 398 ± 33 | 8.3 | 22.2 | 15/15 ref; min day 87 %; indep. spots 100 % | Yes (QC pass) | Partly | Yes | 400–600 (region) | Proposed: skeletal deformation | Review of pure related compounds (carbohydrates) | L6 | Low
NAM | Supporting peak | 1249 | 0.5 | 290 ± 11 | 3.7 | 16.2 | 15/15 ref; min day 72 %; indep. spots 100 % | Yes (QC pass) | Partly | Partly (edge/glass) | — | No assignment proposed (possible amide III / CH deformation region; not verified) | — | — | Uncertain
NAM | Supporting peak | 1589.5 | 0.5 | 263 ± 23 | 8.9 | 14.7 | 15/15 ref; min day 66 %; indep. spots 100 % | Yes (QC pass) | Partly | Partly (edge/glass) | — | No assignment proposed (unassigned; reproducible) | — | — | Uncertain
LTA | Primary marker | 802 | 5.1 | 834 ± 105 | 12.6 | 21.4 | 4/4 ref; min day 89 % | Yes (QC pass) | Yes | Yes | ~813 (ν O-P-O) | Proposed: ν(O-P-O) phosphodiester stretch | Band compilation from biological tissues | L5 | Low
LTA | Primary marker | 861.5 | 1 | 687 ± 74 | 10.8 | 17.6 | 4/4 ref; min day 89 % | Yes (QC pass) | Yes | Yes | 800–950 (region) | Proposed: ν(C-C)/ν(C-O-C) of glycerol / sugar substituents | Review of pure related compounds (carbohydrates) | L6 | Low
LTA | Primary marker | 1092 | 0.5 | 1110 ± 98 | 8.8 | 28.5 | 4/4 ref; min day 100 % | Yes (QC pass) | Yes | Yes | 1090–1095 (ν_s PO2−) | Proposed: ν_s(PO2−) of phosphodiester (poly-glycerophosphate backbone) | Band compilation from biological tissues; Review of pure related compounds (lipids); Biological-sample FTIR (whole cells) | L5, L7, L8 | Moderate
LTA | Supporting peak | 1459 | 0.5 | 721 ± 53 | 7.4 | 18.5 | 4/4 ref; min day 100 % | Yes (QC pass) | Yes | Partly (edge/glass) | 1440–1460 (δ CH2) | Proposed: δ(CH2) scissoring (acyl chains of glycolipid anchor); non-specific | Review of pure related compounds (lipids) | L7 | Low

## 1_File_Inventory
Sheet 1 — File inventory (every raw file in the data folder)
Out-of-scope files (SERS/R6G/AgNO3/HF/substrate studies) are listed for completeness but not copied into the library.
Original file (relative to sukesh folder) | New file name | Biomarker / class | Acquired | Power (%) | Integration (s) | Accumulations | Replicate | Spot | Objective | Laser (nm) | Raman shift min | Raman shift max | Data points | Max raw counts (full range) | Status | Library folder
29-05-26/glass slide empty/glassslide (4).csv | Glass_Blank_20260529_P005pct_T05s_A01_S01.csv | Glass blank | 2026-05-29 17:06:29 | 5 | 5 | 1 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7681 | Valid | Glass_Blank
29-05-26/glass slide empty/glassslide (5).csv | Glass_Blank_20260529_P005pct_T05s_A01_S02.csv | Glass blank | 2026-05-29 17:08:13 | 5 | 5 | 1 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7824 | Valid | Glass_Blank
29-05-26/glass slide empty/glassslide (1).csv | Glass_Blank_20260529_P005pct_T05s_A01_S03.csv | Glass blank | 2026-05-29 17:08:26 | 5 | 5 | 1 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7710 | Valid | Glass_Blank
29-05-26/glass slide empty/glassslide (2).csv | Glass_Blank_20260529_P005pct_T05s_A01_S04.csv | Glass blank | 2026-05-29 17:08:45 | 5 | 5 | 1 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7934 | Valid | Glass_Blank
29-05-26/glass slide empty/glassslide (3).csv | Glass_Blank_20260529_P005pct_T05s_A01_S05.csv | Glass blank | 2026-05-29 17:08:59 | 5 | 5 | 1 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7979 | Valid | Glass_Blank
13-4-26/lps/20-20-5/SP_1.csv | LPS_Bulk_20260413_P010pct_T10s_A05_S01.csv | LPS bulk | 2026-04-13 11:27:46 | 10 | 10 | 5 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5043 | Signal not established | LPS_Bulk
13-4-26/lps/20-20-5/SP_2.csv | LPS_Bulk_20260413_P010pct_T10s_A05_S02.csv | LPS bulk | 2026-04-13 11:29:28 | 10 | 10 | 5 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4839.2 | Signal not established | LPS_Bulk
13-4-26/lps/20-20-5/SP_3.csv | LPS_Bulk_20260413_P020pct_T10s_A05_S01.csv | LPS bulk | 2026-04-13 11:30:42 | 20 | 10 | 5 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6535.8 | Signal not established | LPS_Bulk
13-4-26/lps/20-20-5/SP_4.csv | LPS_Bulk_20260413_P020pct_T10s_A02_S01.csv | LPS bulk | 2026-04-13 11:31:46 | 20 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6072.5 | Signal not established | LPS_Bulk
13-4-26/lps/20-20-5/SP_5.csv | LPS_Bulk_20260413_P020pct_T60s_A30_S01.csv | LPS bulk | 2026-04-13 12:52:30 | 20 | 60 | 30 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 29874.0667 | Signal not established | LPS_Bulk
13-4-26/lps/20-20-5/SP_6.csv | LPS_Bulk_20260413_P010pct_T60s_A30_S01.csv | LPS bulk | 2026-04-13 13:28:32 | 10 | 60 | 30 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 21942.7 | Signal not established | LPS_Bulk
13-4-26/lps/20-20-5/SP_7.csv | LPS_Bulk_20260413_P020pct_T60s_A50_S01.csv | LPS bulk | 2026-04-13 14:24:38 | 20 | 60 | 50 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 31704.56 | Signal not established | LPS_Bulk
13-4-26/lps/20-20-5/SP_8.csv | LPS_Bulk_20260413_P020pct_T60s_A30_S02.csv | LPS bulk | 2026-04-13 15:04:02 | 20 | 60 | 30 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 29108.0667 | Signal not established | LPS_Bulk
13-4-26/lps/20-20-5/SP_9.csv | LPS_Bulk_20260413_P020pct_T60s_A50_S02.csv | LPS bulk | 2026-04-13 17:10:30 | 20 | 60 | 50 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 30282.72 | Signal not established | LPS_Bulk
14-426/lps/SP_1.csv | LPS_Bulk_20260413_P010pct_T10s_A02_S01.csv | LPS bulk | 2026-04-13 19:48:13 | 10 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4299.5 | Signal not established | LPS_Bulk
14-426/lps/SP_2.csv | LPS_Bulk_20260413_P010pct_T10s_A02_S02.csv | LPS bulk | 2026-04-13 19:50:57 | 10 | 10 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4406 | Signal not established | LPS_Bulk
14-426/lps/SP_3.csv | LPS_Bulk_20260413_P010pct_T10s_A02_S03.csv | LPS bulk | 2026-04-13 19:53:05 | 10 | 10 | 2 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4521 | Signal not established | LPS_Bulk
14-426/lps/SP_4.csv | LPS_Bulk_20260413_P020pct_T30s_A60_S01.csv | LPS bulk | 2026-04-13 21:22:24 | 20 | 30 | 60 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17163.1 | Signal not established | LPS_Bulk
14-426/lps/SP_5.csv | LPS_Bulk_20260413_P020pct_T30s_A50_S01.csv | LPS bulk | 2026-04-13 22:26:52 | 20 | 30 | 50 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17087.98 | Signal not established | LPS_Bulk
4-4-26/LTA/45sec diffrent powrset2/SP_1.csv | LTA_Bulk_20260403_P080pct_T45s_A02_S01.csv | LTA bulk | 2026-04-03 07:05:44 | 80 | 45 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 47406 | Valid | LTA_Bulk
4-4-26/LTA/45sec diffrent powrset2/SP_2.csv | LTA_Bulk_20260403_P090pct_T45s_A02_S01.csv | LTA bulk | 2026-04-03 07:07:39 | 90 | 45 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 48531.5 | Valid | LTA_Bulk
4-4-26/LTA/45sec diffrent powrset2/SP_3.csv | LTA_Bulk_20260403_P095pct_T45s_A02_S01.csv | LTA bulk | 2026-04-03 07:11:29 | 95 | 45 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 53846 | Valid | LTA_Bulk
4-4-26/LTA/45sec diffrent powrset2/SP_4.csv | LTA_Bulk_20260403_P080pct_T45s_A02_S02.csv | LTA bulk | 2026-04-03 07:13:42 | 80 | 45 | 2 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 43075 | Valid | LTA_Bulk
4-4-26/LTA/45sec diffrent powrset2/SP_5.csv | LTA_Bulk_20260403_P060pct_T45s_A02_S01.csv | LTA bulk | 2026-04-03 07:17:37 | 60 | 45 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 34145 | Valid | LTA_Bulk
4-4-26/LTA/30sec50power15set/SP_6.csv | LTA_Bulk_20260403_P050pct_T30s_A15_S01.csv | LTA bulk | 2026-04-03 07:36:45 | 50 | 30 | 15 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20222.6667 | Abnormal - review | LTA_Bulk
4-4-26/LTA/10sec,diffrenpower,10set/SP_7.csv | LTA_Bulk_20260403_P050pct_T10s_A10_S01.csv | LTA bulk | 2026-04-03 07:49:35 | 50 | 10 | 10 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8096.3 | Valid | LTA_Bulk
4-4-26/LTA/10sec,diffrenpower,10set/SP_8.csv | LTA_Bulk_20260403_P060pct_T10s_A10_S01.csv | LTA bulk | 2026-04-03 07:51:42 | 60 | 10 | 10 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8643.6 | Valid | LTA_Bulk
4-4-26/LTA/10sec,diffrenpower,10set/SP_9.csv | LTA_Bulk_20260403_P040pct_T10s_A10_S01.csv | LTA bulk | 2026-04-03 07:53:43 | 40 | 10 | 10 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5842.5 | Valid | LTA_Bulk
4-4-26/LTA/10sec,diffrenpower,10set/SP_11.csv | LTA_Bulk_20260403_P070pct_T10s_A10_S01.csv | LTA bulk | 2026-04-03 08:00:17 | 70 | 10 | 10 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9892.5 | Valid | LTA_Bulk
4-4-26/5sec,30,50 power 10 set/SP_12.csv | LTA_Bulk_20260403_P030pct_T05s_A10_S01.csv | LTA bulk | 2026-04-03 08:04:18 | 30 | 5 | 10 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3119 | Low SNR | LTA_Bulk
4-4-26/5sec,30,50 power 10 set/SP_13.csv | LTA_Bulk_20260403_P050pct_T05s_A10_S01.csv | LTA bulk | 2026-04-03 08:05:46 | 50 | 5 | 10 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4454.3 | Low SNR | LTA_Bulk
NAM-15secs and power15-75-2sets/SP_4.csv | NAM_Bulk_20260128_P005pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 17:53:49 | 5 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3698.5 | Low SNR | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_5.csv | NAM_Bulk_20260128_P010pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 17:55:31 | 10 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4679.5 | Low SNR | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_6.csv | NAM_Bulk_20260128_P010pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 17:57:00 | 10 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4519.5 | Low SNR | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_7.csv | NAM_Bulk_20260128_P020pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 17:58:19 | 20 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6796.5 | Low SNR | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_8.csv | NAM_Bulk_20260128_P025pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 17:59:01 | 25 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8157.5 | Low SNR | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_9.csv | NAM_Bulk_20260128_P030pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:01:08 | 30 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9484 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_10.csv | NAM_Bulk_20260128_P030pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 18:04:29 | 30 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9304 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_11.csv | NAM_Bulk_20260128_P035pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:09:20 | 35 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10041 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_12.csv | NAM_Bulk_20260128_P040pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:11:16 | 40 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11643.5 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_13.csv | NAM_Bulk_20260128_P040pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 18:13:01 | 40 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11266.5 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_14.csv | NAM_Bulk_20260128_P045pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:14:13 | 45 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 12142.5 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_15.csv | NAM_Bulk_20260128_P050pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:18:46 | 50 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13073 | Review - higher noise | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_16.csv | NAM_Bulk_20260128_P060pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:20:12 | 60 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15251 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_17.csv | NAM_Bulk_20260128_P055pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:22:17 | 55 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14103 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_18.csv | NAM_Bulk_20260128_P065pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:24:27 | 65 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15625.5 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_19.csv | NAM_Bulk_20260128_P070pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:26:22 | 70 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16923.5 | Valid | NAM_Bulk
NAM-15secs and power15-75-2sets/SP_20.csv | NAM_Bulk_20260128_P075pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 18:28:21 | 75 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17505.5 | Valid | NAM_Bulk
5sec-power5-80 and 2set/SP_21.csv | NAM_Bulk_20260128_P005pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:04:55 | 5 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1808 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_22.csv | NAM_Bulk_20260128_P010pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:05:32 | 10 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2131.5 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_23.csv | NAM_Bulk_20260128_P015pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:05:53 | 15 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2505 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_24.csv | NAM_Bulk_20260128_P020pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:07:23 | 20 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2890.5 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_25.csv | NAM_Bulk_20260128_P025pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:07:53 | 25 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3273 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_26.csv | NAM_Bulk_20260128_P030pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:08:33 | 30 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3542.5 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_27.csv | NAM_Bulk_20260128_P035pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:09:21 | 35 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3907.5 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_28.csv | NAM_Bulk_20260128_P040pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:10:06 | 40 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4158 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_29.csv | NAM_Bulk_20260128_P045pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:10:36 | 45 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4561 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_30.csv | NAM_Bulk_20260128_P050pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:11:18 | 50 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4879.5 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_31.csv | NAM_Bulk_20260128_P055pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:12:09 | 55 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5236 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_32.csv | NAM_Bulk_20260128_P060pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:12:47 | 60 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5553.5 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_33.csv | NAM_Bulk_20260128_P065pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:13:24 | 65 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5946.5 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_34.csv | NAM_Bulk_20260128_P070pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:13:52 | 70 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6262 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_35.csv | NAM_Bulk_20260128_P075pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:14:13 | 75 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6664 | Low SNR | NAM_Bulk
5sec-power5-80 and 2set/SP_36.csv | NAM_Bulk_20260128_P080pct_T05s_A02_S01.csv | NAM bulk | 2026-01-28 19:16:04 | 80 | 5 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6996.5 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_37.csv | NAM_Bulk_20260128_P005pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:21:39 | 5 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2896.5 | Low SNR | NAM_Bulk
10sec-power5-80-2set/SP_38.csv | NAM_Bulk_20260128_P010pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:22:49 | 10 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3332 | Low SNR | NAM_Bulk
10sec-power5-80-2set/SP_39.csv | NAM_Bulk_20260128_P015pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:23:24 | 15 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4248.5 | Low SNR | NAM_Bulk
10sec-power5-80-2set/SP_40.csv | NAM_Bulk_20260128_P020pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:24:09 | 20 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4923 | Low SNR | NAM_Bulk
10sec-power5-80-2set/SP_41.csv | NAM_Bulk_20260128_P020pct_T10s_A02_S02.csv | NAM bulk | 2026-01-28 19:25:00 | 20 | 10 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4946.5 | Low SNR | NAM_Bulk
10sec-power5-80-2set/SP_42.csv | NAM_Bulk_20260128_P025pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:26:14 | 25 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5634 | Low SNR | NAM_Bulk
10sec-power5-80-2set/SP_43.csv | NAM_Bulk_20260128_P030pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:26:57 | 30 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6547 | Low SNR | NAM_Bulk
10sec-power5-80-2set/SP_44.csv | NAM_Bulk_20260128_P035pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:28:23 | 35 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6950 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_45.csv | NAM_Bulk_20260128_P040pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:29:17 | 40 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7227.5 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_46.csv | NAM_Bulk_20260128_P045pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:30:09 | 45 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8021 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_47.csv | NAM_Bulk_20260128_P050pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:31:02 | 50 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8373 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_48.csv | NAM_Bulk_20260128_P055pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:32:25 | 55 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9363 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_49.csv | NAM_Bulk_20260128_P060pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:33:04 | 60 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10106.5 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_50.csv | NAM_Bulk_20260128_P065pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:33:44 | 65 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10655.5 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_51.csv | NAM_Bulk_20260128_P070pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:34:31 | 70 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11417.5 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_52.csv | NAM_Bulk_20260128_P075pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:35:18 | 75 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 12369.5 | Valid | NAM_Bulk
10sec-power5-80-2set/SP_53.csv | NAM_Bulk_20260128_P080pct_T10s_A02_S01.csv | NAM bulk | 2026-01-28 19:36:06 | 80 | 10 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13500 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_54.csv | NAM_Bulk_20260128_P005pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:43:06 | 5 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3583.5 | Low SNR | NAM_Bulk
15sec-power5-80andset2/SP_55.csv | NAM_Bulk_20260128_P010pct_T15s_A02_S03.csv | NAM bulk | 2026-01-28 19:44:46 | 10 | 15 | 2 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4739.5 | Low SNR | NAM_Bulk
15sec-power5-80andset2/SP_56.csv | NAM_Bulk_20260128_P015pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 19:45:53 | 15 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5977 | Low SNR | NAM_Bulk
15sec-power5-80andset2/SP_57.csv | NAM_Bulk_20260128_P020pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:46:53 | 20 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7166.5 | Low SNR | NAM_Bulk
15sec-power5-80andset2/SP_58.csv | NAM_Bulk_20260128_P025pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:47:54 | 25 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8174 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_59.csv | NAM_Bulk_20260128_P030pct_T15s_A02_S03.csv | NAM bulk | 2026-01-28 19:48:39 | 30 | 15 | 2 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9024.5 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_60.csv | NAM_Bulk_20260128_P035pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:49:28 | 35 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10229 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_61.csv | NAM_Bulk_20260128_P040pct_T15s_A02_S03.csv | NAM bulk | 2026-01-28 19:50:14 | 40 | 15 | 2 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11825.5 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_62.csv | NAM_Bulk_20260128_P045pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:51:03 | 45 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 12868.5 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_63.csv | NAM_Bulk_20260128_P050pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:52:04 | 50 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13233.5 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_64.csv | NAM_Bulk_20260128_P055pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:54:54 | 55 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14549 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_65.csv | NAM_Bulk_20260128_P060pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:55:46 | 60 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15617.5 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_66.csv | NAM_Bulk_20260128_P065pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:56:45 | 65 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16585.5 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_67.csv | NAM_Bulk_20260128_P070pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:57:45 | 70 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17121.5 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_68.csv | NAM_Bulk_20260128_P075pct_T15s_A02_S02.csv | NAM bulk | 2026-01-28 19:58:26 | 75 | 15 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18758.5 | Valid | NAM_Bulk
15sec-power5-80andset2/SP_69.csv | NAM_Bulk_20260128_P080pct_T15s_A02_S01.csv | NAM bulk | 2026-01-28 19:59:20 | 80 | 15 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18726 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_70.csv | NAM_Bulk_20260128_P005pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:04:32 | 5 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4606 | Low SNR | NAM_Bulk
20sec-power5-80and set2/SP_71.csv | NAM_Bulk_20260128_P010pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:06:41 | 10 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5726 | Low SNR | NAM_Bulk
20sec-power5-80and set2/SP_72.csv | NAM_Bulk_20260128_P015pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:08:00 | 15 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7541 | Low SNR | NAM_Bulk
20sec-power5-80and set2/SP_73.csv | NAM_Bulk_20260128_P020pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:09:05 | 20 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8936 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_74.csv | NAM_Bulk_20260128_P025pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:10:08 | 25 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10047 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_75.csv | NAM_Bulk_20260128_P030pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:12:13 | 30 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11562.5 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_76.csv | NAM_Bulk_20260128_P035pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:13:18 | 35 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 12732 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_77.csv | NAM_Bulk_20260128_P040pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:14:43 | 40 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14392.5 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_78.csv | NAM_Bulk_20260128_P045pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:15:46 | 45 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15522.5 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_79.csv | NAM_Bulk_20260128_P050pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:16:49 | 50 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15630 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_80.csv | NAM_Bulk_20260128_P055pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:18:05 | 55 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17534 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_81.csv | NAM_Bulk_20260128_P060pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:19:15 | 60 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18572 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_82.csv | NAM_Bulk_20260128_P065pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:20:27 | 65 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20981.5 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_83.csv | NAM_Bulk_20260128_P070pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:21:57 | 70 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 24487 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_84.csv | NAM_Bulk_20260128_P075pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:23:00 | 75 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 26619.5 | Valid | NAM_Bulk
20sec-power5-80and set2/SP_85.csv | NAM_Bulk_20260128_P080pct_T20s_A02_S01.csv | NAM bulk | 2026-01-28 20:24:07 | 80 | 20 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 27033 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_86.csv | NAM_Bulk_20260128_P005pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:29:38 | 5 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5426 | Low SNR | NAM_Bulk
25sec-power5-80 and set2/SP_87.csv | NAM_Bulk_20260128_P010pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:32:18 | 10 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7096.5 | Low SNR | NAM_Bulk
25sec-power5-80 and set2/SP_88.csv | NAM_Bulk_20260128_P015pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:33:41 | 15 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9128 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_89.csv | NAM_Bulk_20260128_P020pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:34:45 | 20 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11269.5 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_90.csv | NAM_Bulk_20260128_P025pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:36:11 | 25 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 12787 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_91.csv | NAM_Bulk_20260128_P030pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:37:22 | 30 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13856 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_92.csv | NAM_Bulk_20260128_P035pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:38:33 | 35 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16323 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_93.csv | NAM_Bulk_20260128_P040pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:39:44 | 40 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18762 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_94.csv | NAM_Bulk_20260128_P045pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:40:54 | 45 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20572 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_95.csv | NAM_Bulk_20260128_P050pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:42:03 | 50 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 21820 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_96.csv | NAM_Bulk_20260128_P055pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:43:07 | 55 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 26457.5 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_97.csv | NAM_Bulk_20260128_P060pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:44:52 | 60 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 27253.5 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_98.csv | NAM_Bulk_20260128_P065pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:46:58 | 65 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 27075.5 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_99.csv | NAM_Bulk_20260128_P070pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:48:28 | 70 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 31104.5 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_100.csv | NAM_Bulk_20260128_P075pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:49:30 | 75 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 33551.5 | Valid | NAM_Bulk
25sec-power5-80 and set2/SP_101.csv | NAM_Bulk_20260128_P080pct_T25s_A02_S01.csv | NAM bulk | 2026-01-28 20:50:39 | 80 | 25 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 33683 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_12.csv | NAM_Bulk_20260403_P070pct_T30s_A02_S01.csv | NAM bulk | 2026-04-03 03:52:56 | 70 | 30 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 32959 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_13.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S01.csv | NAM bulk | 2026-04-03 03:55:22 | 75 | 30 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 31741 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_14.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S02.csv | NAM bulk | 2026-04-03 03:57:55 | 75 | 30 | 2 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 31608 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_15.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S03.csv | NAM bulk | 2026-04-03 04:00:55 | 75 | 30 | 2 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 32607 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_16.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S04.csv | NAM bulk | 2026-04-03 04:06:16 | 75 | 30 | 2 | 4 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 30358.5 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_17.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S05.csv | NAM bulk | 2026-04-03 04:08:05 | 75 | 30 | 2 | 5 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 32505 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_18.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S06.csv | NAM bulk | 2026-04-03 04:12:44 | 75 | 30 | 2 | 6 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 29244.5 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_19.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S07.csv | NAM bulk | 2026-04-03 04:14:47 | 75 | 30 | 2 | 7 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 31359 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_20.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S08.csv | NAM bulk | 2026-04-03 04:20:23 | 75 | 30 | 2 | 8 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 28505.5 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_21.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S09.csv | NAM bulk | 2026-04-03 04:22:32 | 75 | 30 | 2 | 9 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 28498 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_22.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S10.csv | NAM bulk | 2026-04-03 04:24:44 | 75 | 30 | 2 | 10 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 28503 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_23.csv | NAM_Bulk_20260403_P075pct_T30s_A02_S11.csv | NAM bulk | 2026-04-03 04:34:46 | 75 | 30 | 2 | 11 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 30848.5 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_24.csv | NAM_Bulk_20260403_P080pct_T30s_A02_S01.csv | NAM bulk | 2026-04-03 04:42:13 | 80 | 30 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 33083.5 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_25.csv | NAM_Bulk_20260403_P080pct_T30s_A02_S02.csv | NAM bulk | 2026-04-03 04:47:58 | 80 | 30 | 2 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 34213 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_26.csv | NAM_Bulk_20260403_P080pct_T30s_A02_S03.csv | NAM bulk | 2026-04-03 04:51:09 | 80 | 30 | 2 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 31573.5 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_27.csv | NAM_Bulk_20260403_P080pct_T30s_A02_S04.csv | NAM bulk | 2026-04-03 04:54:33 | 80 | 30 | 2 | 4 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 34314 | Valid | NAM_Bulk
4-4-26/NAM multispot/SP_28.csv | NAM_Bulk_20260403_P080pct_T30s_A02_S05.csv | NAM bulk | 2026-04-03 04:56:23 | 80 | 30 | 2 | 5 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 32122 | Valid | NAM_Bulk
4-4-26/35sec80power/SP_1.csv | NAM_Bulk_20260403_P080pct_T35s_A02_S01.csv | NAM bulk | 2026-04-03 05:08:45 | 80 | 35 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 40189.5 | Valid | NAM_Bulk
4-4-26/35sec80power/SP_2.csv | NAM_Bulk_20260403_P080pct_T35s_A02_S02.csv | NAM bulk | 2026-04-03 05:10:24 | 80 | 35 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 41357 | Valid | NAM_Bulk
4-4-26/35sec80power/SP_3.csv | NAM_Bulk_20260403_P080pct_T35s_A02_S03.csv | NAM bulk | 2026-04-03 05:12:07 | 80 | 35 | 2 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 40154 | Valid | NAM_Bulk
4-4-26/35sec80power/SP_4.csv | NAM_Bulk_20260403_P080pct_T35s_A02_S04.csv | NAM bulk | 2026-04-03 05:13:42 | 80 | 35 | 2 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 35991 | Valid | NAM_Bulk
4-4-26/35sec80power/SP_5.csv | NAM_Bulk_20260403_P080pct_T35s_A02_S05.csv | NAM bulk | 2026-04-03 05:15:50 | 80 | 35 | 2 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 37261 | Valid | NAM_Bulk
4-4-26/35sec80power/SP_6.csv | NAM_Bulk_20260403_P080pct_T35s_A02_S06.csv | NAM bulk | 2026-04-03 05:17:52 | 80 | 35 | 2 | 6 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 34361.5 | Valid | NAM_Bulk
4-4-26/40-80 nam/SP_2.csv | NAM_Bulk_20260403_P080pct_T40s_A02_S01.csv | NAM bulk | 2026-04-03 05:25:46 | 80 | 40 | 2 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 43421.5 | Valid | NAM_Bulk
4-4-26/40-80 nam/SP_3.csv | NAM_Bulk_20260403_P080pct_T40s_A02_S02.csv | NAM bulk | 2026-04-03 05:27:52 | 80 | 40 | 2 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 44708.5 | Valid | NAM_Bulk
4-4-26/40-80 nam/SP_4.csv | NAM_Bulk_20260403_P080pct_T40s_A02_S03.csv | NAM bulk | 2026-04-03 05:31:52 | 80 | 40 | 2 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 46215.5 | Valid | NAM_Bulk
4-4-26/40-80 nam/SP_5.csv | NAM_Bulk_20260403_P080pct_T40s_A02_S04.csv | NAM bulk | 2026-04-03 05:35:24 | 80 | 40 | 2 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 45726 | Valid | NAM_Bulk
4-4-26/40-80 nam/SP_6.csv | NAM_Bulk_20260403_P080pct_T40s_A02_S05.csv | NAM bulk | 2026-04-03 05:39:17 | 80 | 40 | 2 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 41911 | Valid | NAM_Bulk
4-4-26/40-80 nam/SP_7.csv | NAM_Bulk_20260403_P080pct_T40s_A02_S06.csv | NAM bulk | 2026-04-03 05:42:59 | 80 | 40 | 2 | 6 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 41741.5 | Valid | NAM_Bulk
4-4-26/40-80 nam/SP_8.csv | NAM_Bulk_20260403_P080pct_T40s_A02_S07.csv | NAM bulk | 2026-04-03 05:44:56 | 80 | 40 | 2 | 7 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 41945.5 | Valid | NAM_Bulk
4-4-26/40-80 nam/SP_9.csv | NAM_Bulk_20260403_P080pct_T40s_A02_S08.csv | NAM bulk | 2026-04-03 05:47:37 | 80 | 40 | 2 | 8 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 45617.5 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_1.csv | NAM_Bulk_20260403_P085pct_T30s_A02_S01.csv | NAM bulk | 2026-04-03 06:04:40 | 85 | 30 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 36382.5 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_2.csv | NAM_Bulk_20260403_P085pct_T30s_A02_S02.csv | NAM bulk | 2026-04-03 06:07:40 | 85 | 30 | 2 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 34413.5 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_3.csv | NAM_Bulk_20260403_P085pct_T30s_A02_S03.csv | NAM bulk | 2026-04-03 06:09:08 | 85 | 30 | 2 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 37693.5 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_4.csv | NAM_Bulk_20260403_P085pct_T30s_A02_S04.csv | NAM bulk | 2026-04-03 06:10:58 | 85 | 30 | 2 | 4 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 38536.5 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_5.csv | NAM_Bulk_20260403_P085pct_T30s_A02_S05.csv | NAM bulk | 2026-04-03 06:12:42 | 85 | 30 | 2 | 5 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 38736 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_6.csv | NAM_Bulk_20260403_P090pct_T30s_A02_S01.csv | NAM bulk | 2026-04-03 06:15:15 | 90 | 30 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 41124.5 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_7.csv | NAM_Bulk_20260403_P090pct_T30s_A02_S02.csv | NAM bulk | 2026-04-03 06:16:47 | 90 | 30 | 2 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 42398 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_8.csv | NAM_Bulk_20260403_P090pct_T30s_A02_S03.csv | NAM bulk | 2026-04-03 06:19:20 | 90 | 30 | 2 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 40849 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_9.csv | NAM_Bulk_20260403_P095pct_T30s_A02_S01.csv | NAM bulk | 2026-04-03 06:21:41 | 95 | 30 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 42700 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_10.csv | NAM_Bulk_20260403_P095pct_T30s_A02_S02.csv | NAM bulk | 2026-04-03 06:23:28 | 95 | 30 | 2 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 42762 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_11.csv | NAM_Bulk_20260403_P095pct_T30s_A02_S03.csv | NAM bulk | 2026-04-03 06:24:54 | 95 | 30 | 2 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 42004 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_12.csv | NAM_Bulk_20260403_P095pct_T30s_A02_S04.csv | NAM bulk | 2026-04-03 06:26:41 | 95 | 30 | 2 | 4 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 41241 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_13.csv | NAM_Bulk_20260403_P095pct_T30s_A02_S05.csv | NAM bulk | 2026-04-03 06:29:31 | 95 | 30 | 2 | 5 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 44984 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_14.csv | NAM_Bulk_20260403_P100pct_T30s_A02_S01.csv | NAM bulk | 2026-04-03 06:31:24 | 100 | 30 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 44275 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_15.csv | NAM_Bulk_20260403_P100pct_T30s_A02_S02.csv | NAM bulk | 2026-04-03 06:34:48 | 100 | 30 | 2 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 45496.5 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_16.csv | NAM_Bulk_20260403_P100pct_T30s_A02_S03.csv | NAM bulk | 2026-04-03 06:36:32 | 100 | 30 | 2 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 44989.5 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_17.csv | NAM_Bulk_20260403_P100pct_T30s_A02_S04.csv | NAM bulk | 2026-04-03 06:38:22 | 100 | 30 | 2 | 4 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 41092 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_18.csv | NAM_Bulk_20260403_P080pct_T30s_A02_S06.csv | NAM bulk | 2026-04-03 06:39:42 | 80 | 30 | 2 | 6 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 33034 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_19.csv | NAM_Bulk_20260403_P085pct_T30s_A02_S06.csv | NAM bulk | 2026-04-03 06:41:07 | 85 | 30 | 2 | 6 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 33980 | Valid | NAM_Bulk
4-4-26/diffrent 3sec 80-100/SP_20.csv | NAM_Bulk_20260403_P100pct_T30s_A02_S05.csv | NAM bulk | 2026-04-03 06:43:20 | 100 | 30 | 2 | 5 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 40053 | Valid | NAM_Bulk
4-4-26/samespot35sec diffrent power/SP_1.csv | NAM_Bulk_20260403_P050pct_T35s_A02_S01.csv | NAM bulk | 2026-04-03 06:52:44 | 50 | 35 | 2 | 1 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 22733 | Valid | NAM_Bulk
4-4-26/samespot35sec diffrent power/SP_2.csv | NAM_Bulk_20260403_P080pct_T35s_A02_S07.csv | NAM bulk | 2026-04-03 06:54:12 | 80 | 35 | 2 | 7 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 34102 | Valid | NAM_Bulk
4-4-26/samespot35sec diffrent power/SP_3.csv | NAM_Bulk_20260403_P090pct_T35s_A02_S01.csv | NAM bulk | 2026-04-03 06:55:38 | 90 | 35 | 2 | 1 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 39101 | Valid | NAM_Bulk
4-4-26/samespot35sec diffrent power/SP_4.csv | NAM_Bulk_20260403_P100pct_T35s_A02_S01.csv | NAM bulk | 2026-04-03 06:57:02 | 100 | 35 | 2 | 1 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 44393 | Valid | NAM_Bulk
4-4-26/samespot35sec diffrent power/SP_5.csv | NAM_Bulk_20260403_P100pct_T35s_A02_S02.csv | NAM bulk | 2026-04-03 06:58:47 | 100 | 35 | 2 | 2 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 44909.5 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM_5P_S1.csv | NAM_Bulk_20260529_P005pct_T05s_A01_S01.csv | NAM bulk | 2026-05-29 17:24:07 | 5 | 5 | 1 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2484 | Low SNR | NAM_Bulk
29-05-26/glass slide empty/NAM_5P_S2.csv | NAM_Bulk_20260529_P005pct_T05s_A01_S02.csv | NAM bulk | 2026-05-29 17:27:00 | 5 | 5 | 1 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2524 | Low SNR | NAM_Bulk
29-05-26/glass slide empty/NAM_10P_S1.csv | NAM_Bulk_20260529_P010pct_T05s_A01_S01.csv | NAM bulk | 2026-05-29 17:35:47 | 10 | 5 | 1 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2882 | Low SNR | NAM_Bulk
29-05-26/glass slide empty/NAM_10P_S2.csv | NAM_Bulk_20260529_P010pct_T05s_A01_S02.csv | NAM bulk | 2026-05-29 17:36:45 | 10 | 5 | 1 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3011 | Low SNR | NAM_Bulk
29-05-26/glass slide empty/NAM_10P_S3.csv | NAM_Bulk_20260529_P010pct_T05s_A01_S03.csv | NAM bulk | 2026-05-29 17:37:09 | 10 | 5 | 1 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3142 | Low SNR | NAM_Bulk
29-05-26/glass slide empty/NAM_10P_S4.csv | NAM_Bulk_20260529_P010pct_T05s_A01_S04.csv | NAM bulk | 2026-05-29 17:38:04 | 10 | 5 | 1 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3001 | Low SNR | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_5AC_S1.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S01.csv | NAM bulk | 2026-05-29 17:52:23 | 70 | 10 | 5 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16030 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_5AC_S2.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S02.csv | NAM bulk | 2026-05-29 17:53:59 | 70 | 10 | 5 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16257.8 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_5AC_S3.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S03.csv | NAM bulk | 2026-05-29 17:55:31 | 70 | 10 | 5 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19107.4 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_5AC_S4.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S04.csv | NAM bulk | 2026-05-29 17:58:47 | 70 | 10 | 5 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18602.2 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_5AC_S5.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S05.csv | NAM bulk | 2026-05-29 18:00:19 | 70 | 10 | 5 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16510 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_3AC_S1.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S01.csv | NAM bulk | 2026-05-29 18:15:11 | 70 | 10 | 3 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14006.6667 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_3AC_S2.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S02.csv | NAM bulk | 2026-05-29 18:16:05 | 70 | 10 | 3 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16690 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_3AC_S3.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S03.csv | NAM bulk | 2026-05-29 18:18:03 | 70 | 10 | 3 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16771 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_3AC_S4.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S04.csv | NAM bulk | 2026-05-29 18:19:59 | 70 | 10 | 3 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16412.3333 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_70P_3AC_S5.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S05.csv | NAM bulk | 2026-05-29 18:22:33 | 70 | 10 | 3 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15882.6667 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_5AC_S1.txtr | NAM_Bulk_20260529_P080pct_T10s_A05_S01.txtr | NAM bulk | 2026-05-29 18:29:49 | 80 | 10 | 5 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17154.8 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_5AC_S2.csv | NAM_Bulk_20260529_P080pct_T10s_A05_S02.csv | NAM bulk | 2026-05-29 18:31:26 | 80 | 10 | 5 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17562 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_5AC_S4.csv | NAM_Bulk_20260529_P080pct_T10s_A05_S03.csv | NAM bulk | 2026-05-29 19:08:27 | 80 | 10 | 5 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16116 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_5AC_S5.csv | NAM_Bulk_20260529_P080pct_T10s_A05_S04.csv | NAM bulk | 2026-05-29 19:10:13 | 80 | 10 | 5 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17593.6 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_5AC_S6.csv | NAM_Bulk_20260529_P080pct_T10s_A05_S05.csv | NAM bulk | 2026-05-29 19:12:05 | 80 | 10 | 5 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17343.2 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_3AC_S1.csv | NAM_Bulk_20260529_P080pct_T10s_A03_S01.csv | NAM bulk | 2026-05-29 19:20:45 | 80 | 10 | 3 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14999 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_3AC_S2.csv | NAM_Bulk_20260529_P080pct_T10s_A03_S02.csv | NAM bulk | 2026-05-29 19:22:29 | 80 | 10 | 3 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17723.6667 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_3AC_S3.csv | NAM_Bulk_20260529_P080pct_T10s_A03_S03.csv | NAM bulk | 2026-05-29 19:23:51 | 80 | 10 | 3 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20839.3333 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_3AC_S4.csv | NAM_Bulk_20260529_P080pct_T10s_A03_S04.csv | NAM bulk | 2026-05-29 19:24:55 | 80 | 10 | 3 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17987.3333 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_3AC_S5.csv | NAM_Bulk_20260529_P080pct_T10s_A03_S05.csv | NAM bulk | 2026-05-29 19:26:09 | 80 | 10 | 3 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16874 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_3AC_S6.csv | NAM_Bulk_20260529_P080pct_T10s_A03_S06.csv | NAM bulk | 2026-05-29 19:27:21 | 80 | 10 | 3 | 6 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18849.3333 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_80P_3AC_S7.csv | NAM_Bulk_20260529_P080pct_T10s_A03_S07.csv | NAM bulk | 2026-05-29 19:29:24 | 80 | 10 | 3 | 7 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17151 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_3AC_S1.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S01.csv | NAM bulk | 2026-05-29 19:33:59 | 90 | 10 | 3 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18741.3333 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_3AC_S2.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S02.csv | NAM bulk | 2026-05-29 19:35:49 | 90 | 10 | 3 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 21654.6667 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_3AC_S3.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S03.csv | NAM bulk | 2026-05-29 19:37:24 | 90 | 10 | 3 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19108.3333 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_3AC_S4.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S04.csv | NAM bulk | 2026-05-29 19:38:37 | 90 | 10 | 3 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18035.6667 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_3AC_S5.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S05.csv | NAM bulk | 2026-05-29 19:39:57 | 90 | 10 | 3 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19578.3333 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_3AC_S6.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S06.csv | NAM bulk | 2026-05-29 19:41:58 | 90 | 10 | 3 | 6 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19075 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_5AC_S1.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S01.csv | NAM bulk | 2026-05-29 19:51:00 | 90 | 10 | 5 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19423.8 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_5AC_S2.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S02.csv | NAM bulk | 2026-05-29 19:52:23 | 90 | 10 | 5 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19562.6 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_5AC_S3.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S03.csv | NAM bulk | 2026-05-29 19:54:15 | 90 | 10 | 5 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18134.8 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_5AC_S4.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S04.csv | NAM bulk | 2026-05-29 19:56:09 | 90 | 10 | 5 | 4 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18456.8 | Valid | NAM_Bulk
29-05-26/glass slide empty/NAM-10S_90P_5AC_S5.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S05.csv | NAM bulk | 2026-05-29 19:57:50 | 90 | 10 | 5 | 5 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20913 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot1.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S06.csv | NAM bulk | 2026-05-29 20:03:10 | 90 | 10 | 5 | 6 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 22019.4 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot2.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S07.csv | NAM bulk | 2026-05-29 20:04:29 | 90 | 10 | 5 | 7 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20380.8 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot3.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S08.csv | NAM bulk | 2026-05-29 20:05:27 | 90 | 10 | 5 | 8 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 22224 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot4.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S09.csv | NAM bulk | 2026-05-29 20:06:29 | 90 | 10 | 5 | 9 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 22095.6 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot5.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S10.csv | NAM bulk | 2026-05-29 20:07:38 | 90 | 10 | 5 | 10 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20870.2 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot6.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S11.csv | NAM bulk | 2026-05-29 20:08:48 | 90 | 10 | 5 | 11 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20235 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot7.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S12.csv | NAM bulk | 2026-05-29 20:09:46 | 90 | 10 | 5 | 12 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20155.4 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot8.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S13.csv | NAM bulk | 2026-05-29 20:10:44 | 90 | 10 | 5 | 13 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19840.4 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot9.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S14.csv | NAM bulk | 2026-05-29 20:11:47 | 90 | 10 | 5 | 14 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19930.6 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot10.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S15.csv | NAM bulk | 2026-05-29 20:12:46 | 90 | 10 | 5 | 15 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19715.8 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-5ac/NAM-samespot11.csv | NAM_Bulk_20260529_P090pct_T10s_A05_S16.csv | NAM bulk | 2026-05-29 20:14:06 | 90 | 10 | 5 | 16 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19195 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot1.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S07.csv | NAM bulk | 2026-05-29 20:19:44 | 90 | 10 | 3 | 7 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 21214 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot2.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S08.csv | NAM bulk | 2026-05-29 20:20:22 | 90 | 10 | 3 | 8 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19283.3333 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot3.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S09.csv | NAM bulk | 2026-05-29 20:21:03 | 90 | 10 | 3 | 9 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19663 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot4.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S10.csv | NAM bulk | 2026-05-29 20:21:40 | 90 | 10 | 3 | 10 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 21738.3333 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot5.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S11.csv | NAM bulk | 2026-05-29 20:22:21 | 90 | 10 | 3 | 11 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 22520 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot6.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S12.csv | NAM bulk | 2026-05-29 20:23:02 | 90 | 10 | 3 | 12 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 22120 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot7.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S13.csv | NAM bulk | 2026-05-29 20:23:37 | 90 | 10 | 3 | 13 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20720 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot8.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S14.csv | NAM bulk | 2026-05-29 20:24:12 | 90 | 10 | 3 | 14 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19680.3333 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot9.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S15.csv | NAM bulk | 2026-05-29 20:24:58 | 90 | 10 | 3 | 15 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20712.6667 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot10.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S16.csv | NAM bulk | 2026-05-29 20:25:34 | 90 | 10 | 3 | 16 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 19426 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot11.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S17.csv | NAM bulk | 2026-05-29 20:26:10 | 90 | 10 | 3 | 17 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18922.3333 | Valid | NAM_Bulk
29-05-26/NAM-90p-10s-3ac/NAM-samespot12.csv | NAM_Bulk_20260529_P090pct_T10s_A03_S18.csv | NAM bulk | 2026-05-29 20:27:02 | 90 | 10 | 3 | 18 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20246.6667 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot1.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S06.csv | NAM bulk | 2026-05-29 20:49:38 | 70 | 10 | 3 | 6 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15450.6667 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot2.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S07.csv | NAM bulk | 2026-05-29 20:51:02 | 70 | 10 | 3 | 7 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16333 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot3.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S08.csv | NAM bulk | 2026-05-29 20:51:48 | 70 | 10 | 3 | 8 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16162.3333 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot4.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S09.csv | NAM bulk | 2026-05-29 20:52:33 | 70 | 10 | 3 | 9 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17714.3333 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot5.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S10.csv | NAM bulk | 2026-05-29 20:53:12 | 70 | 10 | 3 | 10 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17965 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot6.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S11.csv | NAM bulk | 2026-05-29 20:54:11 | 70 | 10 | 3 | 11 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17231.3333 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot7.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S12.csv | NAM bulk | 2026-05-29 20:55:16 | 70 | 10 | 3 | 12 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17076.3333 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot8.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S13.csv | NAM bulk | 2026-05-29 20:56:47 | 70 | 10 | 3 | 13 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16072.3333 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot9.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S14.csv | NAM bulk | 2026-05-29 20:57:32 | 70 | 10 | 3 | 14 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17056.3333 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot10.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S15.csv | NAM bulk | 2026-05-29 20:58:32 | 70 | 10 | 3 | 15 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16633.3333 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-3acc/NAM-samespot11.csv | NAM_Bulk_20260529_P070pct_T10s_A03_S16.csv | NAM bulk | 2026-05-29 20:59:12 | 70 | 10 | 3 | 16 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16610.6667 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot1.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S06.csv | NAM bulk | 2026-05-29 21:06:50 | 70 | 10 | 5 | 6 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16970.2 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot2.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S07.csv | NAM bulk | 2026-05-29 21:07:54 | 70 | 10 | 5 | 7 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18062.2 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot3.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S08.csv | NAM bulk | 2026-05-29 21:09:02 | 70 | 10 | 5 | 8 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17374.4 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot4.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S09.csv | NAM bulk | 2026-05-29 21:10:00 | 70 | 10 | 5 | 9 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16183.2 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot5.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S10.csv | NAM bulk | 2026-05-29 21:10:56 | 70 | 10 | 5 | 10 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15970.4 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot6.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S11.csv | NAM bulk | 2026-05-29 21:11:54 | 70 | 10 | 5 | 11 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15597.2 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot7.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S12.csv | NAM bulk | 2026-05-29 21:12:50 | 70 | 10 | 5 | 12 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15366.8 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot8.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S13.csv | NAM bulk | 2026-05-29 21:14:18 | 70 | 10 | 5 | 13 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14760.4 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot9.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S14.csv | NAM bulk | 2026-05-29 21:15:45 | 70 | 10 | 5 | 14 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14644.4 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot10.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S15.csv | NAM bulk | 2026-05-29 21:16:41 | 70 | 10 | 5 | 15 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14902.4 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot11.csv | NAM_Bulk_20260529_P070pct_T10s_A05_S16.csv | NAM bulk | 2026-05-29 21:18:35 | 70 | 10 | 5 | 16 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14691.8 | Valid | NAM_Bulk
29-05-26/NAM-70p-10s-5ac/NAM-samespot12.txtr | NAM_Bulk_20260529_P070pct_T10s_A05_S17.txtr | NAM bulk | 2026-05-29 21:19:32 | 70 | 10 | 5 | 17 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14505 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot1.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S01.csv | NAM bulk | 2026-05-30 01:15:11 | 70 | 10 | 5 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13882.8 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot2.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S02.csv | NAM bulk | 2026-05-30 01:17:09 | 70 | 10 | 5 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13323.6 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot3.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S03.csv | NAM bulk | 2026-05-30 01:18:45 | 70 | 10 | 5 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15167 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot4.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S04.csv | NAM bulk | 2026-05-30 01:20:14 | 70 | 10 | 5 | 4 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14680.4 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot5.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S05.csv | NAM bulk | 2026-05-30 01:24:57 | 70 | 10 | 5 | 5 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17422.4 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot6.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S06.csv | NAM bulk | 2026-05-30 01:27:31 | 70 | 10 | 5 | 6 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 17829 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot7.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S07.csv | NAM bulk | 2026-05-30 01:29:13 | 70 | 10 | 5 | 7 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 18385.2 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot8.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S08.csv | NAM bulk | 2026-05-30 01:32:52 | 70 | 10 | 5 | 8 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16513 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot9.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S09.csv | NAM bulk | 2026-05-30 01:35:24 | 70 | 10 | 5 | 9 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14855.8 | Valid | NAM_Bulk
29-05-26/70p-5ac-10s-diifrentpoint/spot10.csv | NAM_Bulk_20260530_P070pct_T10s_A05_S10.csv | NAM bulk | 2026-05-30 01:42:56 | 70 | 10 | 5 | 10 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15195.4 | Valid | NAM_Bulk
29-05-26/70p-5ac-30s-diifrentspot/spot1.csv | NAM_Bulk_20260530_P070pct_T30s_A05_S01.csv | NAM bulk | 2026-05-30 02:06:12 | 70 | 30 | 5 | 1 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 45379.2 | Valid | NAM_Bulk
29-05-26/70p-5ac-30s-diifrentspot/spot2.csv | NAM_Bulk_20260530_P070pct_T30s_A05_S02.csv | NAM bulk | 2026-05-30 02:14:47 | 70 | 30 | 5 | 2 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 49024.2 | Valid | NAM_Bulk
29-05-26/70p-5ac-30s-diifrentspot/spot3.csv | NAM_Bulk_20260530_P070pct_T30s_A05_S03.csv | NAM bulk | 2026-05-30 02:24:58 | 70 | 30 | 5 | 3 | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 47222.6 | Valid | NAM_Bulk
29-05-26/70p-60s-5ac-diifrentpoint/spot1.csv | NAM_Bulk_20260530_P070pct_T60s_A05_S01.csv | NAM bulk | 2026-05-30 04:40:46 | 70 | 60 | 5 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Valid – supporting (70%/60 s/5 acc) | NAM_Bulk
29-05-26/70p-60s-5ac-diifrentpoint/spot2.csv | NAM_Bulk_20260530_P070pct_T60s_A05_S02.csv | NAM bulk | 2026-05-30 04:47:59 | 70 | 60 | 5 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65523.2 | Valid – supporting (70%/60 s/5 acc) | NAM_Bulk
29-05-26/70p-60s-5ac-diifrentpoint/spot3.csv | NAM_Bulk_20260530_P070pct_T60s_A05_S03.csv | NAM bulk | 2026-05-30 04:54:55 | 70 | 60 | 5 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65287.2 | Valid – supporting (70%/60 s/5 acc) | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot1.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S01.csv | NAM bulk | 2026-05-30 05:18:20 | 90 | 30 | 5 | 1 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 45311 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot2.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S02.csv | NAM bulk | 2026-05-30 05:24:38 | 90 | 30 | 5 | 2 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 46251.6 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot3.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S03.csv | NAM bulk | 2026-05-30 05:27:18 | 90 | 30 | 5 | 3 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 51634 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot4.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S04.csv | NAM bulk | 2026-05-30 05:29:56 | 90 | 30 | 5 | 4 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 51343.6 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot5.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S05.csv | NAM bulk | 2026-05-30 05:32:52 | 90 | 30 | 5 | 5 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 51005.2 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot6.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S06.csv | NAM bulk | 2026-05-30 05:40:37 | 90 | 30 | 5 | 6 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 45106.4 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot7.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S07.csv | NAM bulk | 2026-05-30 05:50:11 | 90 | 30 | 5 | 7 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 44502 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot8.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S08.csv | NAM bulk | 2026-05-30 05:53:04 | 90 | 30 | 5 | 8 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 49610.4 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot9.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S09.csv | NAM bulk | 2026-05-30 05:55:43 | 90 | 30 | 5 | 9 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 52755.8 | Valid | NAM_Bulk
29-05-26/90p-30s-5acc-samespot/samespot10.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S10.csv | NAM bulk | 2026-05-30 05:59:19 | 90 | 30 | 5 | 10 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 45792.8 | Valid | NAM_Bulk
29-05-26/90p-30s-5ac-diifrentpoint/spot1.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S11.csv | NAM bulk | 2026-05-30 06:11:45 | 90 | 30 | 5 | 11 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 43778.4 | Valid | NAM_Bulk
29-05-26/90p-30s-5ac-diifrentpoint/spot2.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S12.csv | NAM bulk | 2026-05-30 06:14:53 | 90 | 30 | 5 | 12 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 47144.2 | Valid | NAM_Bulk
29-05-26/90p-30s-5ac-diifrentpoint/spot3.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S13.csv | NAM bulk | 2026-05-30 06:19:24 | 90 | 30 | 5 | 13 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 51106 | Valid | NAM_Bulk
29-05-26/90p-30s-5ac-diifrentpoint/spot4.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S14.csv | NAM bulk | 2026-05-30 06:24:42 | 90 | 30 | 5 | 14 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 43173.8 | Valid | NAM_Bulk
29-05-26/90p-30s-5ac-diifrentpoint/spot5.csv | NAM_Bulk_20260530_P090pct_T30s_A05_S15.csv | NAM bulk | 2026-05-30 06:36:34 | 90 | 30 | 5 | 15 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 45705.8 | Valid | NAM_Bulk
29-05-26/90p-60s-5ac-diffrentspot/spot1.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S01.csv | NAM bulk | 2026-05-30 06:57:47 | 90 | 60 | 5 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (abnormal dark / ripple) | Excluded
29-05-26/90p-60s-5ac-diffrentspot/spot2.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S02.csv | NAM bulk | 2026-05-30 07:03:59 | 90 | 60 | 5 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (abnormal dark / ripple) | Excluded
29-05-26/90p-60s-5ac-diffrentspot/spot3.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S03.csv | NAM bulk | 2026-05-30 07:33:11 | 90 | 60 | 5 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (abnormal dark / ripple) | Excluded
29-05-26/90p-60s-5ac-diffrentspot/spot4.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S04.csv | NAM bulk | 2026-05-30 07:40:06 | 90 | 60 | 5 | 4 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (abnormal dark / ripple) | Excluded
29-05-26/90p-60s-5ac-diffrentspot/spot5.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S05.csv | NAM bulk | 2026-05-30 07:48:34 | 90 | 60 | 5 | 5 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (abnormal dark / ripple) | Excluded
29-05-26/90p-60s-5a-samespot/samespot1.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S06.csv | NAM bulk | 2026-05-30 08:27:31 | 90 | 60 | 5 | 6 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (negative counts) | Excluded
29-05-26/90p-60s-5a-samespot/samespot2.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S07.csv | NAM bulk | 2026-05-30 08:41:09 | 90 | 60 | 5 | 7 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (negative counts) | Excluded
29-05-26/90p-60s-5a-samespot/samespot3.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S08.csv | NAM bulk | 2026-05-30 08:46:26 | 90 | 60 | 5 | 8 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (negative counts) | Excluded
29-05-26/90p-60s-5a-samespot/samespot4.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S09.csv | NAM bulk | 2026-05-30 08:53:19 | 90 | 60 | 5 | 9 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (negative counts) | Excluded
29-05-26/90p-60s-5a-samespot/samespot5.csv | NAM_Bulk_20260530_P090pct_T60s_A05_S10.csv | NAM bulk | 2026-05-30 09:01:03 | 90 | 60 | 5 | 10 | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 65535 | Excluded – distorted (negative counts) | Excluded
MB 1ugml/Trial.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-01-15 16:41:40 | 20 | 5 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5725 | Out of scope – not copied
MB 1ugml/SP_5.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-01-15 16:47:14 | 20 | 5 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5536 | Out of scope – not copied
MB 1ugml/SP_3.txtr |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-01-15 16:47:38 | 15 | 5 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4561 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_4.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:52:20 | 5 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3890.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_5.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:52:49 | 15 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5299.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_6.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:53:19 | 25 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6263.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_7.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:53:48 | 35 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7443.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_8.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:54:47 | 45 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8478 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_9.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:55:15 | 55 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9729.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_10.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:56:54 | 65 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11409.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_11.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:58:17 | 65 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11279 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_12.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:59:17 | 65 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11382 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_13.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 16:59:46 | 45 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8642 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_14.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:00:18 | 35 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7848.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_15.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:00:53 | 25 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6369.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_16.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:01:24 | 15 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5202.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_17.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:02:51 | 15 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5304.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_18.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:04:26 | 25 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6485.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_19.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:05:26 | 35 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8121.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_20.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:06:04 | 45 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9088.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_21.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:06:36 | 55 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10318.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_22.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:07:06 | 65 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 12368.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_23.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:07:33 | 85 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15081.5 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_24.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:08:02 | 95 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 15659 | Out of scope – not copied
2-2-26/0.1ug-ml/10sec-multispot-15-95 power-2set/SP_25.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:09:10 | 95 | 10 | 2 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 16255.5 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/SP_33.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:31:31 | 50 | 15 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 20038 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/SP_34.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 17:32:19 | 30 | 15 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13755.5 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/di_1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:46:08 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7281 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/tw__1_3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:46:08 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7281 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/di_2_2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:48:00 | 10 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 23080.5 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/tw_3__11.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:48:00 | 10 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 23080.5 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/tw_1_3.txtr |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:51:31 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10208 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/tw_11_3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:53:28 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10135 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/ww_1_5.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:55:41 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9018.5 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/di_2_6.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:57:56 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8466 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/tw2_7.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 18:59:49 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8925 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/ww2_8.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 19:01:54 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8276 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/bw_2_9.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 19:03:37 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6946 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/di_3_10.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 19:06:20 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9465.5 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/tw_3_11.txtr |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 19:08:33 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10051 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/ww_3_12.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 19:10:29 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10245.5 | Out of scope – not copied
2-2-26/0.1ug-ml/substrate plain/New folder/bw_3_13.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-02-02 19:12:16 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7773.5 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:49:26 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3869.6667 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:50:13 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3950.6667 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:51:24 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3877.6667 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP4.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:52:23 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3865.6667 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP5.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:53:25 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3906.6667 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP6.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:54:36 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3934.6667 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP7.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:56:43 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3900.3333 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP8.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:58:23 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3889.3333 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP9.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 05:59:21 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4040.3333 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP10.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:00:35 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3898.6667 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP11.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:01:35 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3735 | Out of scope – not copied
5-6-2026/agno3 empty wafer data multispot-10s-3ac-20p multispot/SP12.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:03:21 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3683 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_19.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:20:23 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3964.6667 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_20.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:21:16 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4090 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_21.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:22:24 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4032 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_22.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:23:15 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4094 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_23.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:23:55 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4453 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_24.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:24:39 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4337.6667 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_25.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:25:39 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4263.3333 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_26.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:26:30 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4094.6667 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_27.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:27:20 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3879.6667 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_28.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:28:27 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3892 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_29.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:30:06 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3753.3333 | Out of scope – not copied
5-6-2026/pure silicone 10s-20p-3ac multispot/SP_30.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-06 06:31:33 | 20 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3699.3333 | Out of scope – not copied
New folder/SP_1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 12:46:44 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1685.6667 | Out of scope – not copied
New folder/SP_2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 12:47:15 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1670 | Out of scope – not copied
New folder/SP_3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 12:47:51 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1730.6667 | Out of scope – not copied
New folder/diffrenspot-r6g.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 12:48:35 | 5 | 5 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1858.3333 | Out of scope – not copied
10-6-26/Nam-diifrentspot/nam-1mgml -5sec-5p-3ac-diffpoint.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 12:57:14 | 5 | 5 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1803.6667 | Out of scope – not copied
10-6-26/Nam-diifrentspot/nam-1mgml-agno3 -5sec-5p-3ac-diffpoint.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 12:58:09 | 5 | 5 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1780.3333 | Out of scope – not copied
10-6-26/Nam-diifrentspot/nam-1mgml-agno3 -5sec-5p-3ac-diffpoint-3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 12:58:59 | 5 | 5 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1821.6667 | Out of scope – not copied
10-6-26/Nam-diifrentspot/nam-1mgml-agno3 -10sec-5p-3ac-diffpoint-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:31:29 | 5 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2464 | Out of scope – not copied
10-6-26/Nam-diifrentspot/nam-1mgml-agno3 -10sec-5p-3ac-diffpoint-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:33:18 | 5 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2607.6667 | Out of scope – not copied
10-6-26/Nam-diifrentspot/nam-1mgml-agno3 -10sec-5p-3ac-diffpoint-3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:35:55 | 5 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2658 | Out of scope – not copied
10-6-26/r6g-1mgml-10sec-5p-3ac/r6g-1mgml-agno3 -10sec-5p-3ac-samespott-1.txtr |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:42:57 | 5 | 10 | 3 |  | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2879.3333 | Out of scope – not copied
10-6-26/r6g-1mgml-10sec-5p-3ac/r6g-1mgml-agno3 -10sec-5p-3ac-samespott-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:44:55 | 5 | 10 | 3 |  | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2777 | Out of scope – not copied
10-6-26/r6g-1mgml-10sec-5p-3ac/r6g-1mgml-agno3 -10sec-5p-3ac-samespott-3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:45:47 | 5 | 10 | 3 |  | same spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2895.6667 | Out of scope – not copied
10-6-26/r6g-1mgml-10sec-5p-3ac/r6g-1mgml-agno3 -10sec-5p-3ac-diffrentspott-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:46:34 | 5 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2860.3333 | Out of scope – not copied
10-6-26/r6g-1mgml-10sec-5p-3ac/r6g-1mgml-agno3 -10sec-5p-3ac-diffrentspott-2-edge.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:48:19 | 5 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2746.3333 | Out of scope – not copied
10-6-26/r6g-1mgml-10sec-5p-3ac/r6g-1mgml-agno3 -10sec-5p-3ac-diffrent spott-1center.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-10 13:49:43 | 5 | 10 | 3 |  | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2735.6667 | Out of scope – not copied
10-6-26/nam-4m/SP_1 5p-5s-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:15:38 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2063 | Out of scope – not copied
10-6-26/nam-4m/SP_2-5p-5s-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:16:24 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1749.6667 | Out of scope – not copied
10-6-26/nam-4m/SP_3-5p-5s-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:17:15 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3251.6667 | Out of scope – not copied
10-6-26/nam-4m/SP_4-5p-5s-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:18:23 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1965.6667 | Out of scope – not copied
10-6-26/nam-4m/SP_8-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:25:28 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2572 | Out of scope – not copied
10-6-26/nam-4m/SP_10-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:28:50 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3138.6667 | Out of scope – not copied
10-6-26/nam-4m/SP_11-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:29:56 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2304.6667 | Out of scope – not copied
10-6-26/nam-4m/SP_12-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:30:56 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2265.6667 | Out of scope – not copied
10-6-26/nam-4m/SP_13-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:33:49 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3467.3333 | Out of scope – not copied
10-6-26/nam-4m/SP_14-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:35:06 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6050.6667 | Out of scope – not copied
10-6-26/nam-4m/SP_15-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:36:11 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4438.3333 | Out of scope – not copied
10-6-26/nam-4m/SP_16-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:40:48 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7291.4 | Out of scope – not copied
10-6-26/nam-4m/SP_17-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:42:33 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4162.2 | Out of scope – not copied
10-6-26/nam-4m/SP_18-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:43:48 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3450 | Out of scope – not copied
10-6-26/r6g-4m/SP_19-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:47:10 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3080.8 | Out of scope – not copied
10-6-26/r6g-4m/SP_20-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:48:58 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3101.6 | Out of scope – not copied
10-6-26/r6g-4m/SP_21-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:50:21 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3094.2 | Out of scope – not copied
10-6-26/r6g-4m/SP_22-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:53:19 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2949.6667 | Out of scope – not copied
10-6-26/r6g-4m/SP_23-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:54:10 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3081 | Out of scope – not copied
10-6-26/r6g-4m/SP_24-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 04:55:09 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3081 | Out of scope – not copied
10-6-26/r6g-4m/SP_25-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:00:45 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2369.6667 | Out of scope – not copied
10-6-26/r6g-4m/SP_26-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:01:56 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2342.6667 | Out of scope – not copied
10-6-26/r6g-4m/SP_27-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:03:17 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2403.6667 | Out of scope – not copied
10-6-26/r6g-4m/SP_28-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:05:56 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1725 | Out of scope – not copied
10-6-26/r6g-4m/SP_29-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:06:48 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1733.3333 | Out of scope – not copied
10-6-26/r6g-4m/SP_30-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:07:45 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1756 | Out of scope – not copied
10-6-26/nam 6m/SP_36-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:23:39 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3482.6667 | Out of scope – not copied
10-6-26/nam 6m/SP_37-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:26:14 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1786.6667 | Out of scope – not copied
10-6-26/nam 6m/SP_38-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:27:55 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1833.3333 | Out of scope – not copied
10-6-26/nam 6m/SP_39-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:30:48 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2560.3333 | Out of scope – not copied
10-6-26/nam 6m/SP_40-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:32:00 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2634.3333 | Out of scope – not copied
10-6-26/nam 6m/SP_41-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:33:39 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3525.3333 | Out of scope – not copied
10-6-26/nam 6m/SP_42-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:35:52 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4611.6667 | Out of scope – not copied
10-6-26/nam 6m/SP_43-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:37:03 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5631.3333 | Out of scope – not copied
10-6-26/nam 6m/SP_44-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:38:15 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5611.6667 | Out of scope – not copied
10-6-26/nam 6m/SP_45-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:40:57 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5655.4 | Out of scope – not copied
10-6-26/nam 6m/SP_46-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:42:22 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5596 | Out of scope – not copied
10-6-26/nam 6m/SP_48-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:45:38 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4947 | Out of scope – not copied
10-6-26/r6g 6m/SP_49-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:50:13 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2921.2 | Out of scope – not copied
10-6-26/r6g 6m/SP_50-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:51:42 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2981 | Out of scope – not copied
10-6-26/r6g 6m/SP_52-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:55:16 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3090 | Out of scope – not copied
10-6-26/r6g 6m/SP_53-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 05:59:28 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2903.6667 | Out of scope – not copied
10-6-26/r6g 6m/SP_54-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:00:28 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3032.3333 | Out of scope – not copied
10-6-26/r6g 6m/SP_55-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:01:27 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3062.6667 | Out of scope – not copied
10-6-26/r6g 6m/SP_56-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:04:10 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2286.3333 | Out of scope – not copied
10-6-26/r6g 6m/SP_57-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:05:15 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2307.3333 | Out of scope – not copied
10-6-26/r6g 6m/SP_58-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:06:33 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2376 | Out of scope – not copied
10-6-26/r6g 6m/SP_59-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:09:56 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1735.6667 | Out of scope – not copied
10-6-26/r6g 6m/SP_60-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:10:33 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1734 | Out of scope – not copied
10-6-26/r6g 6m/SP_61-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:11:15 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1750.6667 | Out of scope – not copied
10-6-26/nam-5m/SP_64-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:17:54 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2143.3333 | Out of scope – not copied
10-6-26/nam-5m/SP_65-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:18:55 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2157.3333 | Out of scope – not copied
10-6-26/nam-5m/SP_66-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:19:47 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1920.3333 | Out of scope – not copied
10-6-26/nam-5m/SP_67-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:22:38 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2727.6667 | Out of scope – not copied
10-6-26/nam-5m/SP_68-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:24:08 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3284.6667 | Out of scope – not copied
10-6-26/nam-5m/SP_69-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:25:18 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3503.3333 | Out of scope – not copied
10-6-26/nam-5m/SP_70-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:32:21 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4516.6667 | Out of scope – not copied
10-6-26/nam-5m/SP_71-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:33:32 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5093.3333 | Out of scope – not copied
10-6-26/nam-5m/SP_75-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:38:51 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4531.3333 | Out of scope – not copied
10-6-26/nam-5m/SP_78-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:43:23 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4460 | Out of scope – not copied
10-6-26/nam-5m/SP_79-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:50:14 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4002.8 | Out of scope – not copied
10-6-26/nam-5m/SP_80-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:51:59 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4077 | Out of scope – not copied
10-6-26/nam-5m/SP_81-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:53:31 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4560.6 | Out of scope – not copied
10-6-26/r6g-5m/SP_82-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 06:58:32 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2959.4 | Out of scope – not copied
10-6-26/r6g-5m/SP_83-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:00:15 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2902.4 | Out of scope – not copied
10-6-26/r6g-5m/SP_84-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:01:49 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2769.6 | Out of scope – not copied
10-6-26/r6g-5m/SP_85-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:05:02 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2973.6667 | Out of scope – not copied
10-6-26/r6g-5m/SP_86-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:06:43 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2934.3333 | Out of scope – not copied
10-6-26/r6g-5m/SP_87-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:07:33 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2819.3333 | Out of scope – not copied
10-6-26/r6g-5m/SP_88-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:10:59 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2651.6667 | Out of scope – not copied
10-6-26/r6g-5m/SP_89-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:13:07 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2481 | Out of scope – not copied
10-6-26/r6g-5m/SP_90-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:14:11 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2504.6667 | Out of scope – not copied
10-6-26/r6g-5m/SP_91-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:16:18 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1870.6667 | Out of scope – not copied
10-6-26/r6g-5m/SP_92-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:17:02 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1837 | Out of scope – not copied
10-6-26/r6g-5m/SP_93-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:17:46 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1808.3333 | Out of scope – not copied
10-6-26/nam-3m/SP_94-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:20:51 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2472.3333 | Out of scope – not copied
10-6-26/nam-3m/SP_95-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:22:13 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1837 | Out of scope – not copied
10-6-26/nam-3m/SP_96-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:23:54 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1879.3333 | Out of scope – not copied
10-6-26/nam-3m/SP_97-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:26:52 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2779 | Out of scope – not copied
10-6-26/nam-3m/SP_98-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:28:07 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2640.6667 | Out of scope – not copied
10-6-26/nam-3m/SP_99-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:29:07 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2522.3333 | Out of scope – not copied
10-6-26/nam-3m/SP_100-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:31:38 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3045 | Out of scope – not copied
10-6-26/nam-3m/SP_101-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:32:44 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3024.3333 | Out of scope – not copied
10-6-26/nam-3m/SP_102-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:33:47 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2985.3333 | Out of scope – not copied
10-6-26/nam-3m final/SP_103-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:36:44 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3062.8 | Out of scope – not copied
10-6-26/nam-3m final/SP_104-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:38:07 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2870.2 | Out of scope – not copied
10-6-26/nam-3m final/SP_105-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:39:36 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2902.8 | Out of scope – not copied
10-6-26/nam-3m final/SP_106-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:54:49 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2882.3333 | Out of scope – not copied
10-6-26/nam-3m final/SP_107-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:55:43 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2841.3333 | Out of scope – not copied
10-6-26/nam-3m final/SP_108-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 07:57:12 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3166.3333 | Out of scope – not copied
10-6-26/nam-3m final/SP_109-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:00:29 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2700.3333 | Out of scope – not copied
10-6-26/nam-3m final/SP_110-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:01:21 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2686.6667 | Out of scope – not copied
10-6-26/nam-3m final/SP_111-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:02:30 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2647.3333 | Out of scope – not copied
10-6-26/nam-3m final/SP_112-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:04:29 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1875.6667 | Out of scope – not copied
10-6-26/nam-3m final/SP_113-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:05:30 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1825.6667 | Out of scope – not copied
10-6-26/nam-3m final/SP_114-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:06:18 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1926.6667 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_11-5-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:09:17 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1909 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_116-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:10:07 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1989.3333 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_117-5s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:10:50 | 5 | 5 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2020.3333 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_118-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:14:43 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3087.3333 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_119-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:16:24 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2926 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_120-10s-5p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:17:28 | 5 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2859.3333 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_121-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:21:30 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2998.6667 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_122-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:22:48 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3279.6667 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_123-10s-10p-3ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:23:42 | 10 | 10 | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3193 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_124-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:27:22 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3046.4 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_125-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:28:39 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3014.6 | Out of scope – not copied
10-6-26/r6g-3mfinal/SP_126-10s-10p-5ac.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-11 08:29:59 | 10 | 10 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2949.8 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_7.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:48:18 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2798.2 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_8.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:50:20 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3242 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_9.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:51:18 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2181.2 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_10.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:52:12 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3056 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_11.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:53:00 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2667.8 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_12.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:53:59 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2744 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_13.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:55:20 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2528.8 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_14.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:56:20 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3015 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_15.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:57:28 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3072.4 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_16.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:58:42 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2974.2 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_17.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 00:59:51 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2556.8 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_18.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:01:05 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1994 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_19.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:02:11 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2323.6 | Out of scope – not copied
17-06-2026/3m-10ml-hf-5s-10p-5ac/SP_20.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:03:36 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 1935.4 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_21.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:12:03 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2876 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_22.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:13:33 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2658.6 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_23.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:14:41 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3002.6 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_24.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:16:24 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2935.6 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_25.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:17:32 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2609.4 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_26.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:18:45 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2345 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_27.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:19:58 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2833.8 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_28.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:21:06 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2464.6 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_29.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:22:11 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3289.4 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_30.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:23:11 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2870.2 | Out of scope – not copied
17-06-2026/3m-5ml-hf-5s-10p-5ac/SP_31.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:24:20 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2435.4 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_32.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:29:21 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2976.6 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_33.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:30:30 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3709.6 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_34.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:31:24 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3305.4 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_35.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:32:22 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3205.2 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_36.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:33:23 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2766.6 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_37.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:34:32 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2685.8 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_38.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:35:40 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3201.6 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_39.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:37:17 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2701 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_40.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:38:26 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3343.8 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_41.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:39:28 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2824 | Out of scope – not copied
17-06-2026/4m-10ml-hf-5s-10p-5ac/SP_42.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:40:34 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2960 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_43.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:45:28 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3072.6 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_44.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:46:33 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2877.4 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_45.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:48:38 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2722.8 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_46.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:50:00 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3091.8 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_47.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:51:04 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3019.6 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_48.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:52:12 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2776.4 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_49.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:53:25 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3145.6 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_50.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:54:24 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3079 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_51.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:55:24 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3293.4 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_52.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:56:13 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3050.4 | Out of scope – not copied
17-06-2026/4m-5ml-hf-5s-10p-5ac/SP_53.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 01:57:05 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3098 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_54.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:00:43 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4118.2 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_55.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:01:40 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3662.8 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_56.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:02:39 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4183.6 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_57.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:04:06 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2717.6 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_58.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:05:59 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2982 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_59.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:06:55 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3189.4 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_60.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:07:53 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4268.8 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_61.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:08:54 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3581 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_62.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:10:19 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3820.4 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_63.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:11:15 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3877.8 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_64.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:13:06 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3763.8 | Out of scope – not copied
17-06-2026/5m-10ml-hf-5s-10p-5ac/SP_65.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:14:06 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2499.2 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_66.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:17:06 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3922.4 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_67.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:18:18 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3686.4 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_68.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:19:34 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4009.4 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_69.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:20:48 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5006.2 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_70.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:21:38 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4797 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_71.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:23:22 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4804.8 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_72.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:24:42 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4174 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_73.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:25:39 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4448.4 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_74.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:26:26 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3459.2 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_75.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:27:20 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3640.4 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_76.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:28:02 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4569.2 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_77.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:28:48 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3358.6 | Out of scope – not copied
17-06-2026/5m-5ml-hf-5s-10p-5ac/SP_78.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:29:46 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4939.8 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_79.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:32:50 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3509 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_80.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:34:00 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5203.2 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_81.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:34:47 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3402 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_82.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:35:47 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3355.2 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_83.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:36:47 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4713 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_84.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:37:37 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4407.2 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_85.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:38:25 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3163.2 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_86.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:39:25 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3572.8 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_87.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:40:13 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4730.2 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_88.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:40:58 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3251 | Out of scope – not copied
17-06-2026/6m-10ml-hf-5s-10p-5ac/SP_89.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:41:42 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3631.8 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_90.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:44:06 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3695 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_91.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:44:55 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3868.2 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_92.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:45:34 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3993.8 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_93.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:46:17 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3967.6 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_94.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:47:03 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3989.8 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_95.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:47:46 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4164.8 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_96.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:48:28 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3809.4 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_97.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:49:24 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3784.2 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_98.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:50:02 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3880.2 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_99.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:50:44 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3967.2 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_100.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:51:27 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4038.8 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_101.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:52:16 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4009.8 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_102.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:52:58 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3954.8 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/SP_103.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-17 02:53:38 | 10 | 5 | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3669.4 | Out of scope – not copied
17-06-2026/6m-5ml-hf-5s-10p-5ac/CSHM_1mW_Bare_1st.txtr |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-06-24 18:44:25 | 1 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 2811.5 | Out of scope – not copied
04-07-2026/SP_13.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 10:58:19 | 10 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 12768.5 | Out of scope – not copied
04-07-2026/SP_14.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 11:02:26 | 10 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 12345 | Out of scope – not copied
04-07-2026/SP_15.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 11:05:23 | 10 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9243.5 | Out of scope – not copied
04-07-2026/SP_16.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 11:06:52 | 10 | 20 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8348.5 | Out of scope – not copied
04-07-2026/r6g/SP_1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 11:33:09 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4553 | Out of scope – not copied
04-07-2026/r6g/SP_2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 12:40:23 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7310 | Out of scope – not copied
04-07-2026/r6g/SP_3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 12:41:58 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7437.25 | Out of scope – not copied
04-07-2026/r6g/SP_4.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 12:44:32 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8768.25 | Out of scope – not copied
04-07-2026/r6g/SP_5.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 13:13:18 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10594.25 | Out of scope – not copied
04-07-2026/r6g/SP_6.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 13:17:48 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9845 | Out of scope – not copied
04-07-2026/substarte/empty subastrate1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:22:41 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10178.5 | Out of scope – not copied
04-07-2026/substarte/empty subastrate2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:24:54 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9720.25 | Out of scope – not copied
04-07-2026/substarte/empty subastrate3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:26:34 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9965.25 | Out of scope – not copied
04-07-2026/substarte/empty subastrate.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:29:04 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9893.25 | Out of scope – not copied
04-07-2026/substarte/empty subastrate4.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:31:41 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9618.5 | Out of scope – not copied
04-07-2026/namon5hf6mgagno3/1mgml-nam-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:44:39 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6892 | Out of scope – not copied
04-07-2026/namon5hf6mgagno3/1mgml-nam-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:46:55 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7572.75 | Out of scope – not copied
04-07-2026/namon5hf6mgagno3/1mgml-nam-3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:48:44 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7437.5 | Out of scope – not copied
04-07-2026/namon5hf6mgagno3/1mgml-nam-4.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:50:05 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4802 | Out of scope – not copied
04-07-2026/namon5hf6mgagno3/0.1mgml-nam-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:53:19 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7415.25 | Out of scope – not copied
04-07-2026/namon5hf6mgagno3/0.1mgml-nam-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:54:49 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 5099 | Out of scope – not copied
04-07-2026/namon5hf6mgagno3/0.1mgml-nam-3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 15:56:21 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8353.75 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1nM-1sec-10-10p-4acc-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 16:04:25 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9530.25 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1nM-1sec-10-10p-4acc-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 16:08:01 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9839.25 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-10nM-1sec-10-10p-4acc-3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 16:36:12 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10926.5 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-10nM-1sec-10-10p-4acc-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 16:42:01 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9998.25 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-10nM-1sec-10-10p-4acc-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 16:43:20 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10498.25 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-100nM-1sec-10-10p-4acc-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 16:54:07 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9963.25 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-100nM-1sec-10-10p-4acc-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 16:55:01 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9831.75 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-100nM-1sec-10-10p-4acc-3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 16:57:48 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9837 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1microM-1sec-10-10p-4acc-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:14:40 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10847.25 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1microM-1sec-10-10p-4acc-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:21:44 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11249.5 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1microM-1sec-10-10p-4acc-3.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:23:50 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10430.5 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1microM-1sec-10-10p-4acc-4.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:25:04 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10355 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1microM-1sec-10-10p-4acc-5.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:26:59 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10358.25 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-10microM-1sec-10-10p-4acc-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:28:57 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6554.75 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-10microM-1sec-10-10p-4acc-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:30:33 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10156.5 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1mM-1sec-10-10p-4acc-1.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:33:06 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8105.75 | Out of scope – not copied
04-07-2026/r6g-sensor/r6g-1mM-1sec-10-10p-4acc-2.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 17:37:39 | 10 | 10 | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13405.75 | Out of scope – not copied
04-07-2026/rawr6g/SP_17.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:02:20 | 1 | 1 | 1 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14196 | Out of scope – not copied
04-07-2026/rawr6g/SP_18.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:02:31 | 1 | 1 | 1 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14099 | Out of scope – not copied
04-07-2026/rawr6g/SP_19.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:02:37 | 1 | 1 | 1 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13916 | Out of scope – not copied
04-07-2026/rawr6g/SP_20.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:03:56 | 1 | 1 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 13498.5 | Out of scope – not copied
04-07-2026/rawr6g/SP_21.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:05:14 | 2 | 2 | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 35745.5 | Out of scope – not copied
04-07-2026/rawr6g/SP_22.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:05:58 | 2 |  | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9117.5 | Out of scope – not copied
04-07-2026/rawr6g/SP_23.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:06:07 | 2 |  | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9064 | Out of scope – not copied
04-07-2026/rawr6g/SP_24.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:06:14 | 2 |  | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8999.5 | Out of scope – not copied
04-07-2026/rawr6g/SP_25.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:06:19 | 2 |  | 2 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8988 | Out of scope – not copied
04-07-2026/rawr6g/SP_26.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:06:31 | 2 |  | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9008 | Out of scope – not copied
04-07-2026/rawr6g/SP_27.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:06:43 | 2 |  | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9043 | Out of scope – not copied
04-07-2026/rawr6g/SP_28.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:06:51 | 2 |  | 3 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8960.6667 | Out of scope – not copied
04-07-2026/rawr6g/SP_29.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:07:06 | 2 |  | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9027.25 | Out of scope – not copied
04-07-2026/rawr6g/SP_30.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:07:14 | 2 |  | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8956 | Out of scope – not copied
04-07-2026/rawr6g/SP_31.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:07:25 | 1 |  | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6533.75 | Out of scope – not copied
04-07-2026/rawr6g/SP_32.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:07:36 | 1 |  | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6508 | Out of scope – not copied
04-07-2026/rawr6g/SP_33.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:07:49 | 1 |  | 4 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6540.75 | Out of scope – not copied
04-07-2026/rawr6g/SP_34.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:08:00 | 1 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6445.2 | Out of scope – not copied
04-07-2026/rawr6g/SP_35.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:08:20 | 1 |  | 10 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6378.8 | Out of scope – not copied
04-07-2026/rawr6g/SP_36.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:08:47 | 2 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8315.8 | Out of scope – not copied
04-07-2026/rawr6g/SP_37.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:08:56 | 2 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8276.6 | Out of scope – not copied
04-07-2026/rawr6g/SP_38.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:09:05 | 2 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8269.2 | Out of scope – not copied
04-07-2026/rawr6g/SP_39.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:09:17 | 2 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8126.4 | Out of scope – not copied
04-07-2026/rawr6g/SP_40.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:09:31 | 2 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8078.8 | Out of scope – not copied
04-07-2026/rawr6g/SP_41.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:09:41 | 2 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8112.2 | Out of scope – not copied
04-07-2026/rawr6g/SP_42.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:10:02 | 2 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8053.8 | Out of scope – not copied
04-07-2026/rawr6g/SP_43.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:10:14 | 2 |  | 7 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7931.7143 | Out of scope – not copied
04-07-2026/rawr6g/SP_44.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:10:37 | 3 |  | 7 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8847.1429 | Out of scope – not copied
04-07-2026/rawr6g/SP_45.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:10:49 | 3 |  | 7 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8771.7143 | Out of scope – not copied
04-07-2026/rawr6g/SP_46.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:11:02 | 3 |  | 7 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8696.1429 | Out of scope – not copied
04-07-2026/rawr6g/SP_47.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:11:26 | 1 |  | 10 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3646.5 | Out of scope – not copied
04-07-2026/rawr6g/SP_48.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:11:55 | 1 |  | 10 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3663.5 | Out of scope – not copied
04-07-2026/rawr6g/SP_49.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:12:11 | 1 |  | 10 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3631.4 | Out of scope – not copied
04-07-2026/rawr6g/SP_50.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:12:50 | 1 |  | 15 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3640.8 | Out of scope – not copied
04-07-2026/rawr6g/SP_51.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:13:12 | 1 |  | 15 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 3578.4 | Out of scope – not copied
04-07-2026/rawr6g/SP_52.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:14:24 | 10 |  | 15 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8951.5333 | Out of scope – not copied
04-07-2026/rawr6g/SP_53.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:14:48 | 10 |  | 15 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8718.1333 | Out of scope – not copied
04-07-2026/rawr6g/SP_54.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:16:43 | 5 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4900.8 | Out of scope – not copied
04-07-2026/rawr6g/SP_55.csv |  | Not bulk – SERS/substrate experiment (out of scope) | 2026-07-04 20:16:51 | 5 |  | 5 |  | not stated | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 4891.8 | Out of scope – not copied
9-2-2026/20sec20power-variable location LTA/SP_2.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S01.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:22:17 | 20 | 20 | 2 | 1 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8546.5 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_3.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S02.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:23:39 | 20 | 20 | 2 | 2 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 10289.5 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_4.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S03.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:25:15 | 20 | 20 | 2 | 3 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9935.5 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_5.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S04.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:26:46 | 20 | 20 | 2 | 4 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9612.5 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_6.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S05.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:28:27 | 20 | 20 | 2 | 5 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 7964 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_7.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S06.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:30:04 | 20 | 20 | 2 | 6 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9437.5 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_8.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S07.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:33:12 | 20 | 20 | 2 | 7 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 6216 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_10.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S08.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:37:12 | 20 | 20 | 2 | 8 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9514 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_11.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S09.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:39:26 | 20 | 20 | 2 | 9 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 11871 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_12.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S10.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:40:59 | 20 | 20 | 2 | 10 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 9434.5 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_13.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S11.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:42:35 | 20 | 20 | 2 | 11 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 8014 | Contaminated / misidentified | Excluded
9-2-2026/20sec20power-variable location LTA/SP_14.csv | Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S12.csv | Unknown – labelled LTA, polystyrene-like | 2026-02-09 20:43:51 | 20 | 20 | 2 | 12 | different spot | not recorded in file (probe_name empty) | 784.92 | -46.2 | 2841.9 | 2048 | 14278.5 | Contaminated / misidentified | Excluded

## 2_QC
Sheet 2 — Quality control (all bulk / blank / suspect spectra)
File name | Original file | Class | SNR | Max raw counts (full range) | Max raw counts (400–1800) | Max dark-subtracted (400–1800) | Min dark-subtracted (400–1800) | Baseline level (median, counts) | Noise SD (counts) | Saturated (any pixel) | Saturated in 400–1800 | Spikes removed | r to class mean (glass zone excl.) | QC decision | Reason | Used in reference
Glass_Blank_20260529_P005pct_T05s_A01_S01.csv | 29-05-26/glass slide empty/glassslide (4).csv | GLASS | 54.63 | 7681 | 7681 | 6548 | 2118 | 2755.63 | 29.6 | No | No | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
Glass_Blank_20260529_P005pct_T05s_A01_S02.csv | 29-05-26/glass slide empty/glassslide (5).csv | GLASS | 56.74 | 7824 | 7824 | 6695 | 2127 | 2768.58 | 30.25 | No | No | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
Glass_Blank_20260529_P005pct_T05s_A01_S03.csv | 29-05-26/glass slide empty/glassslide (1).csv | GLASS | 50.63 | 7710 | 7710 | 6613 | 2092 | 2753.07 | 32.66 | No | No | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
Glass_Blank_20260529_P005pct_T05s_A01_S04.csv | 29-05-26/glass slide empty/glassslide (2).csv | GLASS | 56.64 | 7934 | 7934 | 6831 | 2193 | 2840.01 | 30.22 | No | No | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
Glass_Blank_20260529_P005pct_T05s_A01_S05.csv | 29-05-26/glass slide empty/glassslide (3).csv | GLASS | 56.46 | 7979 | 7979 | 6854 | 2160 | 2830.48 | 30.98 | No | No | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
LPS_Bulk_20260413_P010pct_T10s_A02_S01.csv | 14-426/lps/SP_1.csv | LPS | 9.3 | 4299.5 | 2985 | 1776.5 | 306 | 849.17 | 18.45 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T10s_A02_S02.csv | 14-426/lps/SP_2.csv | LPS | 4.91 | 4406 | 2078 | 428 | 128.5 | 278.31 | 17.81 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T10s_A02_S03.csv | 14-426/lps/SP_3.csv | LPS | 4.21 | 4521 | 2056 | 430.5 | 105 | 261.15 | 19.38 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T10s_A05_S01.csv | 13-4-26/lps/20-20-5/SP_1.csv | LPS | 9.32 | 5043 | 3364.6 | 634.4 | 225.6 | 358.89 | 14.28 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T10s_A05_S02.csv | 13-4-26/lps/20-20-5/SP_2.csv | LPS | 4.34 | 4839.2 | 3277.4 | 440 | 149.6 | 288.76 | 14.34 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T60s_A30_S01.csv | 13-4-26/lps/20-20-5/SP_6.csv | LPS | 14.75 | 21942.7 | 11756.5 | 1499.63 | -21.5 | 400.17 | 22.24 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T10s_A02_S01.csv | 13-4-26/lps/20-20-5/SP_4.csv | LPS | 4.66 | 6072.5 | 3393 | 556.8 | 205.1 | 393.72 | 17.22 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T10s_A05_S01.csv | 13-4-26/lps/20-20-5/SP_3.csv | LPS | 5.73 | 6535.8 | 3446 | 639.6 | 234.4 | 439.02 | 15.53 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T30s_A50_S01.csv | 14-426/lps/SP_5.csv | LPS | 72.82 | 17087.98 | 4193.54 | 1511.2 | 119.82 | 829.3 | 4.62 | No | No | 1 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T30s_A60_S01.csv | 14-426/lps/SP_4.csv | LPS | 76.75 | 17163.1 | 4909.35 | 1522.37 | 648.68 | 877.65 | 4.3 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T60s_A30_S01.csv | 13-4-26/lps/20-20-5/SP_5.csv | LPS | 29.12 | 29874.0667 | 12219.4333 | 2185.47 | 441.43 | 1225.55 | 17.38 | No | No | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T60s_A30_S02.csv | 13-4-26/lps/20-20-5/SP_8.csv | LPS | 31.51 | 29108.0667 | 12976.2333 | 3634.67 | 587.47 | 1260.35 | 16.87 | No | No | 1 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T60s_A50_S01.csv | 13-4-26/lps/20-20-5/SP_7.csv | LPS | 19.8 | 31704.56 | 10280.74 | 2075.22 | -1497.26 | 210.71 | 28.46 | No | No | 4 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T60s_A50_S02.csv | 13-4-26/lps/20-20-5/SP_9.csv | LPS | 20.73 | 30282.72 | 10605.9 | 4836.24 | 1529.24 | 3191.45 | 45.7 | No | No | 2 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LTA_Bulk_20260403_P030pct_T05s_A10_S01.csv | 4-4-26/5sec,30,50 power 10 set/SP_12.csv | LTA | 6.57 | 3119 | 1902.1 | 251.5 | 68.7 | 162.3 | 8.19 | No | No | 0 | 0.88 | Low SNR | SNR 6.7 < 10 at low power / short integration; genuine LTA spectrum used only for parameter comparison. | No
LTA_Bulk_20260403_P040pct_T10s_A10_S01.csv | 4-4-26/LTA/10sec,diffrenpower,10set/SP_9.csv | LTA | 11.28 | 5842.5 | 2797.1 | 582.8 | 196.1 | 376.02 | 8.8 | No | No | 0 | 0.97 | Valid |  | No
LTA_Bulk_20260403_P050pct_T05s_A10_S01.csv | 4-4-26/5sec,30,50 power 10 set/SP_13.csv | LTA | 8.81 | 4454.3 | 1969.9 | 340.9 | 105 | 217.02 | 7.93 | No | No | 0 | 0.92 | Low SNR | SNR 8.4 < 10 at low power / short integration; genuine LTA spectrum used only for parameter comparison. | No
LTA_Bulk_20260403_P050pct_T10s_A10_S01.csv | 4-4-26/LTA/10sec,diffrenpower,10set/SP_7.csv | LTA | 15.5 | 8096.3 | 2999.9 | 874.7 | 280.7 | 555.52 | 8.68 | No | No | 0 | 0.98 | Valid |  | No
LTA_Bulk_20260403_P050pct_T30s_A15_S01.csv | 4-4-26/LTA/30sec50power15set/SP_6.csv | LTA | 13.7 | 20222.6667 | 9292 | 4192.6 | 988.33 | 2858.17 | 55.05 | No | No | 0 | 0.45 | Abnormal - review | Distorted spectrum (r = 0.45 vs LTA class mean, glass region excluded) with low SNR; not explained by power/time. Not used as reference. | No
LTA_Bulk_20260403_P060pct_T10s_A10_S01.csv | 4-4-26/LTA/10sec,diffrenpower,10set/SP_8.csv | LTA | 17.27 | 8643.6 | 3054.9 | 903.7 | 303.8 | 589.91 | 8.92 | No | No | 0 | 0.98 | Valid |  | No
LTA_Bulk_20260403_P060pct_T45s_A02_S01.csv | 4-4-26/LTA/45sec diffrent powrset2/SP_5.csv | LTA | 23.38 | 34145 | 10105 | 5979.5 | 1818.5 | 3591.98 | 31.44 | No | No | 0 | 0.99 | Valid |  | No
LTA_Bulk_20260403_P070pct_T10s_A10_S01.csv | 4-4-26/LTA/10sec,diffrenpower,10set/SP_11.csv | LTA | 21.58 | 9892.5 | 3150 | 987.8 | 344.3 | 679.38 | 8.44 | No | No | 0 | 0.98 | Valid |  | No
LTA_Bulk_20260403_P080pct_T45s_A02_S01.csv | 4-4-26/LTA/45sec diffrent powrset2/SP_1.csv | LTA | 29.4 | 47406 | 22619 | 20457 | 5087 | 11424.55 | 38.97 | No | No | 0 | 0.94 | Valid |  | Yes
LTA_Bulk_20260403_P080pct_T45s_A02_S02.csv | 4-4-26/LTA/45sec diffrent powrset2/SP_4.csv | LTA | 27.97 | 43075 | 11787 | 8298 | 2553 | 5026.24 | 34.68 | No | No | 0 | 0.99 | Valid |  | Yes
LTA_Bulk_20260403_P090pct_T45s_A02_S01.csv | 4-4-26/LTA/45sec diffrent powrset2/SP_2.csv | LTA | 27.38 | 48531.5 | 16295.5 | 13706 | 3866 | 8013.24 | 41.19 | No | No | 0 | 0.98 | Valid |  | Yes
LTA_Bulk_20260403_P095pct_T45s_A02_S01.csv | 4-4-26/LTA/45sec diffrent powrset2/SP_3.csv | LTA | 29.34 | 53846 | 14848 | 12274.5 | 3583.5 | 7536.68 | 40.79 | No | No | 0 | 0.97 | Valid |  | Yes
NAM_Bulk_20260128_P005pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_21.csv | NAM | 2.54 | 1808 | 1363.5 | 109 | -14 | 28.58 | 13.58 | No | No | 0 | 0.52 | Low SNR | SNR 2.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_37.csv | NAM | 3.58 | 2896.5 | 1782 | 148.5 | -4.5 | 58.73 | 13.42 | No | No | 0 | 0.69 | Low SNR | SNR 3.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_4.csv | NAM | 5.07 | 3698.5 | 3450.5 | 445.5 | 67 | 221.68 | 15.84 | No | No | 0 | 0.72 | Low SNR | SNR 5.0 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_54.csv | NAM | 5.25 | 3583.5 | 2615.5 | 193.5 | 15.5 | 84.9 | 14.01 | No | No | 0 | 0.82 | Low SNR | SNR 5.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_70.csv | NAM | 5 | 4606 | 2628.5 | 219.5 | 16 | 97.06 | 14.86 | No | No | 0 | 0.85 | Low SNR | SNR 4.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_86.csv | NAM | 6.82 | 5426 | 3058 | 270.5 | 24.5 | 116.9 | 16.63 | No | No | 0 | 0.84 | Low SNR | SNR 6.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_22.csv | NAM | 3.43 | 2131.5 | 1387 | 126 | -8 | 42.48 | 13.83 | No | No | 0 | 0.58 | Low SNR | SNR 3.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_38.csv | NAM | 4.64 | 3332 | 1799 | 180.5 | 9 | 76.21 | 12.95 | No | No | 0 | 0.78 | Low SNR | SNR 4.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_5.csv | NAM | 6.32 | 4679.5 | 3461 | 519 | 74 | 229.18 | 16.99 | No | No | 0 | 0.82 | Low SNR | SNR 6.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T15s_A02_S02.csv | NAM-15secs and power15-75-2sets/SP_6.csv | NAM | 5.88 | 4519.5 | 3427.5 | 546 | 88 | 229.58 | 17.87 | No | No | 0 | 0.89 | Low SNR | SNR 6.0 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T15s_A02_S03.csv | 15sec-power5-80andset2/SP_55.csv | NAM | 6.68 | 4739.5 | 2613 | 245 | 7.5 | 110.71 | 14.2 | No | No | 0 | 0.88 | Low SNR | SNR 6.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_71.csv | NAM | 6.25 | 5726 | 3173.5 | 650.5 | 41 | 127.14 | 15.89 | No | No | 0 | 0.91 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_87.csv | NAM | 7.76 | 7096.5 | 3098.5 | 371.5 | 55 | 162.92 | 17.33 | No | No | 0 | 0.94 | Low SNR | SNR 7.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_23.csv | NAM | 2.81 | 2505 | 1397.5 | 146 | 0 | 57.33 | 14.03 | No | No | 0 | 0.62 | Low SNR | SNR 2.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_39.csv | NAM | 5.83 | 4248.5 | 1833.5 | 318 | 39 | 107.23 | 13.87 | No | No | 0 | 0.89 | Low SNR | SNR 5.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T15s_A02_S01.csv | 15sec-power5-80andset2/SP_56.csv | NAM | 7.74 | 5977 | 2713.5 | 344 | 56.5 | 154.75 | 14.83 | No | No | 0 | 0.94 | Low SNR | SNR 7.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_72.csv | NAM | 9.26 | 7541 | 3300.5 | 777.5 | 72.5 | 190.05 | 15.21 | No | No | 0 | 0.96 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_88.csv | NAM | 11.15 | 9128 | 3165 | 522 | 88 | 228.81 | 17.63 | No | No | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P020pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_24.csv | NAM | 4.54 | 2890.5 | 1383 | 191.5 | 12.5 | 76.35 | 14.74 | No | No | 0 | 0.68 | Low SNR | SNR 4.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_40.csv | NAM | 6.49 | 4923 | 1868.5 | 445 | 38.5 | 131.6 | 13.99 | No | No | 0 | 0.89 | Low SNR | SNR 6.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T10s_A02_S02.csv | 10sec-power5-80-2set/SP_41.csv | NAM | 6.33 | 4946.5 | 1884 | 456.5 | 32.5 | 132.39 | 14.99 | No | No | 0 | 0.92 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_7.csv | NAM | 9.66 | 6796.5 | 3598.5 | 760.5 | 146 | 369.31 | 17.65 | No | No | 0 | 0.93 | Low SNR | SNR 9.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_57.csv | NAM | 9.96 | 7166.5 | 2745 | 423.5 | 72.5 | 197.42 | 14.48 | No | No | 0 | 0.95 | Low SNR | SNR 9.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_73.csv | NAM | 11.29 | 8936 | 3331.5 | 808.5 | 98.5 | 227.8 | 15.39 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P020pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_89.csv | NAM | 14.79 | 11269.5 | 3241 | 1069.5 | 117 | 275.39 | 15.77 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P025pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_25.csv | NAM | 4.58 | 3273 | 1408 | 198.5 | 11.5 | 85.33 | 13.44 | No | No | 0 | 0.78 | Low SNR | SNR 4.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P025pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_42.csv | NAM | 8.8 | 5634 | 1913.5 | 481.5 | 44 | 157.25 | 14.07 | No | No | 0 | 0.93 | Low SNR | SNR 8.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P025pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_8.csv | NAM | 9.51 | 8157.5 | 3612 | 899.5 | 179 | 443.94 | 17.8 | No | No | 0 | 0.94 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P025pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_58.csv | NAM | 11.08 | 8174 | 2774 | 461.5 | 87.5 | 214.2 | 15.2 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P025pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_74.csv | NAM | 11.56 | 10047 | 3345 | 822 | 99.5 | 255.4 | 17.34 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P025pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_90.csv | NAM | 15.77 | 12787 | 3220.5 | 1068 | 138 | 309 | 18.31 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P030pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_26.csv | NAM | 5.24 | 3542.5 | 1441.5 | 220 | 11.5 | 97.39 | 13.5 | No | No | 0 | 0.77 | Low SNR | SNR 5.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P030pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_43.csv | NAM | 8.4 | 6547 | 1934 | 463 | 72 | 182.16 | 14.89 | No | No | 0 | 0.94 | Low SNR | SNR 8.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P030pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_9.csv | NAM | 12.14 | 9484 | 3638 | 996 | 216 | 477.94 | 17.37 | No | No | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P030pct_T15s_A02_S02.csv | NAM-15secs and power15-75-2sets/SP_10.csv | NAM | 11.37 | 9304 | 3557.5 | 894.5 | 154.5 | 386.93 | 17.54 | No | No | 0 | 0.94 | Valid |  | No
NAM_Bulk_20260128_P030pct_T15s_A02_S03.csv | 15sec-power5-80andset2/SP_59.csv | NAM | 12.74 | 9024.5 | 2814 | 544.5 | 110 | 252.53 | 14.74 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P030pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_75.csv | NAM | 13.08 | 11562.5 | 3431.5 | 908.5 | 137 | 304.1 | 17.77 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P030pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_91.csv | NAM | 17.59 | 13856 | 3352 | 1140.5 | 167.5 | 369.99 | 18.01 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P035pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_27.csv | NAM | 6.32 | 3907.5 | 1430 | 260.5 | 18.5 | 110.04 | 12.22 | No | No | 0 | 0.84 | Low SNR | SNR 5.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P035pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_44.csv | NAM | 10.04 | 6950 | 1947 | 500.5 | 76.5 | 192.87 | 15.22 | No | No | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P035pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_11.csv | NAM | 11.1 | 10041 | 3466.5 | 824.5 | 129 | 331.99 | 18.69 | No | No | 0 | 0.91 | Valid |  | No
NAM_Bulk_20260128_P035pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_60.csv | NAM | 14.79 | 10229 | 2841.5 | 588 | 120.5 | 277.05 | 14.31 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P035pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_76.csv | NAM | 13.7 | 12732 | 3448 | 925 | 153 | 340.73 | 18.45 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P035pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_92.csv | NAM | 18.27 | 16323 | 3374 | 1173 | 195.5 | 412.46 | 18.83 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P040pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_28.csv | NAM | 6.34 | 4158 | 1454.5 | 265.5 | 8.5 | 121.11 | 13.29 | No | No | 0 | 0.87 | Low SNR | SNR 6.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P040pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_45.csv | NAM | 10.04 | 7227.5 | 1976.5 | 515.5 | 96.5 | 217.21 | 14.61 | No | No | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P040pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_12.csv | NAM | 12.52 | 11643.5 | 3499.5 | 991 | 125.5 | 342.84 | 19.72 | No | No | 0 | 0.9 | Valid |  | No
NAM_Bulk_20260128_P040pct_T15s_A02_S02.csv | NAM-15secs and power15-75-2sets/SP_13.csv | NAM | 12.65 | 11266.5 | 3487.5 | 924 | 105.5 | 313.79 | 19.53 | No | No | 0 | 0.9 | Valid |  | No
NAM_Bulk_20260128_P040pct_T15s_A02_S03.csv | 15sec-power5-80andset2/SP_61.csv | NAM | 14.79 | 11825.5 | 2887.5 | 690 | 148 | 322.18 | 15.54 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P040pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_77.csv | NAM | 15.17 | 14392.5 | 3500 | 977 | 183.5 | 397.59 | 19.57 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P040pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_93.csv | NAM | 20.55 | 18762 | 3421 | 1241 | 222.5 | 475.07 | 18.55 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P045pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_29.csv | NAM | 6.3 | 4561 | 1484.5 | 298 | 24.5 | 132.58 | 13.74 | No | No | 0 | 0.89 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P045pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_46.csv | NAM | 11.55 | 8021 | 1997.5 | 546 | 87.5 | 243.68 | 14.86 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P045pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_14.csv | NAM | 13.06 | 12142.5 | 3454 | 910.5 | 133.5 | 317.54 | 20.24 | No | No | 0 | 0.91 | Valid |  | No
NAM_Bulk_20260128_P045pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_62.csv | NAM | 16.9 | 12868.5 | 2924 | 763.5 | 167 | 348.9 | 15.11 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P045pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_78.csv | NAM | 19.91 | 15522.5 | 3554 | 1031 | 195.5 | 439.27 | 16.02 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P045pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_94.csv | NAM | 22.48 | 20572 | 3484.5 | 1240 | 255 | 509.35 | 19.59 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P050pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_30.csv | NAM | 7.03 | 4879.5 | 1508.5 | 312.5 | 45.5 | 144.69 | 14.16 | No | No | 0 | 0.91 | Low SNR | SNR 7.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P050pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_47.csv | NAM | 10.66 | 8373 | 2017 | 565.5 | 111 | 260.02 | 16.88 | No | No | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P050pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_15.csv | NAM | 13.09 | 13073 | 3458.5 | 912 | 105 | 277.52 | 21.37 | No | No | 0 | 0.84 | Review - higher noise | Correct NAM band pattern but r = 0.84 vs class mean; extra noise structure. | No
NAM_Bulk_20260128_P050pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_63.csv | NAM | 16.12 | 13233.5 | 2938 | 769 | 162 | 359.54 | 16.07 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P050pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_79.csv | NAM | 20.34 | 15630 | 3545.5 | 1022.5 | 203 | 441.83 | 17.36 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P050pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_95.csv | NAM | 22.48 | 21820 | 3509 | 1270.5 | 261.5 | 535 | 19.79 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P055pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_31.csv | NAM | 7.43 | 5236 | 1475.5 | 347 | 48 | 151.68 | 14.1 | No | No | 0 | 0.92 | Low SNR | SNR 7.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P055pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_48.csv | NAM | 12.37 | 9363 | 2309.5 | 628 | 123 | 285.41 | 15.71 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P055pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_17.csv | NAM | 18.64 | 14103 | 3384.5 | 1184.5 | 259.5 | 542.9 | 16.7 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P055pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_64.csv | NAM | 18.02 | 14549 | 3007.5 | 915.5 | 203 | 414.45 | 17.3 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P055pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_80.csv | NAM | 20.23 | 17534 | 3618 | 1095 | 246.5 | 502.07 | 19.06 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P055pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_96.csv | NAM | 26.29 | 26457.5 | 3628 | 1385.5 | 303 | 622.64 | 20.59 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P060pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_32.csv | NAM | 8.06 | 5553.5 | 1506 | 386.5 | 48 | 167.74 | 13.62 | No | No | 0 | 0.92 | Low SNR | SNR 8.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P060pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_49.csv | NAM | 13.58 | 10106.5 | 2349.5 | 663.5 | 134 | 308.33 | 15.63 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P060pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_16.csv | NAM | 18.66 | 15251 | 3528 | 1335 | 307.5 | 646.44 | 18.04 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P060pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_65.csv | NAM | 19.7 | 15617.5 | 3051 | 966.5 | 223 | 454.14 | 16.19 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P060pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_81.csv | NAM | 23.37 | 18572 | 3441.5 | 1177.5 | 247 | 534.18 | 17.79 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P060pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_97.csv | NAM | 26.54 | 27253.5 | 3668.5 | 1453.5 | 322.5 | 646.9 | 20.82 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P065pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_33.csv | NAM | 8.61 | 5946.5 | 1509.5 | 384 | 52 | 172.98 | 14.17 | No | No | 0 | 0.92 | Low SNR | SNR 8.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P065pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_50.csv | NAM | 14.21 | 10655.5 | 2360.5 | 1061 | 150 | 324.53 | 15.89 | No | No | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P065pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_18.csv | NAM | 20.35 | 15625.5 | 3429 | 1263.5 | 274 | 566.6 | 16.14 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P065pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_66.csv | NAM | 21.1 | 16585.5 | 3026.5 | 971.5 | 232 | 457.43 | 16.11 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P065pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_82.csv | NAM | 22.56 | 20981.5 | 3136 | 1233 | 276.5 | 567.5 | 19.58 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P065pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_98.csv | NAM | 28.3 | 27075.5 | 3659.5 | 1524.5 | 337.5 | 672.86 | 20.83 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P070pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_34.csv | NAM | 8.3 | 6262 | 1529.5 | 402.5 | 66 | 184.98 | 14.3 | No | No | 0 | 0.95 | Low SNR | SNR 8.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P070pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_51.csv | NAM | 16.27 | 11417.5 | 2398 | 750.5 | 161 | 346.3 | 15.05 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P070pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_19.csv | NAM | 20.47 | 16923.5 | 2993.5 | 1359 | 305 | 633.59 | 17.68 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P070pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_67.csv | NAM | 21.55 | 17121.5 | 3061 | 1041 | 230.5 | 482.97 | 16.74 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P070pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_83.csv | NAM | 26.26 | 24487 | 3197.5 | 1362 | 299.5 | 636.42 | 18.87 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P070pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_99.csv | NAM | 30.99 | 31104.5 | 3758 | 1642.5 | 374.5 | 739.65 | 20.91 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P075pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_35.csv | NAM | 9.53 | 6664 | 1555.5 | 444.5 | 78.5 | 200.98 | 13.97 | No | No | 0 | 0.94 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P075pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_52.csv | NAM | 16.46 | 12369.5 | 2416.5 | 799 | 176.5 | 377.79 | 16.17 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P075pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_20.csv | NAM | 20.46 | 17505.5 | 3023.5 | 1406 | 317 | 645.62 | 18.53 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P075pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_68.csv | NAM | 22.1 | 18758.5 | 3122 | 1116 | 259 | 522.79 | 17.79 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P075pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_84.csv | NAM | 28.4 | 26619.5 | 3307 | 1502 | 336 | 687.01 | 19.28 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P075pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_100.csv | NAM | 31.76 | 33551.5 | 3820 | 1784.5 | 378 | 796.45 | 21.55 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P080pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_36.csv | NAM | 10.56 | 6996.5 | 1572.5 | 484 | 83 | 219.11 | 12.69 | No | No | 0 | 0.95 | Valid |  | No
NAM_Bulk_20260128_P080pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_53.csv | NAM | 18.12 | 13500 | 2430 | 848 | 185.5 | 392.66 | 15.1 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P080pct_T15s_A02_S01.csv | 15sec-power5-80andset2/SP_69.csv | NAM | 23.28 | 18726 | 2760.5 | 1120.5 | 236 | 516.03 | 16.71 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P080pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_85.csv | NAM | 29.21 | 27033 | 3313.5 | 1552 | 361 | 710.29 | 19.63 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P080pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_101.csv | NAM | 32.47 | 33683 | 3807.5 | 1787.5 | 418 | 810.89 | 21.89 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P050pct_T35s_A02_S01.csv | 4-4-26/samespot35sec diffrent power/SP_1.csv | NAM | 21.27 | 22733 | 4740 | 1522.5 | 439.5 | 836.61 | 23.11 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P070pct_T30s_A02_S01.csv | 4-4-26/NAM multispot/SP_12.csv | NAM | 37.13 | 32959 | 7378.5 | 5711.5 | 1334.5 | 2868.94 | 32.47 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S01.csv | 4-4-26/NAM multispot/SP_13.csv | NAM | 29.53 | 31741 | 5225.5 | 3150.5 | 765.5 | 1548.93 | 25.31 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S02.csv | 4-4-26/NAM multispot/SP_14.csv | NAM | 24.64 | 31608 | 4928.5 | 2550.5 | 684 | 1345.98 | 25.07 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S03.csv | 4-4-26/NAM multispot/SP_15.csv | NAM | 24.54 | 32607 | 7200.5 | 5421.5 | 1409 | 2991.01 | 29.61 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S04.csv | 4-4-26/NAM multispot/SP_16.csv | NAM | 37.04 | 30358.5 | 7529.5 | 5838 | 1333.5 | 2942.88 | 30.4 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S05.csv | 4-4-26/NAM multispot/SP_17.csv | NAM | 30.14 | 32505 | 5583.5 | 3641 | 1013 | 1972.18 | 24.81 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S06.csv | 4-4-26/NAM multispot/SP_18.csv | NAM | 35.21 | 29244.5 | 7523.5 | 5581 | 1415 | 2759.86 | 27.32 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S07.csv | 4-4-26/NAM multispot/SP_19.csv | NAM | 23.95 | 31359 | 6361.5 | 2913 | 929 | 1857 | 25.43 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S08.csv | 4-4-26/NAM multispot/SP_20.csv | NAM | 21.71 | 28505.5 | 6393 | 3122 | 1027.5 | 1898.96 | 26.86 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S09.csv | 4-4-26/NAM multispot/SP_21.csv | NAM | 23.8 | 28498 | 6772 | 3758.5 | 1076.5 | 2074.84 | 31.22 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S10.csv | 4-4-26/NAM multispot/SP_22.csv | NAM | 34.75 | 28503 | 6898.5 | 4234 | 982 | 2064.8 | 28.17 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S11.csv | 4-4-26/NAM multispot/SP_23.csv | NAM | 24.11 | 30848.5 | 5794.5 | 3444.5 | 1415.5 | 2376.19 | 25.71 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S01.csv | 4-4-26/NAM multispot/SP_24.csv | NAM | 24.86 | 33083.5 | 5575 | 3345 | 1357.5 | 2162.55 | 27.13 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S02.csv | 4-4-26/NAM multispot/SP_25.csv | NAM | 19.51 | 34213 | 5710.5 | 3444.5 | 1240.5 | 2239.96 | 24.97 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S03.csv | 4-4-26/NAM multispot/SP_26.csv | NAM | 25.31 | 31573.5 | 5242 | 3053.5 | 778 | 1556.14 | 23.74 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S04.csv | 4-4-26/NAM multispot/SP_27.csv | NAM | 25.22 | 34314 | 5353 | 3078 | 917.5 | 1771.01 | 25.15 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S05.csv | 4-4-26/NAM multispot/SP_28.csv | NAM | 38.41 | 32122 | 5911 | 4229.5 | 992 | 2029.2 | 26.58 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S06.csv | 4-4-26/diffrent 3sec 80-100/SP_18.csv | NAM | 26.86 | 33034 | 4566 | 1970 | 445.5 | 956.06 | 21.53 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S01.csv | 4-4-26/35sec80power/SP_1.csv | NAM | 27.62 | 40189.5 | 5915 | 3155.5 | 1044.5 | 1946.02 | 26.62 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S02.csv | 4-4-26/35sec80power/SP_2.csv | NAM | 28.55 | 41357 | 5722 | 2977 | 996.5 | 1810.33 | 26.13 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S03.csv | 4-4-26/35sec80power/SP_3.csv | NAM | 31.33 | 40154 | 6289.5 | 4134.5 | 1101.5 | 2146.57 | 27.01 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S04.csv | 4-4-26/35sec80power/SP_4.csv | NAM | 39.65 | 35991 | 6193 | 4145.5 | 961.5 | 1975.28 | 28.46 | No | No | 1 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S05.csv | 4-4-26/35sec80power/SP_5.csv | NAM | 29.29 | 37261 | 6969 | 5393 | 2131 | 2864.72 | 30.47 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S06.csv | 4-4-26/35sec80power/SP_6.csv | NAM | 29.38 | 34361.5 | 5754.5 | 3076.5 | 965 | 1810.06 | 24.89 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S07.csv | 4-4-26/samespot35sec diffrent power/SP_2.csv | NAM | 29.91 | 34102 | 5443.5 | 2320 | 666 | 1237.39 | 26.26 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S01.csv | 4-4-26/40-80 nam/SP_2.csv | NAM | 33.06 | 43421.5 | 7796.5 | 3337.5 | 1113 | 2009.68 | 25.43 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S02.csv | 4-4-26/40-80 nam/SP_3.csv | NAM | 30.06 | 44708.5 | 7720.5 | 3537.5 | 943 | 1826.33 | 26.41 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S03.csv | 4-4-26/40-80 nam/SP_4.csv | NAM | 28.89 | 46215.5 | 7760 | 3506 | 1042 | 2009.09 | 29.1 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S04.csv | 4-4-26/40-80 nam/SP_5.csv | NAM | 32.11 | 45726 | 8000.5 | 3896 | 1125.5 | 2123.3 | 30.26 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S05.csv | 4-4-26/40-80 nam/SP_6.csv | NAM | 46.07 | 41911 | 7587.5 | 5512.5 | 1678.5 | 3125.94 | 30.5 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S06.csv | 4-4-26/40-80 nam/SP_7.csv | NAM | 50.76 | 41741.5 | 8600 | 6759 | 1620.5 | 3236.74 | 35.25 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S07.csv | 4-4-26/40-80 nam/SP_8.csv | NAM | 33.65 | 41945.5 | 6023 | 3198.5 | 886.5 | 1673.18 | 26.89 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S08.csv | 4-4-26/40-80 nam/SP_9.csv | NAM | 52.25 | 45617.5 | 8289 | 6448 | 1549 | 3093.77 | 33.27 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S01.csv | 4-4-26/diffrent 3sec 80-100/SP_1.csv | NAM | 23.67 | 36382.5 | 5124.5 | 2891 | 715.5 | 1446.65 | 26.38 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S02.csv | 4-4-26/diffrent 3sec 80-100/SP_2.csv | NAM | 31.29 | 34413.5 | 5614.5 | 3769 | 948.5 | 1899.79 | 27.39 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S03.csv | 4-4-26/diffrent 3sec 80-100/SP_3.csv | NAM | 38.76 | 37693.5 | 6402.5 | 4745.5 | 1177.5 | 2372.38 | 30.44 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S04.csv | 4-4-26/diffrent 3sec 80-100/SP_4.csv | NAM | 45.01 | 38536.5 | 7080 | 5423 | 1314.5 | 2615.99 | 31.75 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S05.csv | 4-4-26/diffrent 3sec 80-100/SP_5.csv | NAM | 41.63 | 38736 | 5622 | 3862 | 925 | 1881.78 | 30.2 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S06.csv | 4-4-26/diffrent 3sec 80-100/SP_19.csv | NAM | 24.3 | 33980 | 4515 | 1976 | 463 | 953.88 | 24.64 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P090pct_T30s_A02_S01.csv | 4-4-26/diffrent 3sec 80-100/SP_6.csv | NAM | 50.92 | 41124.5 | 5657.5 | 4010 | 945 | 1890.84 | 27.58 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P090pct_T30s_A02_S02.csv | 4-4-26/diffrent 3sec 80-100/SP_7.csv | NAM | 54.44 | 42398 | 6867.5 | 5228 | 1303.5 | 2591.58 | 30.4 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P090pct_T30s_A02_S03.csv | 4-4-26/diffrent 3sec 80-100/SP_8.csv | NAM | 45.31 | 40849 | 5479 | 3699 | 885.5 | 1776.21 | 27.08 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P090pct_T35s_A02_S01.csv | 4-4-26/samespot35sec diffrent power/SP_3.csv | NAM | 31.93 | 39101 | 6471 | 2576.5 | 740 | 1320.7 | 27.28 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S01.csv | 4-4-26/diffrent 3sec 80-100/SP_9.csv | NAM | 52.13 | 42700 | 5636.5 | 3997 | 904 | 1801.08 | 27.73 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S02.csv | 4-4-26/diffrent 3sec 80-100/SP_10.csv | NAM | 46.75 | 42762 | 6477.5 | 4838 | 1231 | 2397.56 | 26.55 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S03.csv | 4-4-26/diffrent 3sec 80-100/SP_11.csv | NAM | 41.71 | 42004 | 5653.5 | 3984 | 989 | 1969.97 | 28.95 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S04.csv | 4-4-26/diffrent 3sec 80-100/SP_12.csv | NAM | 44.36 | 41241 | 5497 | 3629.5 | 880.5 | 1782.64 | 26.68 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S05.csv | 4-4-26/diffrent 3sec 80-100/SP_13.csv | NAM | 48.77 | 44984 | 6126 | 4478 | 1023.5 | 2107.29 | 32.46 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S01.csv | 4-4-26/diffrent 3sec 80-100/SP_14.csv | NAM | 57.91 | 44275 | 5813 | 4181.5 | 913.5 | 1936.47 | 27.49 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S02.csv | 4-4-26/diffrent 3sec 80-100/SP_15.csv | NAM | 52.46 | 45496.5 | 5746 | 4127 | 900 | 1862.16 | 30.76 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S03.csv | 4-4-26/diffrent 3sec 80-100/SP_16.csv | NAM | 53.48 | 44989.5 | 5712 | 4080.5 | 889.5 | 1839.57 | 28.98 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S04.csv | 4-4-26/diffrent 3sec 80-100/SP_17.csv | NAM | 33.04 | 41092 | 4952.5 | 2651.5 | 672 | 1315.05 | 22.93 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S05.csv | 4-4-26/diffrent 3sec 80-100/SP_20.csv | NAM | 33.91 | 40053 | 4997.5 | 2765 | 772.5 | 1458.46 | 23.82 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P100pct_T35s_A02_S01.csv | 4-4-26/samespot35sec diffrent power/SP_4.csv | NAM | 37.72 | 44393 | 6485 | 2590.5 | 727 | 1356.77 | 24.49 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T35s_A02_S02.csv | 4-4-26/samespot35sec diffrent power/SP_5.csv | NAM | 35.53 | 44909.5 | 6466.5 | 2572 | 735 | 1333.01 | 26.1 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P005pct_T05s_A01_S01.csv | 29-05-26/glass slide empty/NAM_5P_S1.csv | NAM | 2.46 | 2484 | 1862 | 204 | 13 | 84.64 | 25.59 | No | No | 0 | 0.43 | Low SNR | SNR 2.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P005pct_T05s_A01_S02.csv | 29-05-26/glass slide empty/NAM_5P_S2.csv | NAM | 3.11 | 2524 | 1904 | 286 | -33 | 110.15 | 23.43 | No | No | 0 | 0.46 | Low SNR | SNR 2.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P010pct_T05s_A01_S01.csv | 29-05-26/glass slide empty/NAM_10P_S1.csv | NAM | 2.72 | 2882 | 1918 | 320 | 20 | 128.24 | 24.5 | No | No | 0 | 0.51 | Low SNR | SNR 2.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P010pct_T05s_A01_S02.csv | 29-05-26/glass slide empty/NAM_10P_S2.csv | NAM | 2.76 | 3011 | 1894 | 372 | 45 | 176.61 | 26.88 | No | No | 0 | 0.59 | Low SNR | SNR 2.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P010pct_T05s_A01_S03.csv | 29-05-26/glass slide empty/NAM_10P_S3.csv | NAM | 2.7 | 3142 | 1892 | 367 | 57 | 181.05 | 23.39 | No | No | 0 | 0.52 | Low SNR | SNR 2.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P010pct_T05s_A01_S04.csv | 29-05-26/glass slide empty/NAM_10P_S4.csv | NAM | 2.75 | 3001 | 1929 | 337 | 31 | 160.59 | 24.12 | No | No | 0 | 0.38 | Low SNR | SNR 2.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P070pct_T10s_A03_S01.csv | 29-05-26/glass slide empty/NAM-10S_70P_3AC_S1.csv | NAM | 13.29 | 14006.6667 | 2911.3333 | 819.67 | 302.67 | 512.45 | 16.37 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S02.csv | 29-05-26/glass slide empty/NAM-10S_70P_3AC_S2.csv | NAM | 18.91 | 16690 | 3061.3333 | 1159.67 | 347.33 | 638.13 | 14.65 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S03.csv | 29-05-26/glass slide empty/NAM-10S_70P_3AC_S3.csv | NAM | 16.65 | 16771 | 3216.3333 | 1276.33 | 282.67 | 700.54 | 17.32 | No | No | 0 | 0.93 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S04.csv | 29-05-26/glass slide empty/NAM-10S_70P_3AC_S4.csv | NAM | 12.32 | 16412.3333 | 2504 | 777.33 | 187 | 408.7 | 15.12 | No | No | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S05.csv | 29-05-26/glass slide empty/NAM-10S_70P_3AC_S5.csv | NAM | 12.07 | 15882.6667 | 2444 | 847.67 | 176.67 | 324.38 | 18.83 | No | No | 0 | 0.88 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S06.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot1.csv | NAM | 17.84 | 15450.6667 | 2235 | 687 | 221.67 | 367.73 | 13.95 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S07.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot2.csv | NAM | 16.62 | 16333 | 2211.3333 | 697.33 | 223.67 | 366.78 | 15.95 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S08.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot3.csv | NAM | 18.45 | 16162.3333 | 2193.6667 | 695.33 | 210.33 | 350.41 | 13.56 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S09.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot4.csv | NAM | 17.75 | 17714.3333 | 2200.3333 | 693 | 213 | 359.35 | 14.19 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S10.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot5.csv | NAM | 19.52 | 17965 | 2227.3333 | 705.33 | 227.67 | 363.28 | 13.57 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S11.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot6.csv | NAM | 17.49 | 17231.3333 | 2059 | 669.33 | 16.33 | 343.64 | 14.44 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S12.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot7.csv | NAM | 16.64 | 17076.3333 | 2050 | 637.67 | 211 | 337.68 | 14.44 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S13.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot8.csv | NAM | 17.19 | 16072.3333 | 2047 | 628.33 | 201.33 | 330.77 | 13.67 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S14.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot9.csv | NAM | 16.96 | 17056.3333 | 2065.6667 | 647.33 | 217.33 | 339.03 | 14.71 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S15.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot10.csv | NAM | 16.81 | 16633.3333 | 2059 | 628 | 210 | 333.74 | 14.05 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S16.csv | 29-05-26/NAM-70p-10s-3acc/NAM-samespot11.csv | NAM | 16.09 | 16610.6667 | 2049.6667 | 649.67 | 214 | 335.5 | 15.02 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S01.csv | 29-05-26/glass slide empty/NAM-10S_70P_5AC_S1.csv | NAM | 23.04 | 16030 | 3132.4 | 1043.6 | 333.8 | 611.38 | 11.94 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S02.csv | 29-05-26/glass slide empty/NAM-10S_70P_5AC_S2.csv | NAM | 17.37 | 16257.8 | 3006.2 | 888.2 | 245 | 469.4 | 13.54 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S03.csv | 29-05-26/glass slide empty/NAM-10S_70P_5AC_S3.csv | NAM | 18.38 | 19107.4 | 3033.2 | 963.6 | 337.8 | 594.74 | 12.85 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S04.csv | 29-05-26/glass slide empty/NAM-10S_70P_5AC_S4.csv | NAM | 21.07 | 18602.2 | 2971.4 | 909.4 | 278.4 | 530.86 | 12.24 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S05.csv | 29-05-26/glass slide empty/NAM-10S_70P_5AC_S5.csv | NAM | 17.24 | 16510 | 2958.4 | 829.6 | 314.4 | 545.31 | 12.59 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S06.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot1.csv | NAM | 26.62 | 16970.2 | 2088.6 | 687.6 | 230.4 | 360.95 | 10.17 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S07.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot2.csv | NAM | 21.83 | 18062.2 | 2075.2 | 664.4 | 226 | 348.49 | 11.62 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S08.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot3.csv | NAM | 21.96 | 17374.4 | 2071.4 | 653.2 | 217.6 | 346.52 | 11.59 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S09.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot4.csv | NAM | 20.92 | 16183.2 | 2030.4 | 634.2 | 211.2 | 333.1 | 11.91 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S10.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot5.csv | NAM | 21.3 | 15970.4 | 2047.8 | 651.2 | 202.4 | 333.29 | 11.52 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S11.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot6.csv | NAM | 20.2 | 15597.2 | 2073.2 | 735.2 | 214.6 | 335.25 | 12.32 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S12.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot7.csv | NAM | 21.69 | 15366.8 | 2056.8 | 702.2 | 224.2 | 330.72 | 11.52 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S13.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot8.csv | NAM | 21.56 | 14760.4 | 2065.6 | 687.8 | 211.4 | 329.34 | 11.74 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S14.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot9.csv | NAM | 21.85 | 14644.4 | 2063.6 | 699.2 | 223.6 | 329.57 | 11.35 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S15.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot10.csv | NAM | 19.94 | 14902.4 | 2052.2 | 641.4 | 213.8 | 334.36 | 12.49 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S16.csv | 29-05-26/NAM-70p-10s-5ac/NAM-samespot11.csv | NAM | 21.2 | 14691.8 | 2050.2 | 636.4 | 209 | 329.22 | 11.54 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S17.txtr | 29-05-26/NAM-70p-10s-5ac/NAM-samespot12.txtr | NAM | 22.17 | 14505 | 2057.2 | 653.8 | 211.2 | 331.74 | 11.46 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S01.csv | 29-05-26/glass slide empty/NAM-10S_80P_3AC_S1.csv | NAM | 16.55 | 14999 | 2349 | 901.67 | 288 | 522 | 15.59 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S02.csv | 29-05-26/glass slide empty/NAM-10S_80P_3AC_S2.csv | NAM | 16.3 | 17723.6667 | 2377.3333 | 910.67 | 265.33 | 507.76 | 15.56 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S03.csv | 29-05-26/glass slide empty/NAM-10S_80P_3AC_S3.csv | NAM | 18.93 | 20839.3333 | 2480.3333 | 1132.67 | 446.33 | 681.22 | 15.72 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S04.csv | 29-05-26/glass slide empty/NAM-10S_80P_3AC_S4.csv | NAM | 18.49 | 17987.3333 | 2559.6667 | 1090.67 | 328 | 611.98 | 15.41 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S05.csv | 29-05-26/glass slide empty/NAM-10S_80P_3AC_S5.csv | NAM | 20.34 | 16874 | 3160.3333 | 1866.67 | 518 | 1013.83 | 18.11 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S06.csv | 29-05-26/glass slide empty/NAM-10S_80P_3AC_S6.csv | NAM | 21.43 | 18849.3333 | 3004.6667 | 1593.67 | 625.67 | 1090.5 | 16.16 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S07.csv | 29-05-26/glass slide empty/NAM-10S_80P_3AC_S7.csv | NAM | 16.72 | 17151 | 2461.6667 | 984.67 | 370 | 633.26 | 15 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S01.txtr | 29-05-26/glass slide empty/NAM-10S_80P_5AC_S1.txtr | NAM | 21.09 | 17154.8 | 2387.6 | 936.8 | 211.8 | 441.08 | 12.56 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S02.csv | 29-05-26/glass slide empty/NAM-10S_80P_5AC_S2.csv | NAM | 16.73 | 17562 | 2630.8 | 1102.2 | 470.8 | 790.51 | 11.58 | No | No | 0 | 0.95 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S03.csv | 29-05-26/glass slide empty/NAM-10S_80P_5AC_S4.csv | NAM | 18.98 | 16116 | 2364.4 | 896.6 | 334 | 562.73 | 11.87 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S04.csv | 29-05-26/glass slide empty/NAM-10S_80P_5AC_S5.csv | NAM | 19.56 | 17593.6 | 2501.2 | 961.8 | 360.4 | 660.5 | 11.57 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S05.csv | 29-05-26/glass slide empty/NAM-10S_80P_5AC_S6.csv | NAM | 23.1 | 17343.2 | 2454.4 | 971 | 305.2 | 577.5 | 11.34 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S01.csv | 29-05-26/glass slide empty/NAM-10S_90P_3AC_S1.csv | NAM | 22.93 | 18741.3333 | 2477 | 1126.33 | 315.67 | 597.83 | 16.24 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S02.csv | 29-05-26/glass slide empty/NAM-10S_90P_3AC_S2.csv | NAM | 19.94 | 21654.6667 | 2479.6667 | 1273 | 388.67 | 695.42 | 15.72 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S03.csv | 29-05-26/glass slide empty/NAM-10S_90P_3AC_S3.csv | NAM | 20.55 | 19108.3333 | 2379.6667 | 996.33 | 136.67 | 513.57 | 16.46 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S04.csv | 29-05-26/glass slide empty/NAM-10S_90P_3AC_S4.csv | NAM | 22.61 | 18035.6667 | 2552.6667 | 1191.33 | 237.67 | 607.45 | 15.8 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S05.csv | 29-05-26/glass slide empty/NAM-10S_90P_3AC_S5.csv | NAM | 22.68 | 19578.3333 | 2361.6667 | 1025.67 | 355.33 | 605.03 | 14.67 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S06.csv | 29-05-26/glass slide empty/NAM-10S_90P_3AC_S6.csv | NAM | 17.1 | 19075 | 2329.6667 | 874 | 261 | 492.2 | 14.29 | No | No | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S07.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot1.csv | NAM | 20.54 | 21214 | 2299.6667 | 845.33 | 301.33 | 465.87 | 15.4 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S08.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot2.csv | NAM | 20.24 | 19283.3333 | 2311.6667 | 857.67 | 282.67 | 458.15 | 15.23 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S09.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot3.csv | NAM | 20.94 | 19663 | 2293.6667 | 829 | 290.67 | 450.38 | 14.78 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S10.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot4.csv | NAM | 20.93 | 21738.3333 | 2306.3333 | 864 | 285.67 | 466.16 | 15.17 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S11.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot5.csv | NAM | 19.89 | 22520 | 2332.6667 | 854 | 285.67 | 462.99 | 15.99 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S12.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot6.csv | NAM | 21.93 | 22120 | 2305.6667 | 857.33 | 279 | 453.05 | 14.59 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S13.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot7.csv | NAM | 20.34 | 20720 | 2325 | 828.67 | 285.33 | 443.47 | 15.25 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S14.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot8.csv | NAM | 19.55 | 19680.3333 | 2290 | 839 | 286.67 | 431.19 | 15.6 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S15.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot9.csv | NAM | 20.16 | 20712.6667 | 2291.3333 | 807 | 278.33 | 430.58 | 15.13 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S16.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot10.csv | NAM | 18.66 | 19426 | 2282.3333 | 802 | 272 | 420.98 | 15.84 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S17.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot11.csv | NAM | 19.8 | 18922.3333 | 2281 | 793 | 274 | 415.82 | 14.91 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S18.csv | 29-05-26/NAM-90p-10s-3ac/NAM-samespot12.csv | NAM | 19.3 | 20246.6667 | 2268 | 800 | 279.67 | 423.93 | 15.28 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S01.csv | 29-05-26/glass slide empty/NAM-10S_90P_5AC_S1.csv | NAM | 23.72 | 19423.8 | 2299.8 | 866.6 | 270.4 | 486.49 | 11.35 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S02.csv | 29-05-26/glass slide empty/NAM-10S_90P_5AC_S2.csv | NAM | 26.14 | 19562.6 | 2390.2 | 1034 | 271.2 | 521.88 | 12.81 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S03.csv | 29-05-26/glass slide empty/NAM-10S_90P_5AC_S3.csv | NAM | 31.8 | 18134.8 | 2379.6 | 1146.2 | 320.6 | 616.44 | 11.44 | No | No | 1 | 1 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S04.csv | 29-05-26/glass slide empty/NAM-10S_90P_5AC_S4.csv | NAM | 26.13 | 18456.8 | 2219.2 | 935.4 | 244.4 | 499.24 | 12.21 | No | No | 2 | 1 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S05.csv | 29-05-26/glass slide empty/NAM-10S_90P_5AC_S5.csv | NAM | 27.03 | 20913 | 2301.2 | 1002.8 | 313.8 | 589.74 | 11.36 | No | No | 2 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S06.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot1.csv | NAM | 27.6 | 22019.4 | 2402 | 953.8 | 331.6 | 556.67 | 11.46 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S07.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot2.csv | NAM | 24.75 | 20380.8 | 2372.4 | 906.4 | 304.6 | 518.27 | 12.36 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S08.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot3.csv | NAM | 24.36 | 22224 | 2379.8 | 902 | 307.6 | 515.42 | 13.01 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S09.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot4.csv | NAM | 25.43 | 22095.6 | 2344.6 | 889.6 | 319.8 | 503.77 | 12.25 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S10.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot5.csv | NAM | 23 | 20870.2 | 2329.8 | 863.6 | 300.6 | 487.51 | 13.49 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S11.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot6.csv | NAM | 21.71 | 20235 | 2323 | 849.2 | 298.4 | 477.77 | 14.03 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S12.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot7.csv | NAM | 25.29 | 20155.4 | 2333.2 | 851.4 | 289 | 462.4 | 12.1 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S13.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot8.csv | NAM | 25.24 | 19840.4 | 2324.4 | 833 | 293.8 | 455.58 | 11.88 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S14.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot9.csv | NAM | 25.95 | 19930.6 | 2332.8 | 829.2 | 286.8 | 454.23 | 11.57 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S15.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot10.csv | NAM | 23.24 | 19715.8 | 2328 | 853.6 | 293.2 | 450.26 | 12.94 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S16.csv | 29-05-26/NAM-90p-10s-5ac/NAM-samespot11.csv | NAM | 24.11 | 19195 | 2302.4 | 824.4 | 279.2 | 443.64 | 12.12 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S01.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot1.csv | NAM | 18.98 | 13882.8 | 2315.2 | 793 | 98.6 | 444.22 | 11.85 | No | No | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S02.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot2.csv | NAM | 21.7 | 13323.6 | 2328.4 | 840.4 | -40.8 | 420.43 | 12.66 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S03.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot3.csv | NAM | 23.19 | 15167 | 2342.6 | 857.4 | -18.8 | 470.64 | 11.49 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S04.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot4.csv | NAM | 18.6 | 14680.4 | 2265.2 | 736 | -100 | 409.58 | 11.25 | No | No | 1 | 0.97 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S05.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot5.csv | NAM | 23.56 | 17422.4 | 2386.2 | 913.4 | 17.2 | 490.55 | 12.07 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S06.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot6.csv | NAM | 21.4 | 17829 | 2643.6 | 1177.6 | 177.8 | 583.31 | 12.12 | No | No | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S07.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot7.csv | NAM | 28.53 | 18385.2 | 2754.2 | 1366.6 | 259.8 | 647.36 | 12.74 | No | No | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S08.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot8.csv | NAM | 22.51 | 16513 | 2728.4 | 897.4 | 54.4 | 405.4 | 11.85 | No | No | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S09.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot9.csv | NAM | 16.98 | 14855.8 | 2856.8 | 1062 | 113.2 | 544.47 | 12.23 | No | No | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S10.csv | 29-05-26/70p-5ac-10s-diifrentpoint/spot10.csv | NAM | 23.76 | 15195.4 | 2718.2 | 889.8 | 35.4 | 395.21 | 11.67 | No | No | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260530_P070pct_T30s_A05_S01.csv | 29-05-26/70p-5ac-30s-diifrentspot/spot1.csv | NAM | 54 | 45379.2 | 4619.2 | 2322.2 | 511.4 | 1031.36 | 15.61 | No | No | 0 | 1 | Valid |  | No
NAM_Bulk_20260530_P070pct_T30s_A05_S02.csv | 29-05-26/70p-5ac-30s-diifrentspot/spot2.csv | NAM | 40.31 | 49024.2 | 5854 | 3527.4 | 877 | 1764.45 | 19.09 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T30s_A05_S03.csv | 29-05-26/70p-5ac-30s-diifrentspot/spot3.csv | NAM | 45.87 | 47222.6 | 5926.4 | 2846.4 | 598 | 1180.15 | 17.44 | No | No | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T60s_A05_S01.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot1.csv | NAM | 68.43 | 65535 | 8421.2 | 4413.8 | 1009 | 2007.35 | 20.63 | Yes | No | 0 | 1 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 8421 counts (< 65 535), no negative counts, SNR 35.5, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | No
NAM_Bulk_20260530_P070pct_T60s_A05_S02.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot2.csv | NAM | 56.53 | 65523.2 | 7731.6 | 3290.8 | 751 | 1487.28 | 21.93 | Yes | No | 0 | 1 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 7732 counts (< 65 535), no negative counts, SNR 35.1, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | No
NAM_Bulk_20260530_P070pct_T60s_A05_S03.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot3.csv | NAM | 58.72 | 65287.2 | 9146.6 | 4237.2 | 1414 | 2428.72 | 23.82 | Yes | No | 0 | 1 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 9147 counts (< 65 535), no negative counts, SNR 33.9, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | No
NAM_Bulk_20260530_P090pct_T30s_A05_S01.csv | 29-05-26/90p-30s-5acc-samespot/samespot1.csv | NAM | 52.12 | 45311 | 6083.8 | 2586.8 | 862 | 1440.23 | 18.16 | No | No | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S02.csv | 29-05-26/90p-30s-5acc-samespot/samespot2.csv | NAM | 53.14 | 46251.6 | 5982.8 | 2472.2 | 810 | 1341.15 | 17.74 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S03.csv | 29-05-26/90p-30s-5acc-samespot/samespot3.csv | NAM | 51.93 | 51634 | 5910.2 | 2371 | 802.8 | 1272.08 | 17.66 | No | No | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S04.csv | 29-05-26/90p-30s-5acc-samespot/samespot4.csv | NAM | 52.21 | 51343.6 | 5871.2 | 2348.2 | 764.4 | 1234.78 | 17.95 | No | No | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S05.csv | 29-05-26/90p-30s-5acc-samespot/samespot5.csv | NAM | 57.79 | 51005.2 | 5858.6 | 2359.2 | 788.2 | 1225.22 | 16.03 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S06.csv | 29-05-26/90p-30s-5acc-samespot/samespot6.csv | NAM | 48.54 | 45106.4 | 5886.2 | 2330.4 | 760.6 | 1214 | 18.67 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S07.csv | 29-05-26/90p-30s-5acc-samespot/samespot7.csv | NAM | 53.98 | 44502 | 5869.8 | 2359.4 | 756.4 | 1213.98 | 17.24 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S08.csv | 29-05-26/90p-30s-5acc-samespot/samespot8.csv | NAM | 54.27 | 49610.4 | 5839 | 2308.2 | 764.6 | 1179.25 | 17.16 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S09.csv | 29-05-26/90p-30s-5acc-samespot/samespot9.csv | NAM | 51.58 | 52755.8 | 5859 | 2327.2 | 749.2 | 1179.33 | 18.25 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S10.csv | 29-05-26/90p-30s-5acc-samespot/samespot10.csv | NAM | 53.52 | 45792.8 | 5865.4 | 2334 | 761.6 | 1179.93 | 17.76 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S11.csv | 29-05-26/90p-30s-5ac-diifrentpoint/spot1.csv | NAM | 56.57 | 43778.4 | 5867.4 | 2317.8 | 746.6 | 1170.45 | 16.69 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S12.csv | 29-05-26/90p-30s-5ac-diifrentpoint/spot2.csv | NAM | 47.22 | 47144.2 | 6976.8 | 3637.2 | 1323.8 | 2336.33 | 17.76 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S13.csv | 29-05-26/90p-30s-5ac-diifrentpoint/spot3.csv | NAM | 48.54 | 51106 | 6090.6 | 2640.6 | 634.2 | 1267.16 | 17.19 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S14.csv | 29-05-26/90p-30s-5ac-diifrentpoint/spot4.csv | NAM | 45.72 | 43173.8 | 6369.8 | 3058.6 | 704.2 | 1498.24 | 19.36 | No | No | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S15.csv | 29-05-26/90p-30s-5ac-diifrentpoint/spot5.csv | NAM | 39.25 | 45705.8 | 6498 | 3095.6 | 705.8 | 1611.07 | 23.47 | No | No | 0 | 0.98 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T60s_A05_S01.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot1.csv | NAM | 16.73 | 65535 | 18624.8 | 9798.4 | 1306.8 | 5886.89 | 206.87 | Yes | No | 0 | 0.45 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 18625). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 310 counts (reference ≈ 26), r = 0.46. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S02.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot2.csv | NAM | 18.4 | 65535 | 22234.8 | 14041 | 1563 | 8819.36 | 338.34 | Yes | No | 0 | 0.25 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 22235). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 554 counts (reference ≈ 26), r = 0.26. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S03.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot3.csv | NAM | 21.15 | 65535 | 27829.4 | 17015.2 | 2957.8 | 11751.21 | 352.65 | Yes | No | 2 | 0.15 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 27829). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 622 counts (reference ≈ 26), r = 0.17. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S04.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot4.csv | NAM | 21.56 | 65535 | 28506.4 | 18123.6 | 2225 | 11905.51 | 386.89 | Yes | No | 2 | 0.11 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 28506). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 694 counts (reference ≈ 26), r = 0.12. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S05.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot5.csv | NAM | 20.67 | 65535 | 36617.4 | 27315 | 2696 | 17770.98 | 650.1 | Yes | No | 2 | 0.07 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 36617). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 1075 counts (reference ≈ 26), r = 0.09. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S06.csv | 29-05-26/90p-60s-5a-samespot/samespot1.csv | NAM | 16.02 | 65535 | 34426.4 | 2642 | -2877.2 | -2039.42 | 171.87 | Yes | No | 0 | 0.4 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 88 % of 400–1800 cm-1 (min -2877); r = 0.36. Unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S07.csv | 29-05-26/90p-60s-5a-samespot/samespot2.csv | NAM | 14.94 | 65535 | 30949.6 | 1995.4 | -6048.6 | -4919.33 | 266.96 | Yes | No | 0 | 0.24 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 96 % of 400–1800 cm-1 (min -6049); r = 0.19. Unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S08.csv | 29-05-26/90p-60s-5a-samespot/samespot3.csv | NAM | 15.13 | 65535 | 28975.6 | 1736.4 | -8447.6 | -6061.68 | 295.15 | Yes | No | 0 | 0.2 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 97 % of 400–1800 cm-1 (min -8448); r = 0.17. Unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S09.csv | 29-05-26/90p-60s-5a-samespot/samespot4.csv | NAM | 14.27 | 65535 | 27166.6 | 1286.4 | -9894.6 | -7648.97 | 349.46 | Yes | No | 0 | 0.17 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 98 % of 400–1800 cm-1 (min -9895); r = 0.13. Unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S10.csv | 29-05-26/90p-60s-5a-samespot/samespot5.csv | NAM | 14.29 | 65535 | 25199.6 | 973.8 | -11649.4 | -9432.28 | 396.93 | Yes | No | 0 | 0.13 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 99 % of 400–1800 cm-1 (min -11649); r = 0.09. Unreliable. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S01.csv | 9-2-2026/20sec20power-variable location LTA/SP_2.csv | SUSPECT | 50.59 | 8546.5 | 4954 | 3694 | 699.5 | 1774.58 | 24.5 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S02.csv | 9-2-2026/20sec20power-variable location LTA/SP_3.csv | SUSPECT | 87.92 | 10289.5 | 5852 | 4573 | 710.5 | 1664.04 | 23.46 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S03.csv | 9-2-2026/20sec20power-variable location LTA/SP_4.csv | SUSPECT | 81.45 | 9935.5 | 5664.5 | 4385.5 | 677.5 | 1534.38 | 25.06 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S04.csv | 9-2-2026/20sec20power-variable location LTA/SP_5.csv | SUSPECT | 69.88 | 9612.5 | 5468.5 | 4189.5 | 747 | 1827.9 | 22.55 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S05.csv | 9-2-2026/20sec20power-variable location LTA/SP_6.csv | SUSPECT | 42.95 | 7964 | 4989.5 | 3594 | 688.5 | 1753.5 | 23.93 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S06.csv | 9-2-2026/20sec20power-variable location LTA/SP_7.csv | SUSPECT | 58.35 | 9437.5 | 5353 | 4102.5 | 746.5 | 1874.21 | 24.74 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S07.csv | 9-2-2026/20sec20power-variable location LTA/SP_8.csv | SUSPECT | 26.45 | 6216 | 3791 | 2392.5 | 442 | 1120.88 | 20.95 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S08.csv | 9-2-2026/20sec20power-variable location LTA/SP_10.csv | SUSPECT | 69.41 | 9514 | 5367.5 | 4088.5 | 676.5 | 1667.85 | 23.58 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S09.csv | 9-2-2026/20sec20power-variable location LTA/SP_11.csv | SUSPECT | 83.93 | 11871 | 6771 | 5492 | 935.5 | 2252.96 | 26.07 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S10.csv | 9-2-2026/20sec20power-variable location LTA/SP_12.csv | SUSPECT | 62.39 | 9434.5 | 5357 | 4078 | 739 | 1859.41 | 23.68 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S11.csv | 9-2-2026/20sec20power-variable location LTA/SP_13.csv | SUSPECT | 41.56 | 8014 | 5199.5 | 3844.5 | 697 | 1809.5 | 23.7 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S12.csv | 9-2-2026/20sec20power-variable location LTA/SP_14.csv | SUSPECT | 96.6 | 14278.5 | 7865.5 | 6586.5 | 1015.5 | 2511.85 | 28.82 | No | No | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No

## 3_Reference_Conditions
Sheet 3 — Acquisition conditions and reference selection
Reference: NAM 90 % / 30 s / 5 acc (n = 15); LTA 80–95 % / 45 s / 2 acc (n = 4). LPS: none (pending glass-blank-controlled screening).
NAM 70 % / 60 s / 5 acc (n = 3): laser-line saturation only; fingerprint valid (class A). Retained as a SUPPORTING group, not merged into the reference. NAM 90 % / 60 s / 5 acc (n = 10): distorted (class C), excluded.
Biomarker | Power (%) | Integration (s) | Accumulations | n | n Valid | n saturated | Mean SNR | SNR SD | Mean strongest-marker intensity (counts) | Intensity CV (%) | Mean baseline (counts) | Mean max raw counts (400–1800) | Mean pairwise r | Dates | Reference condition
LPS | 20 | 30 | 60 | 1 | 0 | 0 | 76.75 |  |  |  | 877.65 | 4909.35 |  | 20260413
LPS | 20 | 30 | 50 | 1 | 0 | 0 | 72.82 |  |  |  | 829.3 | 4193.54 |  | 20260413
LPS | 20 | 60 | 30 | 2 | 0 | 0 | 30.31 | 1.69 |  |  | 1242.95 | 12597.83 | 0.99 | 20260413
LPS | 20 | 60 | 50 | 2 | 0 | 0 | 20.27 | 0.65 |  |  | 1701.08 | 10443.32 | -0.19 | 20260413
LPS | 10 | 60 | 30 | 1 | 0 | 0 | 14.75 |  |  |  | 400.17 | 11756.5 |  | 20260413
LPS | 10 | 10 | 5 | 2 | 0 | 0 | 6.83 | 3.52 |  |  | 323.82 | 3321 | 0.83 | 20260413
LPS | 10 | 10 | 2 | 3 | 0 | 0 | 6.14 | 2.76 |  |  | 462.88 | 2373 | 0.56 | 20260413
LPS | 20 | 10 | 5 | 1 | 0 | 0 | 5.73 |  |  |  | 439.02 | 3446 |  | 20260413
LPS | 20 | 10 | 2 | 1 | 0 | 0 | 4.66 |  |  |  | 393.72 | 3393 |  | 20260413
LTA | 95 | 45 | 2 | 1 | 1 | 0 | 29.34 |  | 1196.58 |  | 7536.68 | 14848 |  | 20260403 | YES
LTA | 80 | 45 | 2 | 2 | 2 | 0 | 28.69 | 1.01 | 1057.98 | 11.76 | 8225.4 | 17203 | 0.93 | 20260403 | YES
LTA | 90 | 45 | 2 | 1 | 1 | 0 | 27.38 |  | 1127.7 |  | 8013.24 | 16295.5 |  | 20260403 | YES
LTA | 60 | 45 | 2 | 1 | 1 | 0 | 23.38 |  | 735.05 |  | 3591.98 | 10105 |  | 20260403
LTA | 70 | 10 | 10 | 1 | 1 | 0 | 21.58 |  | 182.17 |  | 679.38 | 3150 |  | 20260403
LTA | 60 | 10 | 10 | 1 | 1 | 0 | 17.27 |  | 154.05 |  | 589.91 | 3054.9 |  | 20260403
LTA | 50 | 10 | 10 | 1 | 1 | 0 | 15.5 |  | 134.45 |  | 555.52 | 2999.9 |  | 20260403
LTA | 50 | 30 | 15 | 1 | 0 | 0 | 13.7 |  | 754.35 |  | 2858.17 | 9292 |  | 20260403
LTA | 40 | 10 | 10 | 1 | 1 | 0 | 11.28 |  | 99.31 |  | 376.02 | 2797.1 |  | 20260403
LTA | 50 | 5 | 10 | 1 | 0 | 0 | 8.81 |  | 69.88 |  | 217.02 | 1969.9 |  | 20260403
LTA | 30 | 5 | 10 | 1 | 0 | 0 | 6.57 |  | 53.78 |  | 162.3 | 1902.1 |  | 20260403
NAM | 70 | 60 | 5 | 3 | 0 | 3 | 61.23 | 6.33 | 1350.19 | 7.09 | 1974.45 | 8433.13 |  | 20260530
NAM | 90 | 30 | 5 | 15 | 15 | 0 | 51.09 | 4.66 | 916.84 | 4.02 | 1357.55 | 6055.24 | 0.98 | 20260530 | YES
NAM | 90 | 30 | 2 | 3 | 3 | 0 | 50.22 | 4.61 | 1428.71 | 15.05 | 2086.21 | 6001.33 | 0.99 | 20260403
NAM | 95 | 30 | 2 | 5 | 5 | 0 | 46.74 | 4.01 | 1332.06 | 13.1 | 2011.71 | 5878.1 | 1 | 20260403
NAM | 70 | 30 | 5 | 3 | 3 | 0 | 46.73 | 6.88 | 804.21 | 4.6 | 1325.32 | 5466.53 | 1 | 20260530
NAM | 100 | 30 | 2 | 5 | 5 | 0 | 46.16 | 11.76 | 1264.11 | 34.84 | 1682.34 | 5444.2 | 0.99 | 20260403
NAM | 80 | 40 | 2 | 8 | 8 | 0 | 38.35 | 9.67 | 1160.58 | 36.1 | 2387.25 | 7722.12 | 0.99 | 20260403
NAM | 70 | 30 | 2 | 1 | 1 | 0 | 37.13 |  | 1205.58 |  | 2868.94 | 7378.5 |  | 20260403
NAM | 100 | 35 | 2 | 2 | 2 | 0 | 36.62 | 1.55 | 925.44 | 0.26 | 1344.89 | 6475.75 | 1 | 20260403
NAM | 85 | 30 | 2 | 6 | 6 | 0 | 34.11 | 9.06 | 991.01 | 35.09 | 1861.75 | 5726.42 | 0.99 | 20260403
NAM | 80 | 25 | 2 | 1 | 1 | 0 | 32.47 |  | 710.73 |  | 810.89 | 3807.5 |  | 20260128
NAM | 90 | 35 | 2 | 1 | 1 | 0 | 31.93 |  | 871.02 |  | 1320.7 | 6471 |  | 20260403
NAM | 75 | 25 | 2 | 1 | 1 | 0 | 31.76 |  | 684.57 |  | 796.45 | 3820 |  | 20260128
NAM | 70 | 25 | 2 | 1 | 1 | 0 | 30.99 |  | 647.79 |  | 739.65 | 3758 |  | 20260128
NAM | 80 | 35 | 2 | 7 | 7 | 0 | 30.82 | 4.06 | 837.84 | 16.93 | 1970.05 | 6040.93 | 0.93 | 20260403
NAM | 80 | 20 | 2 | 1 | 1 | 0 | 29.21 |  | 573.48 |  | 710.29 | 3313.5 |  | 20260128
NAM | 75 | 20 | 2 | 1 | 1 | 0 | 28.4 |  | 547.42 |  | 687.01 | 3307 |  | 20260128
NAM | 65 | 25 | 2 | 1 | 1 | 0 | 28.3 |  | 589.56 |  | 672.86 | 3659.5 |  | 20260128
NAM | 75 | 30 | 2 | 11 | 11 | 0 | 28.13 | 5.46 | 769.26 | 23.17 | 2166.6 | 6382.77 | 0.95 | 20260403
NAM | 80 | 30 | 2 | 6 | 6 | 0 | 26.7 | 6.27 | 666.12 | 27.77 | 1785.82 | 5392.92 | 0.92 | 20260403
NAM | 60 | 25 | 2 | 1 | 1 | 0 | 26.54 |  | 552.55 |  | 646.9 | 3668.5 |  | 20260128
NAM | 55 | 25 | 2 | 1 | 1 | 0 | 26.29 |  | 541.23 |  | 622.64 | 3628 |  | 20260128
NAM | 70 | 20 | 2 | 1 | 1 | 0 | 26.26 |  | 495.58 |  | 636.42 | 3197.5 |  | 20260128
NAM | 90 | 10 | 5 | 16 | 16 | 0 | 25.34 | 2.31 | 309.9 | 6.46 | 502.46 | 2335.15 | 0.99 | 20260529
NAM | 60 | 20 | 2 | 1 | 1 | 0 | 23.37 |  | 415.66 |  | 534.18 | 3441.5 |  | 20260128
NAM | 80 | 15 | 2 | 1 | 1 | 0 | 23.28 |  | 388.99 |  | 516.03 | 2760.5 |  | 20260128
NAM | 65 | 20 | 2 | 1 | 1 | 0 | 22.56 |  | 441.78 |  | 567.5 | 3136 |  | 20260128
NAM | 45 | 25 | 2 | 1 | 1 | 0 | 22.48 |  | 440.32 |  | 509.35 | 3484.5 |  | 20260128
NAM | 50 | 25 | 2 | 1 | 1 | 0 | 22.48 |  | 444.84 |  | 535 | 3509 |  | 20260128
NAM | 70 | 10 | 5 | 27 | 27 | 0 | 21.39 | 2.61 | 254.67 | 11.49 | 429.83 | 2413.8 | 0.97 | 20260529, 20260530
NAM | 75 | 15 | 2 | 2 | 2 | 0 | 21.28 | 1.16 | 386.14 | 2.6 | 584.21 | 3072.75 | 0.99 | 20260128
NAM | 50 | 35 | 2 | 1 | 1 | 0 | 21.27 |  | 491.46 |  | 836.61 | 4740 |  | 20260403
NAM | 70 | 15 | 2 | 2 | 2 | 0 | 21.01 | 0.76 | 361.28 | 0.25 | 558.28 | 3027.25 | 0.99 | 20260128
NAM | 65 | 15 | 2 | 2 | 2 | 0 | 20.73 | 0.52 | 334.21 | 2.38 | 512.02 | 3227.75 | 0.97 | 20260128
NAM | 40 | 25 | 2 | 1 | 1 | 0 | 20.55 |  | 381.19 |  | 475.07 | 3421 |  | 20260128
NAM | 90 | 10 | 3 | 18 | 18 | 0 | 20.45 | 1.46 | 314.11 | 8.6 | 490.78 | 2342.65 | 0.98 | 20260529
NAM | 50 | 20 | 2 | 1 | 1 | 0 | 20.34 |  | 353.01 |  | 441.83 | 3545.5 |  | 20260128
NAM | 55 | 20 | 2 | 1 | 1 | 0 | 20.23 |  | 385.61 |  | 502.07 | 3618 |  | 20260128
NAM | 45 | 20 | 2 | 1 | 1 | 0 | 19.91 |  | 319.01 |  | 439.27 | 3554 |  | 20260128
NAM | 80 | 10 | 5 | 5 | 5 | 0 | 19.89 | 2.38 | 234.41 | 12.6 | 606.46 | 2467.68 | 0.92 | 20260529
NAM | 60 | 15 | 2 | 2 | 2 | 0 | 19.18 | 0.73 | 327.79 | 3.82 | 550.29 | 3289.5 | 0.98 | 20260128
NAM | 80 | 10 | 3 | 7 | 7 | 0 | 18.4 | 1.99 | 294.23 | 15.9 | 722.94 | 2627.57 | 0.96 | 20260529
NAM | 55 | 15 | 2 | 2 | 2 | 0 | 18.33 | 0.44 | 311.46 | 0.08 | 478.68 | 3196 | 0.98 | 20260128
NAM | 35 | 25 | 2 | 1 | 1 | 0 | 18.27 |  | 343.91 |  | 412.46 | 3374 |  | 20260128
NAM | 80 | 10 | 2 | 1 | 1 | 0 | 18.12 |  | 273.55 |  | 392.66 | 2430 |  | 20260128
NAM | 30 | 25 | 2 | 1 | 1 | 0 | 17.59 |  | 316.82 |  | 369.99 | 3352 |  | 20260128
NAM | 90 | 60 | 5 | 10 | 0 | 10 | 17.32 | 2.9 | 6078.95 | 51.26 | 2603.23 | 28053.06 |  | 20260530
NAM | 70 | 10 | 3 | 16 | 16 | 0 | 16.54 | 2.18 | 245.8 | 9.74 | 400.76 | 2345.94 | 0.95 | 20260529
NAM | 75 | 10 | 2 | 1 | 1 | 0 | 16.46 |  | 266.23 |  | 377.79 | 2416.5 |  | 20260128
NAM | 70 | 10 | 2 | 1 | 1 | 0 | 16.27 |  | 244.8 |  | 346.3 | 2398 |  | 20260128
NAM | 25 | 25 | 2 | 1 | 1 | 0 | 15.77 |  | 288.73 |  | 309 | 3220.5 |  | 20260128
NAM | 40 | 20 | 2 | 1 | 1 | 0 | 15.17 |  | 296.78 |  | 397.59 | 3500 |  | 20260128
NAM | 45 | 15 | 2 | 2 | 2 | 0 | 14.98 | 2.72 | 259.87 | 2.41 | 333.22 | 3189 | 0.9 | 20260128
NAM | 20 | 25 | 2 | 1 | 1 | 0 | 14.79 |  | 233.23 |  | 275.39 | 3241 |  | 20260128
NAM | 50 | 15 | 2 | 2 | 1 | 0 | 14.6 | 2.14 | 269.27 | 5.44 | 318.53 | 3198.25 | 0.84 | 20260128
NAM | 65 | 10 | 2 | 1 | 1 | 0 | 14.21 |  | 225.8 |  | 324.53 | 2360.5 |  | 20260128
NAM | 35 | 20 | 2 | 1 | 1 | 0 | 13.7 |  | 252.72 |  | 340.73 | 3448 |  | 20260128
NAM | 60 | 10 | 2 | 1 | 1 | 0 | 13.58 |  | 212.22 |  | 308.33 | 2349.5 |  | 20260128
NAM | 40 | 15 | 2 | 3 | 3 | 0 | 13.32 | 1.27 | 241.25 | 4.13 | 326.27 | 3291.5 | 0.92 | 20260128
NAM | 30 | 20 | 2 | 1 | 1 | 0 | 13.08 |  | 232.53 |  | 304.1 | 3431.5 |  | 20260128
NAM | 35 | 15 | 2 | 2 | 2 | 0 | 12.94 | 2.61 | 209.5 | 1.4 | 304.52 | 3154 | 0.9 | 20260128
NAM | 55 | 10 | 2 | 1 | 1 | 0 | 12.37 |  | 194.45 |  | 285.41 | 2309.5 |  | 20260128
NAM | 30 | 15 | 2 | 3 | 3 | 0 | 12.08 | 0.69 | 199.33 | 5.76 | 372.47 | 3336.5 | 0.94 | 20260128
NAM | 25 | 20 | 2 | 1 | 1 | 0 | 11.56 |  | 200.45 |  | 255.4 | 3345 |  | 20260128
NAM | 45 | 10 | 2 | 1 | 1 | 0 | 11.55 |  | 171.64 |  | 243.68 | 1997.5 |  | 20260128
NAM | 20 | 20 | 2 | 1 | 1 | 0 | 11.29 |  | 173.71 |  | 227.8 | 3331.5 |  | 20260128
NAM | 15 | 25 | 2 | 1 | 1 | 0 | 11.15 |  | 196.62 |  | 228.81 | 3165 |  | 20260128
NAM | 50 | 10 | 2 | 1 | 1 | 0 | 10.66 |  | 179.92 |  | 260.02 | 2017 |  | 20260128
NAM | 80 | 5 | 2 | 1 | 1 | 0 | 10.56 |  | 133.94 |  | 219.11 | 1572.5 |  | 20260128
NAM | 25 | 15 | 2 | 2 | 1 | 0 | 10.3 | 1.11 | 168.88 | 0.32 | 329.07 | 3193 | 0.92 | 20260128
NAM | 35 | 10 | 2 | 1 | 1 | 0 | 10.04 |  | 152.82 |  | 192.87 | 1947 |  | 20260128
NAM | 40 | 10 | 2 | 1 | 1 | 0 | 10.04 |  | 146.76 |  | 217.21 | 1976.5 |  | 20260128
NAM | 20 | 15 | 2 | 2 | 0 | 0 | 9.81 | 0.21 | 157.37 | 11.76 | 283.37 | 3171.75 | 0.9 | 20260128
NAM | 75 | 5 | 2 | 1 | 0 | 0 | 9.53 |  | 133.08 |  | 200.98 | 1555.5 |  | 20260128
NAM | 15 | 20 | 2 | 1 | 0 | 0 | 9.26 |  | 140.78 |  | 190.05 | 3300.5 |  | 20260128
NAM | 25 | 10 | 2 | 1 | 0 | 0 | 8.8 |  | 123.9 |  | 157.25 | 1913.5 |  | 20260128
NAM | 65 | 5 | 2 | 1 | 0 | 0 | 8.61 |  | 121.94 |  | 172.98 | 1509.5 |  | 20260128
NAM | 30 | 10 | 2 | 1 | 0 | 0 | 8.4 |  | 125.04 |  | 182.16 | 1934 |  | 20260128
NAM | 70 | 5 | 2 | 1 | 0 | 0 | 8.3 |  | 118.78 |  | 184.98 | 1529.5 |  | 20260128
NAM | 60 | 5 | 2 | 1 | 0 | 0 | 8.06 |  | 109.73 |  | 167.74 | 1506 |  | 20260128
NAM | 10 | 25 | 2 | 1 | 0 | 0 | 7.76 |  | 134.43 |  | 162.92 | 3098.5 |  | 20260128
NAM | 15 | 15 | 2 | 1 | 0 | 0 | 7.74 |  | 114.72 |  | 154.75 | 2713.5 |  | 20260128
NAM | 55 | 5 | 2 | 1 | 0 | 0 | 7.43 |  | 104.76 |  | 151.68 | 1475.5 |  | 20260128
NAM | 50 | 5 | 2 | 1 | 0 | 0 | 7.03 |  | 99.56 |  | 144.69 | 1508.5 |  | 20260128
NAM | 5 | 25 | 2 | 1 | 0 | 0 | 6.82 |  | 113.32 |  | 116.9 | 3058 |  | 20260128
NAM | 20 | 10 | 2 | 2 | 0 | 0 | 6.41 | 0.11 | 92.9 | 3.1 | 131.99 | 1876.25 | 0.91 | 20260128
NAM | 40 | 5 | 2 | 1 | 0 | 0 | 6.34 |  | 84.33 |  | 121.11 | 1454.5 |  | 20260128
NAM | 35 | 5 | 2 | 1 | 0 | 0 | 6.32 |  | 77.21 |  | 110.04 | 1430 |  | 20260128
NAM | 45 | 5 | 2 | 1 | 0 | 0 | 6.3 |  | 86.56 |  | 132.58 | 1484.5 |  | 20260128
NAM | 10 | 15 | 2 | 3 | 0 | 0 | 6.29 | 0.4 | 102.41 | 6.48 | 189.82 | 3167.17 | 0.75 | 20260128
NAM | 10 | 20 | 2 | 1 | 0 | 0 | 6.25 |  | 99.34 |  | 127.14 | 3173.5 |  | 20260128
NAM | 15 | 10 | 2 | 1 | 0 | 0 | 5.83 |  | 80.87 |  | 107.23 | 1833.5 |  | 20260128
NAM | 30 | 5 | 2 | 1 | 0 | 0 | 5.24 |  | 70.72 |  | 97.39 | 1441.5 |  | 20260128
NAM | 5 | 15 | 2 | 2 | 0 | 0 | 5.16 | 0.12 | 76.96 | 6.27 | 153.29 | 3033 | 0.67 | 20260128
NAM | 5 | 20 | 2 | 1 | 0 | 0 | 5 |  | 74.29 |  | 97.06 | 2628.5 |  | 20260128
NAM | 10 | 10 | 2 | 1 | 0 | 0 | 4.64 |  | 60.01 |  | 76.21 | 1799 |  | 20260128
NAM | 25 | 5 | 2 | 1 | 0 | 0 | 4.58 |  | 61.52 |  | 85.33 | 1408 |  | 20260128
NAM | 20 | 5 | 2 | 1 | 0 | 0 | 4.54 |  | 66.96 |  | 76.35 | 1383 |  | 20260128
NAM | 5 | 10 | 2 | 1 | 0 | 0 | 3.58 |  | 47.97 |  | 58.73 | 1782 |  | 20260128
NAM | 10 | 5 | 2 | 1 | 0 | 0 | 3.43 |  | 47.39 |  | 42.48 | 1387 |  | 20260128
NAM | 15 | 5 | 2 | 1 | 0 | 0 | 2.81 |  | 39.5 |  | 57.33 | 1397.5 |  | 20260128
NAM | 5 | 5 | 1 | 2 | 0 | 0 | 2.78 | 0.46 | 67.9 | 10.29 | 97.4 | 1883 | 0.15 | 20260529
NAM | 10 | 5 | 1 | 4 | 0 | 0 | 2.73 | 0.03 | 67.53 | 6.87 | 161.62 | 1908.25 | 0.58 | 20260529
NAM | 5 | 5 | 2 | 1 | 0 | 0 | 2.54 |  | 34.52 |  | 28.58 | 1363.5 |  | 20260128

## 4_Final_Marker_Table
Sheet 4 — FINAL marker table (bulk powder on glass slide; NAM reference n = 15; LTA preliminary reference n = 4)
Primary marker = reproducible in all reference replicates and across measurement days, SNR ≥ 20 (NAM) / ≥ 15 (LTA), outside or separable from the glass zone. Supporting = reproducible but weaker/less specific. Glass/background = glass-associated. Uncertain = observed, insufficient evidence.
Literature: No directly comparable pure-compound normal Raman reference identified in the searched literature. Assignments are proposed from general Raman group-frequency tables (L5–L7) and are NOT compound-specific proof.
Biomarker | Peak category | Experimental wavenumber (cm-1) | Mean ± SD (cm-1) | Mean intensity (counts, baseline-corr.) | Intensity SD | CV (%) | SNR | Prominence (counts) | Replicate frequency (detected / n) | Literature wavenumber / region (cm-1) | Difference (exp − lit) | Proposed assignment | Literature citation | Confidence | Background status | Height after exploratory glass-template correction
LTA | Primary marker | 802.2 | 802.2 ± 5.1 | 834 | 105 | 12.6 | 21.4 | 891 | 4/4 | ~813 (ν O-P-O) | ≈ −11 | Proposed: ν(O-P-O) phosphodiester stretch | L5 | Low | Outside glass zone | 783
LTA | Primary marker | 861.5 | 861.5 ± 1.0 | 687 | 74 | 10.8 | 17.6 | 408 | 4/4 | 800–950 (region) | within region | Proposed: ν(C-C)/ν(C-O-C) of glycerol / sugar substituents | L6 | Low | Outside glass zone | 638
LTA | Primary marker | 1091.8 | 1091.8 ± 0.5 | 1110 | 98 | 8.8 | 28.5 | 1180 | 4/4 | 1090–1095 (ν_s PO2−) | ≈ 0 to +2 | Proposed: ν_s(PO2−) of phosphodiester (poly-glycerophosphate backbone) | L5, L7, L8 | Moderate | Outside glass zone | 1051
LTA | Supporting peak | 1458.8 | 1458.8 ± 0.5 | 721 | 53 | 7.4 | 18.5 | 778 | 4/4 | 1440–1460 (δ CH2) | ≈ 0 to +9 | Proposed: δ(CH2) scissoring (acyl chains of glycolipid anchor); non-specific | L7 | Low | Adjacent to glass zone; ~18 % of height glass-related (exploratory) | 593
LTA | Uncertain | 541.2 | 541.2 ± 5.2 | 440 | 121 | 27.5 | 11.3 | 465 | 4/4 | 400–600 (region) | within region | Unassigned (skeletal region; CV 27 %) | — | Uncertain | Outside glass zone | 392
LTA | Glass/background | 1367.2 | 1367.2 ± 2.2 | 511 | 35 | 6.8 | 13.2 | 121 | 1/4 | — | — | Glass-associated feature (inside glass zone; removed by glass-template correction; 2/9 detected) | — | Not applicable | Glass-associated feature | -68
NAM | Primary marker | 541.0 | 541.0 ± 0.0 | 715 | 33 | 4.6 | 39.9 | 736 | 15/15 | 400–600 (region) | within region | Proposed: pyranose-ring skeletal deformation δ(C-C-O)/δ(C-C-C) | L6 | Low | Outside glass zone | 671
NAM | Primary marker | 829.0 | 829.0 ± 0.0 | 625 | 30 | 4.8 | 34.8 | 571 | 15/15 | 800–950 (region) | within region | Proposed: ν(C-O-C)/ring mode in the carbohydrate anomeric region | L6 | Low | Outside glass zone | 580
NAM | Primary marker | 870.5 | 870.5 ± 0.5 | 872 | 37 | 4.3 | 48.6 | 820 | 15/15 | 800–950 (region) | within region | Proposed: ν(C-O-C)/ν(C-C) ring mode (anomeric region) | L6 | Low | Outside glass zone | 826
NAM | Primary marker | 929.0 | 929.0 ± 0.0 | 917 | 37 | 4 | 51.1 | 945 | 15/15 | 800–950 (region) | within region | Proposed: ν(C-O-C)/ν(C-C); possible lactyl C-COOH contribution | L6, L5 | Low | Outside glass zone | 870
NAM | Primary marker | 955.0 | 955.0 ± 0.0 | 862 | 44 | 5.1 | 48.1 | 830 | 15/15 | 800–950 (upper edge) | ≈ +5 beyond region | Proposed: ν(C-C) / CH3 rocking (N-acetyl or lactyl methyl) – tentative | L5, L6 | Low | Outside glass zone | 822
NAM | Primary marker | 1022.7 | 1022.7 ± 0.5 | 749 | 47 | 6.2 | 41.8 | 783 | 15/15 | 1000–1200 (region) | within region | Proposed: ν(C-O)/ν(C-C) (C-OH stretching) of sugar ring | L6 | Low | Outside glass zone | 712
NAM | Primary marker | 1127.0 | 1127.0 ± 0.0 | 581 | 30 | 5.2 | 32.4 | 519 | 15/15 | 1000–1200 (region) | within region | Proposed: ν(C-O-C) (3-O-lactyl ether / ring) or ν(C-C) | L6 | Low | Outside glass zone | 545
NAM | Supporting peak | 457.0 | 457.0 ± 0.4 | 547 | 32 | 5.8 | 30.5 | 560 | 15/15 | 400–600 (region) | within region | Proposed: pyranose-ring skeletal deformation | L6 | Low | Outside glass zone | 517
NAM | Supporting peak | 487.0 | 487.0 ± 0.0 | 398 | 33 | 8.3 | 22.2 | 395 | 15/15 | 400–600 (region) | within region | Proposed: skeletal deformation | L6 | Low | Outside glass zone | 378
NAM | Supporting peak | 1248.7 | 1248.7 ± 0.5 | 290 | 11 | 3.7 | 16.2 | 318 | 15/15 | — | — | No assignment proposed (possible amide III / CH deformation region; not verified) | — | Uncertain | Outside glass zone | 245
NAM | Supporting peak | 1589.5 | 1589.5 ± 0.5 | 263 | 23 | 8.9 | 14.7 | 268 | 15/15 | — | — | No assignment proposed (unassigned; reproducible) | — | Uncertain | Outside glass zone | 198
NAM | Uncertain | 583.8 | 583.8 ± 0.6 | 268 | 44 | 16.3 | 15.1 | 241 | 14/15 | — | — | Unassigned | — | Uncertain | Outside glass zone | 230
NAM | Uncertain | 1058.9 | 1058.9 ± 0.3 | 285 | 18 | 6.4 | 15.8 | 150 | 15/15 | — | — | Unassigned (detected in only 31 % of valid spectra) | — | Uncertain | Outside glass zone | 254
NAM | Uncertain | 1155.0 | 1155.0 ± 0.0 | 366 | 21 | 5.8 | 20.4 | 266 | 15/15 | — | — | Unassigned | — | Uncertain | Outside glass zone | 318
NAM | Uncertain | 1344.9 | 1344.9 ± 1.3 | 626 | 107 | 17.1 | 35.2 | 609 | 15/15 | — | — | Sharp band inside glass zone; intensity confounded by glass (CV 17 %) | — | Uncertain | Inside glass zone (1308–1455 cm-1) | 153
NAM | Uncertain | 1392.5 | 1392.5 ± 0.5 | 607 | 111 | 18.3 | 34.1 | 364 | 15/15 | — | — | Sharp band inside glass zone; intensity confounded by glass (CV 18 %) | — | Uncertain | Inside glass zone (1308–1455 cm-1) | 142
NAM | Uncertain | 1450.0 | 1450.0 ± 0.5 | 499 | 50 | 10 | 27.9 | 308 | 15/15 | 1440–1460 (CH2/CH3 def.) | ≈ 0 | Proposed: δ(CH2/CH3); inside glass zone | L7 | Uncertain | Inside glass zone (1308–1455 cm-1) | 343

## 5_Excluded_Data
Sheet 5 — Excluded, flagged and low-SNR spectra (none deleted)
File name | Original file | Class | Power (%) | Integration (s) | Acc | Decision | Reason | Library folder
LTA_Bulk_20260403_P050pct_T30s_A15_S01.csv | 4-4-26/LTA/30sec50power15set/SP_6.csv | LTA | 50 | 30 | 15 | Abnormal - review | Distorted spectrum (r = 0.45 vs LTA class mean, glass region excluded) with low SNR; not explained by power/time. Not used as reference. | LTA_Bulk
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S01.csv | 9-2-2026/20sec20power-variable location LTA/SP_2.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S02.csv | 9-2-2026/20sec20power-variable location LTA/SP_3.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S03.csv | 9-2-2026/20sec20power-variable location LTA/SP_4.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S04.csv | 9-2-2026/20sec20power-variable location LTA/SP_5.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S05.csv | 9-2-2026/20sec20power-variable location LTA/SP_6.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S06.csv | 9-2-2026/20sec20power-variable location LTA/SP_7.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S07.csv | 9-2-2026/20sec20power-variable location LTA/SP_8.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S08.csv | 9-2-2026/20sec20power-variable location LTA/SP_10.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S09.csv | 9-2-2026/20sec20power-variable location LTA/SP_11.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S10.csv | 9-2-2026/20sec20power-variable location LTA/SP_12.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S11.csv | 9-2-2026/20sec20power-variable location LTA/SP_13.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S12.csv | 9-2-2026/20sec20power-variable location LTA/SP_14.csv | SUSPECT | 20 | 20 | 2 | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S01.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot1.csv | NAM | 90 | 60 | 5 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 18625). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 310 counts (reference ≈ 26), r = 0.46. Fingerprint unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S02.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot2.csv | NAM | 90 | 60 | 5 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 22235). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 554 counts (reference ≈ 26), r = 0.26. Fingerprint unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S03.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot3.csv | NAM | 90 | 60 | 5 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 27829). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 622 counts (reference ≈ 26), r = 0.17. Fingerprint unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S04.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot4.csv | NAM | 90 | 60 | 5 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 28506). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 694 counts (reference ≈ 26), r = 0.12. Fingerprint unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S05.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot5.csv | NAM | 90 | 60 | 5 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 36617). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 1075 counts (reference ≈ 26), r = 0.09. Fingerprint unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S06.csv | 29-05-26/90p-60s-5a-samespot/samespot1.csv | NAM | 90 | 60 | 5 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 88 % of 400–1800 cm-1 (min -2877); r = 0.36. Unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S07.csv | 29-05-26/90p-60s-5a-samespot/samespot2.csv | NAM | 90 | 60 | 5 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 96 % of 400–1800 cm-1 (min -6049); r = 0.19. Unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S08.csv | 29-05-26/90p-60s-5a-samespot/samespot3.csv | NAM | 90 | 60 | 5 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 97 % of 400–1800 cm-1 (min -8448); r = 0.17. Unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S09.csv | 29-05-26/90p-60s-5a-samespot/samespot4.csv | NAM | 90 | 60 | 5 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 98 % of 400–1800 cm-1 (min -9895); r = 0.13. Unreliable. | Excluded
NAM_Bulk_20260530_P090pct_T60s_A05_S10.csv | 29-05-26/90p-60s-5a-samespot/samespot5.csv | NAM | 90 | 60 | 5 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 99 % of 400–1800 cm-1 (min -11649); r = 0.09. Unreliable. | Excluded
LTA_Bulk_20260403_P030pct_T05s_A10_S01.csv | 4-4-26/5sec,30,50 power 10 set/SP_12.csv | LTA | 30 | 5 | 10 | Low SNR | SNR 6.7 < 10 at low power / short integration; genuine LTA spectrum used only for parameter comparison. | LTA_Bulk
LTA_Bulk_20260403_P050pct_T05s_A10_S01.csv | 4-4-26/5sec,30,50 power 10 set/SP_13.csv | LTA | 50 | 5 | 10 | Low SNR | SNR 8.4 < 10 at low power / short integration; genuine LTA spectrum used only for parameter comparison. | LTA_Bulk
NAM_Bulk_20260128_P005pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_21.csv | NAM | 5 | 5 | 2 | Low SNR | SNR 2.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P005pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_37.csv | NAM | 5 | 10 | 2 | Low SNR | SNR 3.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P005pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_4.csv | NAM | 5 | 15 | 2 | Low SNR | SNR 5.0 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P005pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_54.csv | NAM | 5 | 15 | 2 | Low SNR | SNR 5.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P005pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_70.csv | NAM | 5 | 20 | 2 | Low SNR | SNR 4.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P005pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_86.csv | NAM | 5 | 25 | 2 | Low SNR | SNR 6.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P010pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_22.csv | NAM | 10 | 5 | 2 | Low SNR | SNR 3.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P010pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_38.csv | NAM | 10 | 10 | 2 | Low SNR | SNR 4.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P010pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_5.csv | NAM | 10 | 15 | 2 | Low SNR | SNR 6.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P010pct_T15s_A02_S02.csv | NAM-15secs and power15-75-2sets/SP_6.csv | NAM | 10 | 15 | 2 | Low SNR | SNR 6.0 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P010pct_T15s_A02_S03.csv | 15sec-power5-80andset2/SP_55.csv | NAM | 10 | 15 | 2 | Low SNR | SNR 6.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P010pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_71.csv | NAM | 10 | 20 | 2 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P010pct_T25s_A02_S01.csv | 25sec-power5-80 and set2/SP_87.csv | NAM | 10 | 25 | 2 | Low SNR | SNR 7.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P015pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_23.csv | NAM | 15 | 5 | 2 | Low SNR | SNR 2.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P015pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_39.csv | NAM | 15 | 10 | 2 | Low SNR | SNR 5.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P015pct_T15s_A02_S01.csv | 15sec-power5-80andset2/SP_56.csv | NAM | 15 | 15 | 2 | Low SNR | SNR 7.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P015pct_T20s_A02_S01.csv | 20sec-power5-80and set2/SP_72.csv | NAM | 15 | 20 | 2 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P020pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_24.csv | NAM | 20 | 5 | 2 | Low SNR | SNR 4.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P020pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_40.csv | NAM | 20 | 10 | 2 | Low SNR | SNR 6.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P020pct_T10s_A02_S02.csv | 10sec-power5-80-2set/SP_41.csv | NAM | 20 | 10 | 2 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P020pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_7.csv | NAM | 20 | 15 | 2 | Low SNR | SNR 9.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P020pct_T15s_A02_S02.csv | 15sec-power5-80andset2/SP_57.csv | NAM | 20 | 15 | 2 | Low SNR | SNR 9.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P025pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_25.csv | NAM | 25 | 5 | 2 | Low SNR | SNR 4.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P025pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_42.csv | NAM | 25 | 10 | 2 | Low SNR | SNR 8.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P025pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_8.csv | NAM | 25 | 15 | 2 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P030pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_26.csv | NAM | 30 | 5 | 2 | Low SNR | SNR 5.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P030pct_T10s_A02_S01.csv | 10sec-power5-80-2set/SP_43.csv | NAM | 30 | 10 | 2 | Low SNR | SNR 8.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P035pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_27.csv | NAM | 35 | 5 | 2 | Low SNR | SNR 5.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P040pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_28.csv | NAM | 40 | 5 | 2 | Low SNR | SNR 6.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P045pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_29.csv | NAM | 45 | 5 | 2 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P050pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_30.csv | NAM | 50 | 5 | 2 | Low SNR | SNR 7.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P055pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_31.csv | NAM | 55 | 5 | 2 | Low SNR | SNR 7.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P060pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_32.csv | NAM | 60 | 5 | 2 | Low SNR | SNR 8.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P065pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_33.csv | NAM | 65 | 5 | 2 | Low SNR | SNR 8.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P070pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_34.csv | NAM | 70 | 5 | 2 | Low SNR | SNR 8.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P075pct_T05s_A02_S01.csv | 5sec-power5-80 and 2set/SP_35.csv | NAM | 75 | 5 | 2 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260529_P005pct_T05s_A01_S01.csv | 29-05-26/glass slide empty/NAM_5P_S1.csv | NAM | 5 | 5 | 1 | Low SNR | SNR 2.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260529_P005pct_T05s_A01_S02.csv | 29-05-26/glass slide empty/NAM_5P_S2.csv | NAM | 5 | 5 | 1 | Low SNR | SNR 2.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260529_P010pct_T05s_A01_S01.csv | 29-05-26/glass slide empty/NAM_10P_S1.csv | NAM | 10 | 5 | 1 | Low SNR | SNR 2.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260529_P010pct_T05s_A01_S02.csv | 29-05-26/glass slide empty/NAM_10P_S2.csv | NAM | 10 | 5 | 1 | Low SNR | SNR 2.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260529_P010pct_T05s_A01_S03.csv | 29-05-26/glass slide empty/NAM_10P_S3.csv | NAM | 10 | 5 | 1 | Low SNR | SNR 2.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260529_P010pct_T05s_A01_S04.csv | 29-05-26/glass slide empty/NAM_10P_S4.csv | NAM | 10 | 5 | 1 | Low SNR | SNR 2.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | NAM_Bulk
NAM_Bulk_20260128_P050pct_T15s_A02_S01.csv | NAM-15secs and power15-75-2sets/SP_15.csv | NAM | 50 | 15 | 2 | Review - higher noise | Correct NAM band pattern but r = 0.84 vs class mean; extra noise structure. | NAM_Bulk
LPS_Bulk_20260413_P010pct_T10s_A02_S01.csv | 14-426/lps/SP_1.csv | LPS | 10 | 10 | 2 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P010pct_T10s_A02_S02.csv | 14-426/lps/SP_2.csv | LPS | 10 | 10 | 2 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P010pct_T10s_A02_S03.csv | 14-426/lps/SP_3.csv | LPS | 10 | 10 | 2 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P010pct_T10s_A05_S01.csv | 13-4-26/lps/20-20-5/SP_1.csv | LPS | 10 | 10 | 5 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P010pct_T10s_A05_S02.csv | 13-4-26/lps/20-20-5/SP_2.csv | LPS | 10 | 10 | 5 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P010pct_T60s_A30_S01.csv | 13-4-26/lps/20-20-5/SP_6.csv | LPS | 10 | 60 | 30 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P020pct_T10s_A02_S01.csv | 13-4-26/lps/20-20-5/SP_4.csv | LPS | 20 | 10 | 2 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P020pct_T10s_A05_S01.csv | 13-4-26/lps/20-20-5/SP_3.csv | LPS | 20 | 10 | 5 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P020pct_T30s_A50_S01.csv | 14-426/lps/SP_5.csv | LPS | 20 | 30 | 50 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P020pct_T30s_A60_S01.csv | 14-426/lps/SP_4.csv | LPS | 20 | 30 | 60 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P020pct_T60s_A30_S01.csv | 13-4-26/lps/20-20-5/SP_5.csv | LPS | 20 | 60 | 30 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P020pct_T60s_A30_S02.csv | 13-4-26/lps/20-20-5/SP_8.csv | LPS | 20 | 60 | 30 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P020pct_T60s_A50_S01.csv | 13-4-26/lps/20-20-5/SP_7.csv | LPS | 20 | 60 | 50 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
LPS_Bulk_20260413_P020pct_T60s_A50_S02.csv | 13-4-26/lps/20-20-5/SP_9.csv | LPS | 20 | 60 | 50 | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | LPS_Bulk
NAM_Bulk_20260530_P070pct_T60s_A05_S01.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot1.csv | NAM | 70 | 60 | 5 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 8421 counts (< 65 535), no negative counts, SNR 35.5, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | NAM_Bulk
NAM_Bulk_20260530_P070pct_T60s_A05_S02.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot2.csv | NAM | 70 | 60 | 5 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 7732 counts (< 65 535), no negative counts, SNR 35.1, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | NAM_Bulk
NAM_Bulk_20260530_P070pct_T60s_A05_S03.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot3.csv | NAM | 70 | 60 | 5 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 9147 counts (< 65 535), no negative counts, SNR 33.9, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | NAM_Bulk

## 6_Literature
Sheet 6 — Literature used
No peer-reviewed normal-Raman spectrum of pure N-acetylmuramic acid or isolated LTA powder was located in this search; NAM/LTA assignments rely on group frequencies (L5–L7).
ID | Full citation | DOI | Biomarker | Technique | Organism / source | Access note
L1 | Wu X., Zhao Y., Zughaier S.M. Highly sensitive detection and differentiation of endotoxins derived from bacterial pathogens by surface-enhanced Raman scattering. Biosensors 2021, 11(7), 234. | 10.3390/bios11070234 | LPS / LOS | SERS | Neisseria meningitidis LOS/LPS; enteric LPS | Band table read from publisher page (MDPI)
L2 | Yang Y., Xu B., Haverstick J., Ibtehaz N., Muszyński A., Chen X., Chowdhury M.E.H., Zughaier S.M., Zhao Y. Differentiation and classification of bacterial endotoxins based on surface enhanced Raman scattering and advanced machine learning. Nanoscale 2022, 14(24), 8806–8817. | 10.1039/d2nr01277d | LPS (multiple species), LTA (B. subtilis), PGN | SERS, 785 nm, Ag nanorod array | Commercial endotoxins | Read from publisher PDF (RSC); detailed assignments in their Table S6 (not accessed)
L3 | Saikia D., Tanan C.B., Dhananjaya G., Hungund B.S., Mahanta N., Singh S.P. Validating phosphoethanolamine modification as a potential spectral marker of colistin resistance. Analyst 2025, 150, 281–289. | 10.1039/d4an01228c | LPS / lipid A (in whole cells) | Raman, 532 nm, whole bacterial cells | A. baumannii, E. coli (whole cells) | Read from publisher PDF (RSC)
L4 | Kumar A., Islam M.R., Zughaier S.M., Chen X., Zhao Y. Precision classification and quantitative analysis of bacteria biomarkers via surface-enhanced Raman spectroscopy and machine learning. Spectrochim. Acta A 2024, 320, 124627. | 10.1016/j.saa.2024.124627 | LTA (among 6 biomarkers) | SERS | Commercial LTA | Abstract only; LTA band positions not accessible
L5 | Movasaghi Z., Rehman S., Rehman I.U. Raman spectroscopy of biological tissues. Appl. Spectrosc. Rev. 2007, 42(5), 493–541. | 10.1080/05704920701551530 | General (group frequencies) | Raman review | — | Citation verified (Crossref)
L6 | Wiercigroch E., Szafraniec E., Czamara K., Pacia M.Z., Majzner K., Kochan K., Kaczor A., Baranska M., Malek K. Raman and infrared spectroscopy of carbohydrates: a review. Spectrochim. Acta A 2017, 185, 317–335. | 10.1016/j.saa.2017.05.045 | Carbohydrates (group frequencies) | Raman/IR review | — | Citation verified (Crossref)
L8 | Grunert T., Jovanovic D., Sirisarn W., et al. Analysis of Staphylococcus aureus wall teichoic acid glycoepitopes by Fourier Transform Infrared Spectroscopy provides novel insights into the staphylococcal glycocode. Sci. Rep. 2018, 8, 1889. | 10.1038/s41598-018-20222-6 | Wall teichoic acid (related to LTA backbone chemistry) | FTIR, intact S. aureus cells | S. aureus whole cells | Read from publisher page (Nature)
L9 | Zou B., Zhao X., Liao J., Zhao Y. Near-infrared activity analysis and identification of fungal and bacterial markers. Microchem. J. 2025, 208, 112552. | 10.1016/j.microc.2024.112552 | N-acetylmuramic acid, N-acetylglucosamine | NIR + DFT-calculated IR | Pure compounds (NIR); calculated IR | Abstract/summary only; no Raman band positions reported
L7 | Czamara K., Majzner K., Pacia M.Z., Kochan K., Kaczor A., Baranska M. Raman spectroscopy of lipids: a review. J. Raman Spectrosc. 2015, 46, 4–20. | 10.1002/jrs.4607 | Lipids / phospholipids (group frequencies) | Raman review | — | Citation verified (Crossref)
L10 | Bansil R., Yannas I.V., Stanley H.E. Raman spectroscopy: a structural probe of glycosaminoglycans. Biochim. Biophys. Acta 1978, 541, 535–542. | 10.1016/0304-4165(78)90163-0 | Glycosaminoglycans (N-acetylated amino-sugar polymers; related to NAM chemistry) | Normal Raman, pure polysaccharides | Purified glycosaminoglycans | Citation verified (Crossref); full text not accessed – band values not extracted
L11 | Study of N-acetylamino saccharides with synchrotron-based ultraviolet resonance Raman spectroscopy: in combination with ATR far-ultraviolet spectroscopy. J. Phys. Chem. Lett. 2025, 16(7), 1769–. | 10.1021/acs.jpclett.4c03435 | N-acetylamino saccharides (compounds not verified; NAM inclusion unknown) | UV resonance Raman (not 785 nm normal Raman) | Pure saccharides | Identified via Crossref/ACS; full text not accessible (403) – not used for assignments

## 7_Literature_Bands
Sheet 7 — Literature band list
Biomarker / group | Ref ID | Technique | Reported band (cm-1) | Reported assignment | Comparison with our data | Note
LPS | L1 | SERS | 550 | β(CH2) in ring | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L1 | SERS | 680 | δ(C-O-C); fatty acid | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L1 | SERS | 850 | ν(C-O-C) 1,4-glycosidic link | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L1 | SERS | 988 | β(CH) carbohydrates | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L1 | SERS | 1131 | ν(C-C) fatty acid | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L1 | SERS | 1330 | δ(CH) phospholipid | No LPS experimental peak established ¦ lies in glass zone | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L1 | SERS | 1450 | CH3/CH2 deformation | No LPS experimental peak established ¦ lies in glass zone | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L1 | SERS (enteric LPS) | 543 | saccharide and lipid | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L1 | SERS (enteric LPS) | 981 | saccharide and lipid | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L2 | SERS 785 nm | 735 | common LPS peak (assignment in source SI) | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L2 | SERS 785 nm | 794 | common LPS peak | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L2 | SERS 785 nm | 894 | common LPS peak | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L2 | SERS 785 nm | 917 | common LPS peak | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L2 | SERS 785 nm | 1003 | common LPS peak | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L2 | SERS 785 nm | 1333 | common LPS peak (C-H deformation) | No LPS experimental peak established ¦ lies in glass zone | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L2 | SERS 785 nm | 1592 | common LPS peak | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L2 | SERS 785 nm | 1614 | common LPS peak | No LPS experimental peak established | SERS bands can differ from bulk Raman (surface selection rules)
LPS | L3 | Raman 532 nm, whole cells | 1103 | lipopolysaccharide (A. baumannii) | No LPS experimental peak established | Whole-cell context
LTA | L2 | SERS 785 nm | 894 | C-O-H deformation (shared with chitin) | closest exp. 861 (-33 cm-1) – no match | SERS bands can differ from bulk Raman (surface selection rules)
LTA | L2 | SERS 785 nm | 1333 | C-H deformation (shared with PGN) | closest exp. 1459 (+126 cm-1) – no match | SERS bands can differ from bulk Raman (surface selection rules)
Phosphodiester (LTA backbone) | L5 | Raman group frequency | 1090 | ν_s(PO2−) phosphodiester | closest exp. 1092 (+2 cm-1) – consistent | Group frequency
Phosphodiester (LTA backbone) | L5 | Raman group frequency | 813 | ν(O-P-O) phosphodiester | closest exp. 802 (-11 cm-1) – consistent | Group frequency
Lipid acyl chains | L7 | Raman group frequency | 1440–1460 | δ(CH2) scissoring | region – see Sheet 4 | Group frequency
Carbohydrate (NAM sugar ring) | L6 | Raman group frequency | < 600 | pyranose ring skeletal deformations | region – see Sheet 4 | Group frequency
Carbohydrate (NAM sugar ring) | L6 | Raman group frequency | 800–950 | anomeric region: ν(C-O-C), δ(C-C-H) | region – see Sheet 4 | Group frequency
Carbohydrate (NAM sugar ring) | L6 | Raman group frequency | 1000–1200 | ν(C-O), ν(C-C), δ(C-O-H) | region – see Sheet 4 | Group frequency

## 8_All_Reproducible_Peaks
Sheet 8 — All peaks detected in reference replicates (prominence ≥ 5 × noise)
Full list for traceability; the report highlights only the selected marker peaks.
Biomarker | Peak (cm-1) | Replicates detected | n | Frequency (%) | Relative intensity | In glass zone
NAM | 457 | 15 | 15 | 100 | 0.6 | No
NAM | 487 | 15 | 15 | 100 | 0.43 | No
NAM | 521 | 15 | 15 | 100 | 0.22 | No
NAM | 541 | 15 | 15 | 100 | 0.78 | No
NAM | 584 | 14 | 15 | 93.33 | 0.29 | No
NAM | 614 | 15 | 15 | 100 | 0.17 | No
NAM | 645 | 15 | 15 | 100 | 0.22 | No
NAM | 702 | 15 | 15 | 100 | 0.22 | No
NAM | 743 | 15 | 15 | 100 | 0.16 | No
NAM | 829 | 15 | 15 | 100 | 0.68 | No
NAM | 871 | 15 | 15 | 100 | 0.95 | No
NAM | 900 | 8 | 15 | 53.33 | 0.19 | No
NAM | 929 | 15 | 15 | 100 | 1 | No
NAM | 955 | 15 | 15 | 100 | 0.94 | No
NAM | 1023 | 15 | 15 | 100 | 0.82 | No
NAM | 1059 | 15 | 15 | 100 | 0.31 | No
NAM | 1084 | 15 | 15 | 100 | 0.28 | No
NAM | 1127 | 15 | 15 | 100 | 0.63 | No
NAM | 1155 | 15 | 15 | 100 | 0.4 | No
NAM | 1249 | 15 | 15 | 100 | 0.32 | No
NAM | 1319 | 15 | 15 | 100 | 0.56 | Yes
NAM | 1345 | 15 | 15 | 100 | 0.68 | Yes
NAM | 1392 | 15 | 15 | 100 | 0.66 | Yes
NAM | 1450 | 15 | 15 | 100 | 0.54 | Yes
NAM | 1589 | 15 | 15 | 100 | 0.29 | No
NAM | 1701 | 15 | 15 | 100 | 0.15 | No
LTA | 416 | 2 | 4 | 50 | 0.25 | No
LTA | 539 | 3 | 4 | 75 | 0.38 | No
LTA | 549 | 1 | 4 | 25 | 0.3 | No
LTA | 629 | 1 | 4 | 25 | 0.24 | No
LTA | 642 | 1 | 4 | 25 | 0.2 | No
LTA | 652 | 1 | 4 | 25 | 0.2 | No
LTA | 725 | 1 | 4 | 25 | 0.21 | No
LTA | 798 | 2 | 4 | 50 | 0.73 | No
LTA | 806 | 2 | 4 | 50 | 0.71 | No
LTA | 862 | 4 | 4 | 100 | 0.62 | No
LTA | 1010 | 2 | 4 | 50 | 0.2 | No
LTA | 1092 | 4 | 4 | 100 | 1 | No
LTA | 1260 | 1 | 4 | 25 | 0.16 | No
LTA | 1328 | 1 | 4 | 25 | 0.46 | Yes
LTA | 1370 | 1 | 4 | 25 | 0.44 | Yes
LTA | 1459 | 4 | 4 | 100 | 0.65 | No
LTA | 1576 | 1 | 4 | 25 | 0.19 | No
LTA | 1583 | 1 | 4 | 25 | 0.19 | No
LTA | 1613 | 1 | 4 | 25 | 0.16 | No
LTA | 1672 | 1 | 4 | 25 | 0.11 | No
LTA | 1751 | 3 | 4 | 75 | 0.18 | No

## 9_Reproducibility
Sheet 9 — Spectral reproducibility (Pearson r of normalised spectra, 400–1800 cm-1)
Correlation measures reproducibility, not molecular identity.
Biomarker | n | Mean pairwise r | Min pairwise r | Mean r to mean spectrum | Min r to mean spectrum
NAM | 15 | 0.98 | 0.879 | 0.991 | 0.95
LTA | 4 | 0.955 | 0.927 | 0.983 | 0.974

## 10_Bulk_Intensity
Sheet 10 — Bulk powder Raman intensity of the reference sets (400–1800 cm-1)
Raw counts include dark (~1200 counts) and background; marker intensity is baseline-corrected peak height.
Biomarker | Condition | n | Max raw counts mean | Max raw counts SD | Max dark-subtracted mean | Max dark-subtracted SD | Baseline mean (counts) | Strongest marker intensity mean | Strongest marker intensity SD | CV (%) | SNR mean | SNR SD | Dates
NAM | 90% / 30 s / 5 acc | 15 | 6055.2 | 323.5 | 2569.8 | 391.9 | 1357.5 | 916.8 | 36.8 | 4 | 51.1 | 4.7 | 20260530
LTA | 80-95% / 45 s / 2 acc | 4 | 16387.4 | 4559.8 | 13683.9 | 5061.9 | 8000.2 | 1110.1 | 97.8 | 8.8 | 28.5 | 1 | 20260403

## 11_Future_Acquisition
Sheet 11 — Future measurement quick reference
Biomarker | Starting condition | Notes
NAM | 785 nm, 90 %, 30 s, 5 acc | Replicate at reference condition (independent spots/slides); matched glass blank
LTA | 785 nm, 80 %, 45 s, 2 acc | Preliminary reference (n = 4): add ≥ 6 independent spots (ideally a second day) to reach n ≥ 10; matched glass blank
LPS | 785 nm, 70 / 80 / 90 %, 30 s, 5 acc (screening) | ≥ 5 independent spots per power on glass slide; select reference only after comparison with matched blank
Glass blank | Same settings as the corresponding sample | Same instrument, wavelength, power, integration, accumulations, objective and slide type
If saturation occurs | Reduce power and/or integration | Check raw counts in 400–1800 cm-1 and at the laser line; document the change

## 12_NAM_60s_ReAudit
Sheet 12 — Individual re-audit of the 13 NAM 60 s × 5 acc spectra (A valid / B saturated / C distorted / D other)
No spectrum is saturated inside 400–1800 cm-1 (class B = 0). Saturation occurs only at the laser line (≈ −2 to −4 cm-1).
file | original | power | category | decision | saturated_pixels | saturated_shift | max_raw_400_1800 | dark_median | min_dark_subtracted | neg_fraction_pct | noise | SNR | r_vs_reference | reason
NAM_Bulk_20260530_P070pct_T60s_A05_S01.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot1.csv | 70 | A – Valid candidate | Valid – supporting (70%/60 s/5 acc) | 1 | -2.2 to -2.2 | 8421 | 1959 | 1009 | 0 | 35 | 35.5 | 0.998 | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 8421 counts (< 65 535), no negative counts, SNR 35.5, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference).
NAM_Bulk_20260530_P070pct_T60s_A05_S02.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot2.csv | 70 | A – Valid candidate | Valid – supporting (70%/60 s/5 acc) | 1 | -2.2 to -2.2 | 7732 | 1959 | 751 | 0 | 31.1 | 35.1 | 0.998 | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 7732 counts (< 65 535), no negative counts, SNR 35.1, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference).
NAM_Bulk_20260530_P070pct_T60s_A05_S03.csv | 29-05-26/70p-60s-5ac-diifrentpoint/spot3.csv | 70 | A – Valid candidate | Valid – supporting (70%/60 s/5 acc) | 1 | -2.2 to -2.2 | 9147 | 1959 | 1414 | 0 | 35.8 | 33.9 | 0.998 | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 9147 counts (< 65 535), no negative counts, SNR 33.9, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference).
NAM_Bulk_20260530_P090pct_T60s_A05_S01.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot1.csv | 90 | C – Distorted | Excluded – distorted (abnormal dark / ripple) | 2 | -4.2 to -2.2 | 18625 | 5910 | 1307 | 0 | 309.9 | 11.2 | 0.457 | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 18625). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 310 counts (reference ≈ 26), r = 0.46. Fingerprint unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S02.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot2.csv | 90 | C – Distorted | Excluded – distorted (abnormal dark / ripple) | 2 | -4.2 to -2.2 | 22235 | 5910 | 1563 | 0 | 553.9 | 11.2 | 0.258 | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 22235). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 554 counts (reference ≈ 26), r = 0.26. Fingerprint unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S03.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot3.csv | 90 | C – Distorted | Excluded – distorted (abnormal dark / ripple) | 2 | -4.2 to -2.2 | 27829 | 5910 | 2958 | 0 | 622.1 | 11.9 | 0.167 | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 27829). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 622 counts (reference ≈ 26), r = 0.17. Fingerprint unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S04.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot4.csv | 90 | C – Distorted | Excluded – distorted (abnormal dark / ripple) | 2 | -4.2 to -2.2 | 28506 | 5910 | 2225 | 0 | 694.2 | 12 | 0.125 | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 28506). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 694 counts (reference ≈ 26), r = 0.12. Fingerprint unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S05.csv | 29-05-26/90p-60s-5ac-diffrentspot/spot5.csv | 90 | C – Distorted | Excluded – distorted (abnormal dark / ripple) | 2 | -4.2 to -2.2 | 36617 | 5910 | 2696 | 0 | 1075 | 12.4 | 0.087 | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 36617). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 1075 counts (reference ≈ 26), r = 0.09. Fingerprint unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S06.csv | 29-05-26/90p-60s-5a-samespot/samespot1.csv | 90 | C – Distorted | Excluded – distorted (negative counts) | 2 | -4.2 to -2.2 | 34426 | 26587 | -2877 | 88.4 | 285.7 | 8.6 | 0.363 | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 88 % of 400–1800 cm-1 (min -2877); r = 0.36. Unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S07.csv | 29-05-26/90p-60s-5a-samespot/samespot2.csv | 90 | C – Distorted | Excluded – distorted (negative counts) | 2 | -4.2 to -2.2 | 30950 | 26587 | -6049 | 96 | 457.4 | 7.7 | 0.194 | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 96 % of 400–1800 cm-1 (min -6049); r = 0.19. Unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S08.csv | 29-05-26/90p-60s-5a-samespot/samespot3.csv | 90 | C – Distorted | Excluded – distorted (negative counts) | 2 | -4.2 to -2.2 | 28976 | 26587 | -8448 | 97.2 | 523.4 | 7.4 | 0.169 | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 97 % of 400–1800 cm-1 (min -8448); r = 0.17. Unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S09.csv | 29-05-26/90p-60s-5a-samespot/samespot4.csv | 90 | C – Distorted | Excluded – distorted (negative counts) | 2 | -4.2 to -2.2 | 27167 | 26587 | -9895 | 98.4 | 594.5 | 7.3 | 0.128 | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 98 % of 400–1800 cm-1 (min -9895); r = 0.13. Unreliable.
NAM_Bulk_20260530_P090pct_T60s_A05_S10.csv | 29-05-26/90p-60s-5a-samespot/samespot5.csv | 90 | C – Distorted | Excluded – distorted (negative counts) | 2 | -4.2 to -2.2 | 25200 | 26587 | -11649 | 98.9 | 696.6 | 7.1 | 0.095 | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 99 % of 400–1800 cm-1 (min -11649); r = 0.09. Unreliable.

## 13_NAM_70pct_vs_Reference
Sheet 13 — NAM 70 %/60 s/5 acc (n = 3, supporting group) vs reference 90 %/30 s/5 acc (n = 15)
Decision: 70 % spectra are ADDITIONAL SUPPORTING NAM spectra forming a separate acquisition-condition group; NOT merged into the reference (different power and integration time; n = 3; single session). Positions agree within ≤ 1 cm-1; r to reference = 0.998.
Group | Peak (nominal) | n | Position mean | Position SD | Intensity mean | Intensity SD | CV (%) | SNR | Prominence | Prominence/noise | Detected | Relative intensity | Glass zone
REF 90%/30s/5acc | 458 | 15 | 457 | 0.38 | 546.79 | 31.88 | 5.83 | 30.48 | 560.33 | 31.24 | 15 | 0.6 | No
REF 90%/30s/5acc | 541 | 15 | 541 | 0 | 715.17 | 32.57 | 4.55 | 39.86 | 736.02 | 41.01 | 15 | 0.78 | No
REF 90%/30s/5acc | 830 | 15 | 829 | 0 | 624.94 | 30.2 | 4.83 | 34.85 | 570.88 | 31.83 | 15 | 0.68 | No
REF 90%/30s/5acc | 870 | 15 | 870.53 | 0.52 | 871.56 | 37.07 | 4.25 | 48.57 | 820.46 | 45.74 | 15 | 0.95 | No
REF 90%/30s/5acc | 929 | 15 | 929 | 0 | 916.84 | 36.85 | 4.02 | 51.09 | 944.82 | 52.63 | 15 | 1 | No
REF 90%/30s/5acc | 955 | 15 | 955 | 0 | 862.12 | 44.31 | 5.14 | 48.07 | 830.02 | 46.29 | 15 | 0.94 | No
REF 90%/30s/5acc | 1023 | 15 | 1022.73 | 0.46 | 748.96 | 46.79 | 6.25 | 41.79 | 782.53 | 43.65 | 15 | 0.82 | No
REF 90%/30s/5acc | 1127 | 15 | 1127 | 0 | 581.26 | 30.13 | 5.18 | 32.41 | 519.3 | 28.99 | 15 | 0.63 | No
70%/60s/5acc | 458 | 3 | 457.33 | 0.58 | 814.56 | 47.44 | 5.82 | 36.93 | 836.2 | 37.91 | 3 | 0.6 | No
70%/60s/5acc | 541 | 3 | 541 | 0 | 1072.49 | 75.51 | 7.04 | 48.63 | 1104.59 | 50.1 | 3 | 0.79 | No
70%/60s/5acc | 830 | 3 | 829.33 | 0.58 | 923.01 | 76.37 | 8.27 | 41.86 | 846 | 38.35 | 3 | 0.68 | No
70%/60s/5acc | 870 | 3 | 871 | 0 | 1300.53 | 81.63 | 6.28 | 59 | 1237.95 | 56.15 | 3 | 0.96 | No
70%/60s/5acc | 929 | 3 | 929 | 0 | 1350.19 | 95.67 | 7.09 | 61.23 | 1388.29 | 62.95 | 3 | 1 | No
70%/60s/5acc | 955 | 3 | 955 | 0 | 1267.67 | 132.82 | 10.48 | 57.45 | 1219.22 | 55.27 | 3 | 0.94 | No
70%/60s/5acc | 1023 | 3 | 1022.67 | 0.58 | 1106.52 | 80.05 | 7.23 | 50.18 | 1145.18 | 51.91 | 3 | 0.82 | No
70%/60s/5acc | 1127 | 3 | 1127 | 0 | 859.84 | 77.33 | 8.99 | 39 | 768.94 | 34.86 | 3 | 0.64 | No
Group | n | SNR mean | SNR SD | Mean pairwise r | Mean r to reference mean (glass zone excl.) | Strongest marker height (counts) | SD | Baseline (counts)
REF 90%/30s/5acc | 15 | 51.092 | 4.658 | 0.98 | 0.997 | 916.84 | 36.846 | 1357.547
70%/60s/5acc | 3 | 61.228 | 6.331 | 0.974 | 0.999 | 1350.188 | 95.668 | 1974.45

## 14_NAM_CrossDay_Detection
Sheet 14 — NAM peak detection frequency across measurement days (all Valid NAM spectra; prominence ≥ 5 × noise)
Used for Criterion 1 (reproducibility) of the NAM peak hierarchy.
Peak (nominal) | Detected % 20260128 | Detected % 20260403 | Detected % 20260529 | Detected % 20260530 | Mean detection %
458 | 97 | 100 | 100 | 100 | 99
487 | 87 | 100 | 100 | 100 | 97
541 | 100 | 100 | 100 | 100 | 100
584 | 70 | 96 | 85 | 79 | 83
830 | 100 | 100 | 100 | 100 | 100
870 | 100 | 100 | 100 | 100 | 100
929 | 100 | 100 | 100 | 100 | 100
955 | 100 | 100 | 100 | 100 | 100
1023 | 100 | 100 | 100 | 100 | 100
1059 | 0 | 82 | 1 | 64 | 37
1127 | 98 | 100 | 100 | 100 | 100
1155 | 52 | 100 | 82 | 96 | 83
1249 | 72 | 100 | 97 | 100 | 92
1345 | 74 | 100 | 97 | 100 | 93
1392 | 84 | 100 | 96 | 100 | 95
1450 | 95 | 95 | 86 | 100 | 94
1589 | 66 | 100 | 97 | 100 | 91

## 15_GO_NOGO
Sheet 15 — Formal GO / RE-ACQUIRE decision per biomarker (whole replicate group)
Biomarker | Existing n | Quality | Reproducibility | Marker consistency | Background | Decision | Reason
NAM | 15 at reference condition (5 independent points + 10 repeats of one spot); 224 valid NAM spectra on 4 days | High: SNR 51.1 ± 4.7; no saturation in 400–1800; stable baseline (~1360 counts) | High: pairwise r 0.994 (min 0.976); mean≈median (r 0.9991); 10 independent spots at 70 %/10 s: r 0.987 | 7/7 primary markers in 15/15 reference spectra, ≥ 98 % of valid spectra on every day and 10/10 independent spots; CV 4–6 % | Markers outside glass zone; 93–95 % of height retained after glass-template correction; hot pixels corrected | FINAL (28 Sep 2026, operator decision: no further NAM acquisition). GO for identity, markers and reference spectrum; independent spots at the reference condition n = 6 (< 10 target), accepted because 31 further independent spots at other conditions confirm the spectrum (sheets 26–32). | Only 5 of the 15 reference spectra are independent spots and all come from one session, so spot-to-spot and day-to-day intensity variation at 90 %/30 s/5 acc is not yet established.
LTA | 4 at reference condition (spot independence not recorded); 9 valid in total; 1 day | Adequate: SNR 28.5 ± 1.0; strong fluorescence background (~8 000 counts) | Moderate: pairwise r 0.952 (min 0.923); mean≈median (r 0.9990) but n = 4 | 3 primary markers in 4/4 reference and ≥ 8/9 valid spectra; CV 9–13 % | 1367 glass-associated; 1459 on glass shoulder (82 % retained); absolute intensities depend on ALS settings (−40 % to +25 %) | RE-ACQUIRE | n = 4 < 10, single day, spot independence unknown; intensities strongly baseline-dependent.
LPS | 14 (10–20 % power only); 0 at a suitable condition | Insufficient: strongest feature is the glass band; no LPS-specific band above noise after glass comparison | Poor: pairwise r 0.33 (range −0.47 to 1.0) | No reproducible candidate band | Dominated by glass band (1308–1455); fixed hot pixels at 1095.7 and 1189.5 cm-1 | RE-ACQUIRE | No LPS fingerprint established; power too low; no matched blanks.

## 16_Acquisition_Decision
Sheet 16 — Final acquisition decision and recommended re-acquisition conditions
Biomarker | Existing data sufficient? | Why | Re-acquisition needed? | Recommended condition | Target n
NAM | PARTLY – YES for identity/markers; NO for final intensity reference | Excellent quality and marker reproducibility, but only 5 independent spots at the reference condition, one session | NO (closed 28 Sep 2026) – top-up not performed by decision; limitation documented | 785 nm, 90 %, 30 s, 5 acc (unchanged; no saturation, SNR ≈ 51) | 10–15 independent spots (≥ 2 slides, ideally 2 days) + 3 matched glass blanks
LTA | NO | n = 4, single day, spot independence unknown; baseline-sensitive intensities | YES | 785 nm, 80 %, 45 s, 2 acc (lowest power in the 80–95 % plateau; SNR ≈ 27–28; max raw ≈ 16 000 counts) | 10–15 independent spots + 3 matched glass blanks
LPS | NO | Only 10–20 % power; spectrum dominated by glass band; no reproducible LPS band | NO (completed 28 Sep 2026: 10 spots at 60 %/40 s/5 acc) | Screening: 785 nm, 70 / 80 / 90 %, 30 s, 5 acc; ≥ 5 spots per power + ≥ 3 matched blanks per power; then 10–15 spots at the selected power | 10–15 at the selected condition

## 17_Final_Selection
Sheet 17 — Final spectrum selection (reference groups). NAM: 15 (FINAL, 28 Sep 2026); LTA: 4 (< 10 → re-acquire); LPS: 10 (ESTABLISHED 28 Sep 2026, sheets 38–42)
Selection rule: every spectrum acquired at the reference condition that passes QC is retained; none were dropped to improve statistics. The set is NOT filled with spectra from other conditions.
biomarker | file | spot | SNR | max_raw_400_1800 | saturated_400_1800 | spikes | r_to_group | baseline_mean | selected | reason
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S01.csv | same spot (repeat) | 52.1 | 6084 | No | 0 | 0.999 | 1405 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S02.csv | same spot (repeat) | 53.1 | 5983 | No | 0 | 0.998 | 1312 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S03.csv | same spot (repeat) | 51.9 | 5910 | No | 0 | 0.999 | 1258 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S04.csv | same spot (repeat) | 52.2 | 5871 | No | 0 | 0.999 | 1231 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S05.csv | same spot (repeat) | 57.8 | 5859 | No | 0 | 0.999 | 1224 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S06.csv | same spot (repeat) | 48.5 | 5886 | No | 0 | 0.999 | 1208 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S07.csv | same spot (repeat) | 54 | 5870 | No | 0 | 0.999 | 1211 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S08.csv | same spot (repeat) | 54.3 | 5839 | No | 0 | 0.999 | 1184 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S09.csv | same spot (repeat) | 51.6 | 5859 | No | 0 | 0.999 | 1189 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S10.csv | same spot (repeat) | 53.5 | 5865 | No | 0 | 0.999 | 1196 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S11.csv | different point | 56.6 | 5867 | No | 0 | 0.999 | 1177 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S12.csv | different point | 47.2 | 6977 | No | 0 | 0.994 | 2293 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S13.csv | different point | 48.5 | 6091 | No | 0 | 0.998 | 1311 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S14.csv | different point | 45.7 | 6370 | No | 0 | 0.995 | 1565 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
NAM | NAM_Bulk_20260530_P090pct_T30s_A05_S15.csv | different point | 39.2 | 6498 | No | 0 | 0.985 | 1632 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
LTA | LTA_Bulk_20260403_P080pct_T45s_A02_S01.csv | not recorded | 29.4 | 22619 | No | 0 | 0.974 | 11895 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
LTA | LTA_Bulk_20260403_P080pct_T45s_A02_S02.csv | not recorded | 28 | 11787 | No | 0 | 0.983 | 5247 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
LTA | LTA_Bulk_20260403_P090pct_T45s_A02_S01.csv | not recorded | 27.4 | 16296 | No | 0 | 0.987 | 8350 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)
LTA | LTA_Bulk_20260403_P095pct_T45s_A02_S01.csv | not recorded | 29.3 | 14848 | No | 0 | 0.982 | 7732 | YES | Passes all QC criteria; all spectra at the reference condition retained (no cherry-picking)

## 18_Mean_vs_Median
Sheet 18 — Mean vs median representative spectrum and group reproducibility
Decision rule: if r(mean, median) ≥ 0.999 and ¦mean − median¦ < 1 SD at every wavenumber and marker, the spectra are consistent and the MEAN is used (median reported alongside). Otherwise the median is used and outliers are investigated.
NAM_REF and LTA_REF meet the rule → mean selected. LPS_ALL: r pairs down to −0.47 → no representative spectrum.
Group | n | SNR mean | SNR SD | Pairwise r mean (glass zone excl.) | Pairwise r min | r(mean, median) | max ¦mean−median¦ / max(mean) | Fraction of points ¦mean−median¦ > 1 SD | Max ¦mean−median¦/SD at markers | Median pointwise CV (%)
NAM_REF | 15 | 51.0921 | 4.6575 | 0.9944 | 0.9757 | 0.9991 | 0.0555 | 0 | 0.5446 | 7.8963
NAM_REF_diffpoint | 5 | 47.4598 | 6.216 | 0.9895 | 0.9757 | 0.9901 | 0.1271 | 0 | 0.2636 | 10.5155
NAM_REF_samespot | 10 | 52.9082 | 2.3616 | 0.9985 | 0.9971 | 1 | 0.0091 | 0 | 0.4144 | 2.1817
NAM_VAL_70_10_5_diffpoint | 10 | 21.9207 | 3.2658 | 0.9874 | 0.9785 | 0.9993 | 0.0338 | 0 | 0.2887 | 18.1001
NAM_SUP_70_60_5 | 3 | 61.2276 | 6.3312 | 0.9981 | 0.9973 | 0.9954 | 0.0897 | 0 | 0.5548 | 9.0199
LTA_REF | 4 | 28.5234 | 1.0093 | 0.9516 | 0.9232 | 0.999 | 0.0461 | 0 | 0.4838 | 12.1036
LTA_ALLVALID | 9 | 22.5674 | 6.633 | 0.9501 | 0.8673 | 0.9892 | 0.1738 | 0 | 0.2777 | 77.9067
LPS_ALL | 14 | 21.9968 | 24.1604 | 0.3316 | -0.4699 | 0.9711 | 0.2589 | 0 |  | 155.0689
GLASS | 5 | 55.0205 | 2.6016 | 0.8979 | 0.8362 | 0.9999 | 0.0192 | 0 |  | 2.9246

## 19_Peak_Criteria
Sheet 19 — Peak promotion criteria (primary = C1 + C2 + C3 + C4)
C1: detected in all reference spectra and ≥95 % of valid spectra on every day (NAM) or ≥8/9 valid spectra (LTA) and all independent spots (NAM). Thresholds are project conventions.
Supporting: detected in all reference spectra and ≥90 % overall, but fails at least one primary criterion. Uncertain: otherwise. Glass/background: inside glass zone and removed by glass correction.
Biomarker | Peak (nominal) | Position mean ± SD | Detected in reference | Min detection on any day (%) / LTA all-valid (%) | Detected in independent-spot set (%) | SNR | CV (%) | Prominence (counts) | Inside glass zone | Height retained after glass-template correction (%) | Distance from range edge (cm-1) | C1 reproducible | C2 SNR≥15 & CV≤15 % | C3 background-separated (not glass zone, ≥90 % retained, ≥75 cm-1 from edge) | C4 literature-plausible | Category
NAM | 541 | 541.0 ± 0.0 | 15/15 | 100 | 100 | 39.9 | 4.6 | 736 | No | 93.9 | 141 | Yes | Yes | Yes | Yes | Primary marker
NAM | 830 | 829.0 ± 0.0 | 15/15 | 100 | 100 | 34.8 | 4.8 | 570.9 | No | 92.8 | 429 | Yes | Yes | Yes | Yes | Primary marker
NAM | 870 | 870.5 ± 0.5 | 15/15 | 100 | 100 | 48.6 | 4.3 | 820.5 | No | 94.8 | 470.5 | Yes | Yes | Yes | Yes | Primary marker
NAM | 929 | 929.0 ± 0.0 | 15/15 | 100 | 100 | 51.1 | 4 | 944.8 | No | 94.9 | 529 | Yes | Yes | Yes | Yes | Primary marker
NAM | 955 | 955.0 ± 0.0 | 15/15 | 100 | 100 | 48.1 | 5.1 | 830 | No | 95.4 | 555 | Yes | Yes | Yes | Yes | Primary marker
NAM | 1023 | 1022.7 ± 0.5 | 15/15 | 100 | 100 | 41.8 | 6.2 | 782.5 | No | 95.1 | 622.7 | Yes | Yes | Yes | Yes | Primary marker
NAM | 1127 | 1127.0 ± 0.0 | 15/15 | 98 | 100 | 32.4 | 5.2 | 519.3 | No | 93.7 | 673 | Yes | Yes | Yes | Yes | Primary marker
NAM | 458 | 457.0 ± 0.4 | 15/15 | 97 | 100 | 30.5 | 5.8 | 560.3 | No | 94.5 | 57 | Yes | Yes | No | Yes | Supporting peak
NAM | 487 | 487.0 ± 0.0 | 15/15 | 87 | 100 | 22.2 | 8.3 | 394.5 | No | 94.9 | 87 | No | Yes | Yes | Yes | Supporting peak
NAM | 1249 | 1248.7 ± 0.5 | 15/15 | 72 | 100 | 16.2 | 3.7 | 318 | No | 84.6 | 551.3 | No | Yes | No | No | Supporting peak
NAM | 584 | 583.8 ± 0.6 | 14/15 | 70 | 50 | 15.1 | 16.3 | 240.6 | No | 86 | 183.8 | No | No | No | No | Uncertain
NAM | 1059 | 1058.9 ± 0.3 | 15/15 | 0 | 0 | 15.8 | 6.4 | 150.1 | No | 89.1 | 658.9 | No | Yes | No | No | Uncertain
NAM | 1155 | 1155.0 ± 0.0 | 15/15 | 52 | 90 | 20.4 | 5.8 | 265.7 | No | 87 | 645 | No | Yes | No | No | Uncertain
NAM | 1589 | 1589.5 ± 0.5 | 15/15 | 66 | 100 | 14.7 | 8.9 | 268.1 | No | 75.4 | 210.5 | No | No | No | No | Supporting peak
NAM | 1346 | 1344.9 ± 1.3 | 15/15 | 74 | 100 | 35.2 | 17.1 | 609.1 | Yes | 24.4 | 455.1 | No | No | No | No | Uncertain
NAM | 1393 | 1392.5 ± 0.5 | 15/15 | 84 | 100 | 34.1 | 18.3 | 363.6 | Yes | 23.4 | 407.5 | No | No | No | No | Uncertain
NAM | 1451 | 1450.0 ± 0.5 | 15/15 | 86 | 100 | 27.9 | 10 | 307.5 | Yes | 68.8 | 350 | No | Yes | No | No | Uncertain
LTA | 802 | 802.2 ± 5.1 | 4/4 | 88.9 |  | 21.4 | 12.6 | 891 | No | 93.9 | 402.2 | Yes | Yes | Yes | Yes | Primary marker
LTA | 861 | 861.5 ± 1.0 | 4/4 | 88.9 |  | 17.6 | 10.8 | 408 | No | 92.8 | 461.5 | Yes | Yes | Yes | Yes | Primary marker
LTA | 1092 | 1091.8 ± 0.5 | 4/4 | 100 |  | 28.5 | 8.8 | 1179.7 | No | 94.6 | 691.8 | Yes | Yes | Yes | Yes | Primary marker
LTA | 1459 | 1458.8 ± 0.5 | 4/4 | 100 |  | 18.5 | 7.4 | 777.6 | No | 82.2 | 341.2 | Yes | Yes | No | Yes | Supporting peak
LTA | 541 | 541.2 ± 5.2 | 4/4 | 77.8 |  | 11.3 | 27.5 | 465.2 | No | 89 | 141.2 | No | No | No | No | Uncertain
LTA | 1367 | 1367.2 ± 2.2 | 1/4 | 11.1 |  | 13.2 | 6.8 | 121.3 | Yes | -13.4 | 432.8 | No | No | No | No | Glass/background

## 20_Chemical_Assignment
Sheet 20 — Chemical / functional-group interpretation of final markers (TENTATIVE assignments)
NAM (MurNAc) = 2-acetamido-3-O-[(R)-1-carboxyethyl]-2-deoxy-D-glucopyranose: pyranose ring, N-acetyl amide, 3-O-lactyl ether, carboxylic acid. LTA = poly(glycerol-phosphate) chain with D-Ala/glycosyl substituents on a diacylglycerol glycolipid anchor.
Biomarker | Category | Raman_shift | Mean_SD | Intensity | SNR | Literature | Molecular_vibration | Functional_group | Chemical_component | Biomarker_relevance | Citation | Confidence | Statement
NAM | Primary marker | 541 | 541.0 ± 0.0 | 715 ± 33 counts (rel. 0.78) | 39.9 | 400–600 (region) | Skeletal ring deformation δ(C–C–O)/δ(C–C–C) | Pyranose-ring skeleton | Glucopyranose ring of MurNAc | Sugar-ring fingerprint (shared by hexosamines; not NAM-unique alone) | L6 | Low | Tentative assignment.
NAM | Primary marker | 829 | 829.0 ± 0.0 | 625 ± 30 counts (rel. 0.68) | 34.8 | 800–950 (region) | ν(C–O–C) ring and/or δ(C1–H) (anomeric region) | Pyranose ring / anomeric centre | Ring C–O–C and anomeric C1 | Carbohydrate ring mode | L6 | Low | Tentative assignment.
NAM | Primary marker | 870.5 | 870.5 ± 0.5 | 872 ± 37 counts (rel. 0.95) | 48.6 | 800–950 (region) | ν(C–C)/ν(C–O–C) ring stretching | Pyranose ring | Glucopyranose ring | Carbohydrate ring mode | L6 | Low | Tentative assignment.
NAM | Primary marker | 929 | 929.0 ± 0.0 | 917 ± 37 counts (rel. 1.00) | 51.1 | 800–950 (region) | ν(C–O–C)/ν(C–C); lactyl contribution not verified | Ring C–O–C / C–C; possibly lactyl C–C(OOH) | Ring and/or 3-O-lactyl side chain | Possibly distinguishes MurNAc (lactyl ether) from GlcNAc – unverified | L6, L5 | Low | Tentative assignment.
NAM | Primary marker | 955 | 955.0 ± 0.0 | 862 ± 44 counts (rel. 0.94) | 48.1 | ~950–960 (edge of 800–950 region) | ν(C–C) and/or CH3 rocking (N-acetyl or lactyl methyl) | C–C skeleton / CH3 | Ring C–C; N-acetyl / lactyl CH3 | General; tentative | L5, L6 | Low | Tentative assignment.
NAM | Primary marker | 1023 | 1022.7 ± 0.5 | 749 ± 47 counts (rel. 0.82) | 41.8 | 1000–1200 (region) | ν(C–O)/ν(C–C), C–OH stretching | C–O (C–OH) | Ring hydroxyl-bearing carbons (C4, C6) | Carbohydrate C–O band | L6 | Low | Tentative assignment.
NAM | Primary marker | 1127 | 1127.0 ± 0.0 | 581 ± 30 counts (rel. 0.63) | 32.4 | 1000–1200 (region) | ν(C–O–C) (ring / 3-O-lactyl ether) or ν(C–C) | C–O–C ether / C–C | Ring oxygen and ether linkage | Carbohydrate C–O–C band | L6 | Low | Tentative assignment.
LTA | Primary marker | 802 | 802.2 ± 5.1 | 834 ± 105 counts (rel. 0.75) | 21.4 | ~813 | ν(O–P–O) symmetric stretch | Phosphodiester | Poly(glycerol-phosphate) backbone | Backbone-specific for teichoic-acid-type polymers (with 1092) | L5 | Low | Tentative assignment.
LTA | Primary marker | 861.5 | 861.5 ± 1.0 | 687 ± 74 counts (rel. 0.62) | 17.6 | 800–950 (region) | ν(C–C)/ν(C–O–C) | Polyol / glycosyl C–C, C–O–C | Glycerol units and sugar substituents | General | L6 | Low | Tentative assignment.
LTA | Primary marker | 1092 | 1091.8 ± 0.5 | 1110 ± 98 counts (rel. 1.00) | 28.5 | 1090–1095 | ν_s(PO2−) symmetric stretch | Phosphodiester (PO2−) | Poly(glycerol-phosphate) backbone | Strongest LTA band; phosphodiester backbone | L5, L7 (+L8 IR, whole cells) | Moderate | Tentative assignment.
LTA | Supporting peak | 1459 | 1458.8 ± 0.5 | 721 ± 53 counts (rel. 0.65) | 18.5 | 1440–1460 | δ(CH2) scissoring | CH2 (acyl chains, glycerol) | Glycolipid anchor fatty-acyl chains; glycerol CH2 | Non-specific (all lipids / CH2-rich compounds) | L7 | Low | Tentative assignment.

## 21_Background_Check
Sheet 21 — Background-origin screen for each final marker (glass, instrument, fluorescence, cosmic ray, contamination, processing)
biomarker | peak | glass | instrument | fluorescence | cosmic | contamination | processing | verdict
NAM | 541 | glass blank at 8 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
NAM | 830 | glass blank at 8 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
NAM | 870 | glass blank at 9 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
NAM | 929 | glass blank at 9 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
NAM | 955 | glass blank at 8 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
NAM | 1023 | glass blank at 7 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602); nearest PS band 8 cm-1 away | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
NAM | 1127 | glass blank at 7 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
LTA | 802 | glass blank at 9 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
LTA | 861 | glass blank at 8 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
LTA | 1092 | glass blank at 10 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background
LTA | 1459 | glass blank at 22 % of its max | not in empty-slide blank; not a fixed-pixel feature (same position across powers/days) | broad fluorescence removed by ALS; band width ≈ 10–50 cm-1 is Raman-like, not fluorescence | present in 100 % of spectra (cosmic spikes are random, single-spectrum) | no polystyrene pattern (620/1001/1031/1602) | present in 100 % of unsmoothed (despiked) spectra → not created by SG smoothing | Not background

## 22_QC_Extended
Sheet 22 — Post-processing QC for every bulk/blank/suspect spectrum (pipeline v3)
Acceptance: not saturated in 400–1800 (16-bit detector limit, physical); SNR ≥ 10 (project convention); r to class mean ≥ 0.85 with glass zone excluded (project convention); no contamination pattern; no distortion (negative counts / abnormal dark). Spikes are corrected, not a rejection criterion.
Spectrum | Class | Power (%) | Int (s) | Acc | Spot | SNR | Noise (counts, robust) | Max dark-sub. (400–1800) | Min dark-sub. (400–1800) | Max raw (400–1800) | Max raw (full) | Saturated laser line only | Saturated in 400–1800 | Baseline mean | Baseline range | Baseline slope (counts/100 cm-1) | Dynamic range (max/noise) | Strongest-marker prominence | Spikes removed | Correlation with group mean | QC | Reason | In reference
Glass_Blank_20260529_P005pct_T05s_A01_S01.csv | GLASS | 5 | 5 | 1 | not stated | 54.63 | 29.6 | 6548 | 2118 | 7681 | 7681 | No | No | 3113.5 | 2907.55 | 165.83 | 54.68 | 1660.93 | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
Glass_Blank_20260529_P005pct_T05s_A01_S02.csv | GLASS | 5 | 5 | 1 | not stated | 56.74 | 30.25 | 6695 | 2127 | 7824 | 7824 | No | No | 3141.42 | 2907.05 | 167.51 | 56.82 | 1783.63 | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
Glass_Blank_20260529_P005pct_T05s_A01_S03.csv | GLASS | 5 | 5 | 1 | not stated | 50.63 | 32.66 | 6613 | 2092 | 7710 | 7710 | No | No | 3113.65 | 2868.66 | 166.2 | 50.67 | 1712.37 | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
Glass_Blank_20260529_P005pct_T05s_A01_S04.csv | GLASS | 5 | 5 | 1 | not stated | 56.64 | 30.22 | 6831 | 2193 | 7934 | 7934 | No | No | 3214.52 | 2957.61 | 171.84 | 56.68 | 1755.5 | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
Glass_Blank_20260529_P005pct_T05s_A01_S05.csv | GLASS | 5 | 5 | 1 | not stated | 56.46 | 30.98 | 6854 | 2160 | 7979 | 7979 | No | No | 3204.93 | 2946.58 | 171.42 | 56.55 | 1790.74 | 0 |  | Valid | Empty glass slide background (5% / 5 s / 1 acc) - not matched to sample settings. | No
LPS_Bulk_20260413_P010pct_T10s_A02_S01.csv | LPS | 10 | 10 | 2 | not stated | 9.3 | 18.45 | 1776.5 | 306 | 2985 | 4299.5 | No | No | 887.96 | 1289.94 | -88.75 | 9.37 | 136.63 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T10s_A02_S02.csv | LPS | 10 | 10 | 2 | not stated | 4.91 | 17.81 | 428 | 128.5 | 2078 | 4406 | No | No | 275.34 | 208.81 | -8.75 | 4.92 | 94.75 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T10s_A02_S03.csv | LPS | 10 | 10 | 2 | not stated | 4.21 | 19.38 | 430.5 | 105 | 2056 | 4521 | No | No | 257.88 | 187.97 | -6.33 | 4.24 | 89.6 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T10s_A05_S01.csv | LPS | 10 | 10 | 5 | not stated | 9.32 | 14.28 | 634.4 | 225.6 | 3364.6 | 5043 | No | No | 371.84 | 252.09 | 0.77 | 9.36 | 142.57 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T10s_A05_S02.csv | LPS | 10 | 10 | 5 | not stated | 4.34 | 14.34 | 440 | 149.6 | 3277.4 | 4839.2 | No | No | 281.84 | 214.38 | -10.36 | 4.38 | 70.14 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P010pct_T60s_A30_S01.csv | LPS | 10 | 60 | 30 | not stated | 14.75 | 22.24 | 1499.63 | -21.5 | 11756.5 | 21942.7 | No | No | 475.99 | 456.55 | 21.66 | 14.77 | 344.84 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T10s_A02_S01.csv | LPS | 20 | 10 | 2 | not stated | 4.66 | 17.22 | 556.8 | 205.1 | 3393 | 6072.5 | No | No | 378.15 | 262.97 | -10.89 | 4.71 | 88.83 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T10s_A05_S01.csv | LPS | 20 | 10 | 5 | not stated | 5.73 | 15.53 | 639.6 | 234.4 | 3446 | 6535.8 | No | No | 422.29 | 295.36 | -13.65 | 5.78 | 103.84 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T30s_A50_S01.csv | LPS | 20 | 30 | 50 | not stated | 72.82 | 4.62 | 1511.2 | 119.82 | 4193.54 | 17087.98 | No | No | 882.84 | 566.22 | 8.71 | 72.85 | 346.72 | 1 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T30s_A60_S01.csv | LPS | 20 | 30 | 60 | not stated | 76.75 | 4.3 | 1522.37 | 648.68 | 4909.35 | 17163.1 | No | No | 912.07 | 569 | 3.68 | 76.76 | 341.67 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T60s_A30_S01.csv | LPS | 20 | 60 | 30 | not stated | 29.12 | 17.38 | 2185.47 | 441.43 | 12219.43 | 29874.07 | No | No | 1271.45 | 808.9 | 2.12 | 29.12 | 528.86 | 0 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T60s_A30_S02.csv | LPS | 20 | 60 | 30 | not stated | 31.51 | 16.87 | 3634.67 | 587.47 | 12976.23 | 29108.07 | No | No | 1329.2 | 772.65 | 10.18 | 31.52 | 560.19 | 1 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T60s_A50_S01.csv | LPS | 20 | 60 | 50 | not stated | 19.8 | 28.46 | 2075.22 | -1497.26 | 10280.74 | 31704.56 | No | No | 311.76 | 887.3 | 60.85 | 19.81 | 583.8 | 4 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LPS_Bulk_20260413_P020pct_T60s_A50_S02.csv | LPS | 20 | 60 | 50 | not stated | 20.73 | 45.7 | 4836.24 | 1529.24 | 10605.9 | 30282.72 | No | No | 3071.18 | 2045.1 | -79.54 | 20.73 | 944.58 | 2 |  | Signal not established | Valid acquisition, but no reproducible LPS band resolved; dominant ~1360 cm-1 feature matches glass-slide background. Not used as a reference. | No
LTA_Bulk_20260403_P030pct_T05s_A10_S01.csv | LTA | 30 | 5 | 10 | not stated | 6.57 | 8.19 | 251.5 | 68.7 | 1902.1 | 3119 | No | No | 159.93 | 151.83 | -8.3 | 6.66 | 58 | 0 | 0.87 | Low SNR | SNR 6.7 < 10 at low power / short integration; genuine LTA spectrum used only for parameter comparison. | No
LTA_Bulk_20260403_P040pct_T10s_A10_S01.csv | LTA | 40 | 10 | 10 | different spot | 11.28 | 8.8 | 582.8 | 196.1 | 2797.1 | 5842.5 | No | No | 379.67 | 366.75 | -20.65 | 11.34 | 105.3 | 0 | 0.94 | Valid |  | No
LTA_Bulk_20260403_P050pct_T05s_A10_S01.csv | LTA | 50 | 5 | 10 | not stated | 8.81 | 7.93 | 340.9 | 105 | 1969.9 | 4454.3 | No | No | 216.7 | 200.77 | -11.01 | 8.82 | 76.45 | 0 | 0.89 | Low SNR | SNR 8.4 < 10 at low power / short integration; genuine LTA spectrum used only for parameter comparison. | No
LTA_Bulk_20260403_P050pct_T10s_A10_S01.csv | LTA | 50 | 10 | 10 | different spot | 15.5 | 8.68 | 874.7 | 280.7 | 2999.9 | 8096.3 | No | No | 564.61 | 560.31 | -33.61 | 15.5 | 142.22 | 0 | 0.96 | Valid |  | No
LTA_Bulk_20260403_P050pct_T30s_A15_S01.csv | LTA | 50 | 30 | 15 | not stated | 13.7 | 55.05 | 4192.6 | 988.33 | 9292 | 20222.67 | No | No | 2597.71 | 2395.07 | -135.1 | 18.99 | 889.33 | 0 | 0.49 | Abnormal - review | Distorted spectrum (r = 0.45 vs LTA class mean, glass region excluded) with low SNR; not explained by power/time. Not used as reference. | No
LTA_Bulk_20260403_P060pct_T10s_A10_S01.csv | LTA | 60 | 10 | 10 | different spot | 17.27 | 8.92 | 903.7 | 303.8 | 3054.9 | 8643.6 | No | No | 594.5 | 570.44 | -33.34 | 17.36 | 160.63 | 0 | 0.95 | Valid |  | No
LTA_Bulk_20260403_P060pct_T45s_A02_S01.csv | LTA | 60 | 45 | 2 | different spot | 23.38 | 31.44 | 5979.5 | 1818.5 | 10105 | 34145 | No | No | 3758.91 | 4019.82 | -250.48 | 23.39 | 760.72 | 0 | 0.98 | Valid |  | No
LTA_Bulk_20260403_P070pct_T10s_A10_S01.csv | LTA | 70 | 10 | 10 | different spot | 21.58 | 8.44 | 987.8 | 344.3 | 3150 | 9892.5 | No | No | 667.71 | 609.82 | -33.94 | 21.61 | 194.05 | 0 | 0.97 | Valid |  | No
LTA_Bulk_20260403_P080pct_T45s_A02_S01.csv | LTA | 80 | 45 | 2 | different spot | 29.4 | 38.97 | 20457 | 5087 | 22619 | 47406 | No | No | 11895.19 | 15071.53 | -1026.19 | 29.52 | 1218.46 | 0 | 0.97 | Valid |  | Yes
LTA_Bulk_20260403_P080pct_T45s_A02_S02.csv | LTA | 80 | 45 | 2 | different spot | 27.97 | 34.68 | 8298 | 2553 | 11787 | 43075 | No | No | 5246.62 | 5620.15 | -353.41 | 27.98 | 1023.74 | 0 | 0.98 | Valid |  | Yes
LTA_Bulk_20260403_P090pct_T45s_A02_S01.csv | LTA | 90 | 45 | 2 | different spot | 27.38 | 41.19 | 13706 | 3866 | 16295.5 | 48531.5 | No | No | 8350.32 | 9675.14 | -638.24 | 27.44 | 1207.82 | 0 | 0.99 | Valid |  | Yes
LTA_Bulk_20260403_P095pct_T45s_A02_S01.csv | LTA | 95 | 45 | 2 | different spot | 29.34 | 40.79 | 12274.5 | 3583.5 | 14848 | 53846 | No | No | 7732 | 8425.95 | -537.8 | 29.39 | 1268.63 | 0 | 0.98 | Valid |  | Yes
NAM_Bulk_20260128_P005pct_T05s_A02_S01.csv | NAM | 5 | 5 | 2 | not stated | 2.54 | 13.58 | 109 | -14 | 1363.5 | 1808 | No | No | 31.25 | 42.47 | -3.09 | 2.55 | 40.45 | 0 | 0.52 | Low SNR | SNR 2.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T10s_A02_S01.csv | NAM | 5 | 10 | 2 | not stated | 3.58 | 13.42 | 148.5 | -4.5 | 1782 | 2896.5 | No | No | 61.04 | 62.21 | -4.67 | 3.64 | 53.36 | 0 | 0.67 | Low SNR | SNR 3.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T15s_A02_S01.csv | NAM | 5 | 15 | 2 | not stated | 5.07 | 15.84 | 445.5 | 67 | 3450.5 | 3698.5 | No | No | 233.62 | 318.44 | -20.59 | 5.11 | 90.55 | 0 | 0.71 | Low SNR | SNR 5.0 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T15s_A02_S02.csv | NAM | 5 | 15 | 2 | not stated | 5.25 | 14.01 | 193.5 | 15.5 | 2615.5 | 3583.5 | No | No | 85.18 | 95.18 | -5.97 | 5.27 | 80.47 | 0 | 0.82 | Low SNR | SNR 5.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T20s_A02_S01.csv | NAM | 5 | 20 | 2 | not stated | 5 | 14.86 | 219.5 | 16 | 2628.5 | 4606 | No | No | 96.62 | 94.06 | -6.46 | 5.01 | 81.57 | 0 | 0.84 | Low SNR | SNR 4.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P005pct_T25s_A02_S01.csv | NAM | 5 | 25 | 2 | not stated | 6.82 | 16.63 | 270.5 | 24.5 | 3058 | 5426 | No | No | 112.78 | 100.81 | -5.97 | 6.84 | 122.1 | 0 | 0.82 | Low SNR | SNR 6.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T05s_A02_S01.csv | NAM | 10 | 5 | 2 | not stated | 3.43 | 13.83 | 126 | -8 | 1387 | 2131.5 | No | No | 43.62 | 56.46 | -3.49 | 3.46 | 55.16 | 0 | 0.58 | Low SNR | SNR 3.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T10s_A02_S01.csv | NAM | 10 | 10 | 2 | not stated | 4.64 | 12.95 | 180.5 | 9 | 1799 | 3332 | No | No | 78.72 | 87.22 | -5.83 | 4.76 | 65.67 | 0 | 0.77 | Low SNR | SNR 4.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T15s_A02_S01.csv | NAM | 10 | 15 | 2 | not stated | 6.32 | 16.99 | 519 | 74 | 3461 | 4679.5 | No | No | 248.81 | 366.39 | -24.3 | 6.33 | 118.85 | 0 | 0.8 | Low SNR | SNR 6.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T15s_A02_S02.csv | NAM | 10 | 15 | 2 | not stated | 5.88 | 17.87 | 546 | 88 | 3427.5 | 4519.5 | No | No | 255.88 | 325.24 | -22.3 | 5.89 | 113.33 | 0 | 0.88 | Low SNR | SNR 6.0 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T15s_A02_S03.csv | NAM | 10 | 15 | 2 | not stated | 6.68 | 14.2 | 245 | 7.5 | 2613 | 4739.5 | No | No | 112.09 | 109.71 | -7.45 | 6.69 | 102.94 | 0 | 0.88 | Low SNR | SNR 6.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T20s_A02_S01.csv | NAM | 10 | 20 | 2 | not stated | 6.25 | 15.89 | 650.5 | 41 | 3173.5 | 5726 | No | No | 129.66 | 121.8 | -8.69 | 6.27 | 108.76 | 0 | 0.9 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P010pct_T25s_A02_S01.csv | NAM | 10 | 25 | 2 | not stated | 7.76 | 17.33 | 371.5 | 55 | 3098.5 | 7096.5 | No | No | 163.49 | 151.44 | -9.92 | 7.85 | 142.04 | 0 | 0.93 | Low SNR | SNR 7.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T05s_A02_S01.csv | NAM | 15 | 5 | 2 | not stated | 2.81 | 14.03 | 146 | 0 | 1397.5 | 2505 | No | No | 60.76 | 80.47 | -5.3 | 2.84 | 42.21 | 0 | 0.61 | Low SNR | SNR 2.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T10s_A02_S01.csv | NAM | 15 | 10 | 2 | not stated | 5.83 | 13.87 | 318 | 39 | 1833.5 | 4248.5 | No | No | 111.7 | 108.6 | -7.78 | 5.84 | 87.1 | 0 | 0.87 | Low SNR | SNR 5.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T15s_A02_S01.csv | NAM | 15 | 15 | 2 | not stated | 7.74 | 14.83 | 344 | 56.5 | 2713.5 | 5977 | No | No | 158.75 | 151.75 | -10.34 | 7.75 | 121.99 | 0 | 0.93 | Low SNR | SNR 7.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T20s_A02_S01.csv | NAM | 15 | 20 | 2 | not stated | 9.26 | 15.21 | 777.5 | 72.5 | 3300.5 | 7541 | No | No | 191.02 | 174.7 | -11.98 | 9.34 | 154.06 | 0 | 0.95 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P015pct_T25s_A02_S01.csv | NAM | 15 | 25 | 2 | not stated | 11.15 | 17.63 | 522 | 88 | 3165 | 9128 | No | No | 231.01 | 211.98 | -13.64 | 11.31 | 207.46 | 0 | 0.95 | Valid |  | No
NAM_Bulk_20260128_P020pct_T05s_A02_S01.csv | NAM | 20 | 5 | 2 | not stated | 4.54 | 14.74 | 191.5 | 12.5 | 1383 | 2890.5 | No | No | 78.99 | 95.57 | -6.2 | 4.54 | 74.34 | 0 | 0.67 | Low SNR | SNR 4.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T10s_A02_S01.csv | NAM | 20 | 10 | 2 | not stated | 6.49 | 13.99 | 445 | 38.5 | 1868.5 | 4923 | No | No | 137.73 | 145.75 | -9.77 | 6.51 | 97.21 | 0 | 0.88 | Low SNR | SNR 6.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T10s_A02_S02.csv | NAM | 20 | 10 | 2 | not stated | 6.33 | 14.99 | 456.5 | 32.5 | 1884 | 4946.5 | No | No | 141.72 | 146.8 | -10.27 | 6.35 | 103.36 | 0 | 0.91 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T15s_A02_S01.csv | NAM | 20 | 15 | 2 | not stated | 9.66 | 17.65 | 760.5 | 146 | 3598.5 | 6796.5 | No | No | 394.41 | 509.34 | -34.48 | 9.71 | 184.51 | 0 | 0.92 | Low SNR | SNR 9.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T15s_A02_S02.csv | NAM | 20 | 15 | 2 | not stated | 9.96 | 14.48 | 423.5 | 72.5 | 2745 | 7166.5 | No | No | 201.57 | 188.94 | -12.97 | 9.97 | 154.64 | 0 | 0.94 | Low SNR | SNR 9.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P020pct_T20s_A02_S01.csv | NAM | 20 | 20 | 2 | not stated | 11.29 | 15.39 | 808.5 | 98.5 | 3331.5 | 8936 | No | No | 233.73 | 227.72 | -14.69 | 11.38 | 179.15 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P020pct_T25s_A02_S01.csv | NAM | 20 | 25 | 2 | not stated | 14.79 | 15.77 | 1069.5 | 117 | 3241 | 11269.5 | No | No | 279.5 | 252.7 | -16.06 | 14.95 | 244.89 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P025pct_T05s_A02_S01.csv | NAM | 25 | 5 | 2 | not stated | 4.58 | 13.44 | 198.5 | 11.5 | 1408 | 3273 | No | No | 88.51 | 101.16 | -6.95 | 4.58 | 68.1 | 0 | 0.77 | Low SNR | SNR 4.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P025pct_T10s_A02_S01.csv | NAM | 25 | 10 | 2 | not stated | 8.8 | 14.07 | 481.5 | 44 | 1913.5 | 5634 | No | No | 164.04 | 170.29 | -11.23 | 8.85 | 130.69 | 0 | 0.91 | Low SNR | SNR 8.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P025pct_T15s_A02_S01.csv | NAM | 25 | 15 | 2 | not stated | 9.51 | 17.8 | 899.5 | 179 | 3612 | 8157.5 | No | No | 474.45 | 581.63 | -39.01 | 9.53 | 182.64 | 0 | 0.93 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P025pct_T15s_A02_S02.csv | NAM | 25 | 15 | 2 | not stated | 11.08 | 15.2 | 461.5 | 87.5 | 2774 | 8174 | No | No | 223.21 | 218.39 | -14.76 | 11.12 | 179.99 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P025pct_T20s_A02_S01.csv | NAM | 25 | 20 | 2 | not stated | 11.56 | 17.34 | 822 | 99.5 | 3345 | 10047 | No | No | 257.74 | 245.53 | -16.36 | 11.67 | 213.48 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P025pct_T25s_A02_S01.csv | NAM | 25 | 25 | 2 | not stated | 15.77 | 18.31 | 1068 | 138 | 3220.5 | 12787 | No | No | 314.84 | 294.04 | -18.54 | 15.93 | 306.2 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P030pct_T05s_A02_S01.csv | NAM | 30 | 5 | 2 | not stated | 5.24 | 13.5 | 220 | 11.5 | 1441.5 | 3542.5 | No | No | 99.69 | 121.02 | -8.06 | 5.27 | 77.8 | 0 | 0.76 | Low SNR | SNR 5.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P030pct_T10s_A02_S01.csv | NAM | 30 | 10 | 2 | not stated | 8.4 | 14.89 | 463 | 72 | 1934 | 6547 | No | No | 189.57 | 191.66 | -13.32 | 8.43 | 137.75 | 0 | 0.93 | Low SNR | SNR 8.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P030pct_T15s_A02_S01.csv | NAM | 30 | 15 | 2 | not stated | 12.14 | 17.37 | 996 | 216 | 3638 | 9484 | No | No | 514.5 | 626.73 | -43.07 | 12.24 | 221.83 | 0 | 0.95 | Valid |  | No
NAM_Bulk_20260128_P030pct_T15s_A02_S02.csv | NAM | 30 | 15 | 2 | not stated | 11.37 | 17.54 | 894.5 | 154.5 | 3557.5 | 9304 | No | No | 427.39 | 539.18 | -37.48 | 11.38 | 216.77 | 0 | 0.93 | Valid |  | No
NAM_Bulk_20260128_P030pct_T15s_A02_S03.csv | NAM | 30 | 15 | 2 | not stated | 12.74 | 14.74 | 544.5 | 110 | 2814 | 9024.5 | No | No | 262.83 | 254.26 | -17.15 | 12.78 | 200.37 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P030pct_T20s_A02_S01.csv | NAM | 30 | 20 | 2 | not stated | 13.08 | 17.77 | 908.5 | 137 | 3431.5 | 11562.5 | No | No | 310.2 | 298.03 | -19.49 | 13.25 | 242.76 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P030pct_T25s_A02_S01.csv | NAM | 30 | 25 | 2 | not stated | 17.59 | 18.01 | 1140.5 | 167.5 | 3352 | 13856 | No | No | 374.8 | 323.29 | -21.55 | 17.71 | 331.6 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P035pct_T05s_A02_S01.csv | NAM | 35 | 5 | 2 | not stated | 6.32 | 12.22 | 260.5 | 18.5 | 1430 | 3907.5 | No | No | 116.76 | 138.27 | -9.24 | 6.34 | 84.49 | 0 | 0.83 | Low SNR | SNR 5.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P035pct_T10s_A02_S01.csv | NAM | 35 | 10 | 2 | not stated | 10.04 | 15.22 | 500.5 | 76.5 | 1947 | 6950 | No | No | 201.63 | 201.84 | -13.58 | 10.08 | 159.3 | 0 | 0.95 | Valid |  | No
NAM_Bulk_20260128_P035pct_T15s_A02_S01.csv | NAM | 35 | 15 | 2 | not stated | 11.1 | 18.69 | 824.5 | 129 | 3466.5 | 10041 | No | No | 370.76 | 501.65 | -35.04 | 11.11 | 221.4 | 0 | 0.89 | Valid |  | No
NAM_Bulk_20260128_P035pct_T15s_A02_S02.csv | NAM | 35 | 15 | 2 | not stated | 14.79 | 14.31 | 588 | 120.5 | 2841.5 | 10229 | No | No | 286.3 | 278.86 | -18.48 | 14.84 | 221.08 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P035pct_T20s_A02_S01.csv | NAM | 35 | 20 | 2 | not stated | 13.7 | 18.45 | 925 | 153 | 3448 | 12732 | No | No | 346.89 | 311.46 | -21.36 | 13.84 | 263.28 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P035pct_T25s_A02_S01.csv | NAM | 35 | 25 | 2 | not stated | 18.27 | 18.83 | 1173 | 195.5 | 3374 | 16323 | No | No | 416.91 | 356.84 | -23.97 | 18.41 | 361.84 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P040pct_T05s_A02_S01.csv | NAM | 40 | 5 | 2 | not stated | 6.34 | 13.29 | 265.5 | 8.5 | 1454.5 | 4158 | No | No | 129.21 | 155.7 | -9.85 | 6.4 | 90.9 | 0 | 0.86 | Low SNR | SNR 6.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P040pct_T10s_A02_S01.csv | NAM | 40 | 10 | 2 | not stated | 10.04 | 14.61 | 515.5 | 96.5 | 1976.5 | 7227.5 | No | No | 230.97 | 231.62 | -16 | 10.12 | 157.59 | 0 | 0.94 | Valid |  | No
NAM_Bulk_20260128_P040pct_T15s_A02_S01.csv | NAM | 40 | 15 | 2 | not stated | 12.52 | 19.72 | 991 | 125.5 | 3499.5 | 11643.5 | No | No | 392.53 | 530.71 | -37.04 | 12.53 | 260.52 | 0 | 0.88 | Valid |  | No
NAM_Bulk_20260128_P040pct_T15s_A02_S02.csv | NAM | 40 | 15 | 2 | not stated | 12.65 | 19.53 | 924 | 105.5 | 3487.5 | 11266.5 | No | No | 364.9 | 496.27 | -34.52 | 12.72 | 261.91 | 0 | 0.88 | Valid |  | No
NAM_Bulk_20260128_P040pct_T15s_A02_S03.csv | NAM | 40 | 15 | 2 | not stated | 14.79 | 15.54 | 690 | 148 | 2887.5 | 11825.5 | No | No | 330.96 | 318.93 | -20.99 | 14.9 | 245.33 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P040pct_T20s_A02_S01.csv | NAM | 40 | 20 | 2 | not stated | 15.17 | 19.57 | 977 | 183.5 | 3500 | 14392.5 | No | No | 407.59 | 363.37 | -24.8 | 15.25 | 308.5 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P040pct_T25s_A02_S01.csv | NAM | 40 | 25 | 2 | not stated | 20.55 | 18.55 | 1241 | 222.5 | 3421 | 18762 | No | No | 480.84 | 397.79 | -27.29 | 20.71 | 392.51 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P045pct_T05s_A02_S01.csv | NAM | 45 | 5 | 2 | not stated | 6.3 | 13.74 | 298 | 24.5 | 1484.5 | 4561 | No | No | 139.96 | 156.38 | -10.49 | 6.35 | 91.43 | 0 | 0.88 | Low SNR | SNR 6.3 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P045pct_T10s_A02_S01.csv | NAM | 45 | 10 | 2 | not stated | 11.55 | 14.86 | 546 | 87.5 | 1997.5 | 8021 | No | No | 253.01 | 246.63 | -16.78 | 11.59 | 182.54 | 0 | 0.95 | Valid |  | No
NAM_Bulk_20260128_P045pct_T15s_A02_S01.csv | NAM | 45 | 15 | 2 | not stated | 13.06 | 20.24 | 910.5 | 133.5 | 3454 | 12142.5 | No | No | 363.32 | 498.07 | -34.58 | 13.15 | 276.49 | 0 | 0.88 | Valid |  | No
NAM_Bulk_20260128_P045pct_T15s_A02_S02.csv | NAM | 45 | 15 | 2 | not stated | 16.9 | 15.11 | 763.5 | 167 | 2924 | 12868.5 | No | No | 359.65 | 332.87 | -22.74 | 16.95 | 264.89 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P045pct_T20s_A02_S01.csv | NAM | 45 | 20 | 2 | not stated | 19.91 | 16.02 | 1031 | 195.5 | 3554 | 15522.5 | No | No | 444.51 | 392.05 | -27.03 | 20.08 | 339.54 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P045pct_T25s_A02_S01.csv | NAM | 45 | 25 | 2 | not stated | 22.48 | 19.59 | 1240 | 255 | 3484.5 | 20572 | No | No | 517.84 | 454.34 | -29.84 | 22.61 | 459.38 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P050pct_T05s_A02_S01.csv | NAM | 50 | 5 | 2 | not stated | 7.03 | 14.16 | 312.5 | 45.5 | 1508.5 | 4879.5 | No | No | 152.36 | 164.54 | -11.39 | 7.08 | 107.29 | 0 | 0.91 | Low SNR | SNR 7.4 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P050pct_T10s_A02_S01.csv | NAM | 50 | 10 | 2 | not stated | 10.66 | 16.88 | 565.5 | 111 | 2017 | 8373 | No | No | 268.75 | 266.35 | -18.06 | 10.71 | 193.68 | 0 | 0.95 | Valid |  | No
NAM_Bulk_20260128_P050pct_T15s_A02_S01.csv | NAM | 50 | 15 | 2 | not stated | 13.09 | 21.37 | 912 | 105 | 3458.5 | 13073 | No | No | 326.72 | 497.32 | -33.84 | 13.12 | 296.97 | 0 | 0.82 | Review - higher noise | Correct NAM band pattern but r = 0.84 vs class mean; extra noise structure. | No
NAM_Bulk_20260128_P050pct_T15s_A02_S02.csv | NAM | 50 | 15 | 2 | not stated | 16.12 | 16.07 | 769 | 162 | 2938 | 13233.5 | No | No | 369.39 | 347.61 | -23.43 | 16.25 | 268.16 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P050pct_T20s_A02_S01.csv | NAM | 50 | 20 | 2 | not stated | 20.34 | 17.36 | 1022.5 | 203 | 3545.5 | 15630 | No | No | 447.33 | 398.26 | -26.77 | 20.47 | 363 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P050pct_T25s_A02_S01.csv | NAM | 50 | 25 | 2 | not stated | 22.48 | 19.79 | 1270.5 | 261.5 | 3509 | 21820 | No | No | 541.43 | 461.23 | -30.89 | 22.65 | 462.9 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P055pct_T05s_A02_S01.csv | NAM | 55 | 5 | 2 | not stated | 7.43 | 14.1 | 347 | 48 | 1475.5 | 5236 | No | No | 163.26 | 180.44 | -12.63 | 7.53 | 111.44 | 0 | 0.92 | Low SNR | SNR 7.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P055pct_T10s_A02_S01.csv | NAM | 55 | 10 | 2 | not stated | 12.37 | 15.71 | 628 | 123 | 2309.5 | 9363 | No | No | 297.04 | 278.19 | -19.9 | 12.48 | 207.24 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P055pct_T15s_A02_S01.csv | NAM | 55 | 15 | 2 | not stated | 18.64 | 16.7 | 1184.5 | 259.5 | 3384.5 | 14103 | No | No | 581.79 | 626.32 | -43.09 | 18.79 | 323.39 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P055pct_T15s_A02_S02.csv | NAM | 55 | 15 | 2 | not stated | 18.02 | 17.3 | 915.5 | 203 | 3007.5 | 14549 | No | No | 432.61 | 397 | -27.36 | 18.06 | 326.08 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P055pct_T20s_A02_S01.csv | NAM | 55 | 20 | 2 | not stated | 20.23 | 19.06 | 1095 | 246.5 | 3618 | 17534 | No | No | 507.77 | 445.6 | -30.3 | 20.35 | 397.49 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P055pct_T25s_A02_S01.csv | NAM | 55 | 25 | 2 | not stated | 26.29 | 20.59 | 1385.5 | 303 | 3628 | 26457.5 | No | No | 629.79 | 537.91 | -35.52 | 26.45 | 561.92 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P060pct_T05s_A02_S01.csv | NAM | 60 | 5 | 2 | not stated | 8.06 | 13.62 | 386.5 | 48 | 1506 | 5553.5 | No | No | 178.12 | 186.4 | -13 | 8.16 | 119.45 | 0 | 0.92 | Low SNR | SNR 8.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P060pct_T10s_A02_S01.csv | NAM | 60 | 10 | 2 | not stated | 13.58 | 15.63 | 663.5 | 134 | 2349.5 | 10106.5 | No | No | 321.2 | 318.26 | -21.27 | 13.63 | 219.11 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P060pct_T15s_A02_S01.csv | NAM | 60 | 15 | 2 | not stated | 18.66 | 18.04 | 1335 | 307.5 | 3528 | 15251 | No | No | 685.04 | 746.39 | -50.71 | 18.76 | 352.01 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P060pct_T15s_A02_S02.csv | NAM | 60 | 15 | 2 | not stated | 19.7 | 16.19 | 966.5 | 223 | 3051 | 15617.5 | No | No | 464.23 | 416.16 | -28.38 | 19.87 | 332.8 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P060pct_T20s_A02_S01.csv | NAM | 60 | 20 | 2 | not stated | 23.37 | 17.79 | 1177.5 | 247 | 3441.5 | 18572 | No | No | 544.06 | 490.23 | -32.86 | 23.42 | 433.24 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P060pct_T25s_A02_S01.csv | NAM | 60 | 25 | 2 | not stated | 26.54 | 20.82 | 1453.5 | 322.5 | 3668.5 | 27253.5 | No | No | 653.42 | 564.55 | -37.18 | 26.75 | 567.4 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P065pct_T05s_A02_S01.csv | NAM | 65 | 5 | 2 | not stated | 8.61 | 14.17 | 384 | 52 | 1509.5 | 5946.5 | No | No | 180 | 194.91 | -12.61 | 8.64 | 130.86 | 0 | 0.91 | Low SNR | SNR 8.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P065pct_T10s_A02_S01.csv | NAM | 65 | 10 | 2 | not stated | 14.21 | 15.89 | 1061 | 150 | 2360.5 | 10655.5 | No | No | 337.28 | 330.96 | -22.32 | 14.32 | 233.82 | 1 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P065pct_T15s_A02_S01.csv | NAM | 65 | 15 | 2 | not stated | 20.35 | 16.14 | 1263.5 | 274 | 3429 | 15625.5 | No | No | 611.1 | 656.43 | -46.64 | 20.4 | 347.98 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260128_P065pct_T15s_A02_S02.csv | NAM | 65 | 15 | 2 | not stated | 21.1 | 16.11 | 971.5 | 232 | 3026.5 | 16585.5 | No | No | 470.18 | 422.8 | -28.93 | 21.14 | 353.41 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P065pct_T20s_A02_S01.csv | NAM | 65 | 20 | 2 | not stated | 22.56 | 19.58 | 1233 | 276.5 | 3136 | 20981.5 | No | No | 575.64 | 498.13 | -33.84 | 22.7 | 454.03 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P065pct_T25s_A02_S01.csv | NAM | 65 | 25 | 2 | not stated | 28.3 | 20.83 | 1524.5 | 337.5 | 3659.5 | 27075.5 | No | No | 678.15 | 575.88 | -37.95 | 28.53 | 608.49 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260128_P070pct_T05s_A02_S01.csv | NAM | 70 | 5 | 2 | not stated | 8.3 | 14.3 | 402.5 | 66 | 1529.5 | 6262 | No | No | 196.36 | 207.16 | -14.25 | 8.36 | 125.32 | 0 | 0.94 | Low SNR | SNR 8.6 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P070pct_T10s_A02_S01.csv | NAM | 70 | 10 | 2 | not stated | 16.27 | 15.05 | 750.5 | 161 | 2398 | 11417.5 | No | No | 359.72 | 339.87 | -23.2 | 16.35 | 257.95 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P070pct_T15s_A02_S01.csv | NAM | 70 | 15 | 2 | not stated | 20.47 | 17.68 | 1359 | 305 | 2993.5 | 16923.5 | No | No | 676.39 | 687.56 | -47.65 | 20.51 | 379.7 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P070pct_T15s_A02_S02.csv | NAM | 70 | 15 | 2 | not stated | 21.55 | 16.74 | 1041 | 230.5 | 3061 | 17121.5 | No | No | 495.81 | 454.91 | -30.31 | 21.71 | 377.13 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P070pct_T20s_A02_S01.csv | NAM | 70 | 20 | 2 | not stated | 26.26 | 18.87 | 1362 | 299.5 | 3197.5 | 24487 | No | No | 641.75 | 563.3 | -37.76 | 26.36 | 509.63 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P070pct_T25s_A02_S01.csv | NAM | 70 | 25 | 2 | not stated | 30.99 | 20.91 | 1642.5 | 374.5 | 3758 | 31104.5 | No | No | 748.47 | 635.14 | -42.14 | 31.17 | 667.09 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P075pct_T05s_A02_S01.csv | NAM | 75 | 5 | 2 | not stated | 9.53 | 13.97 | 444.5 | 78.5 | 1555.5 | 6664 | No | No | 208.41 | 216.1 | -14.85 | 9.57 | 139.44 | 0 | 0.93 | Low SNR | SNR 9.2 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260128_P075pct_T10s_A02_S01.csv | NAM | 75 | 10 | 2 | not stated | 16.46 | 16.17 | 799 | 176.5 | 2416.5 | 12369.5 | No | No | 389.61 | 366.58 | -25.11 | 16.58 | 274.79 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260128_P075pct_T15s_A02_S01.csv | NAM | 75 | 15 | 2 | not stated | 20.46 | 18.53 | 1406 | 317 | 3023.5 | 17505.5 | No | No | 689.01 | 718.16 | -49.14 | 20.59 | 393.68 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P075pct_T15s_A02_S02.csv | NAM | 75 | 15 | 2 | not stated | 22.1 | 17.79 | 1116 | 259 | 3122 | 18758.5 | No | No | 537.64 | 486.25 | -32.67 | 22.24 | 405.99 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P075pct_T20s_A02_S01.csv | NAM | 75 | 20 | 2 | not stated | 28.4 | 19.28 | 1502 | 336 | 3307 | 26619.5 | No | No | 694.03 | 599.24 | -40.34 | 28.56 | 562.36 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P075pct_T25s_A02_S01.csv | NAM | 75 | 25 | 2 | not stated | 31.76 | 21.55 | 1784.5 | 378 | 3820 | 33551.5 | No | No | 803.41 | 688.15 | -45 | 32.01 | 707.77 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P080pct_T05s_A02_S01.csv | NAM | 80 | 5 | 2 | not stated | 10.56 | 12.69 | 484 | 83 | 1572.5 | 6996.5 | No | No | 230.65 | 234.7 | -16.64 | 10.58 | 142.73 | 0 | 0.95 | Valid |  | No
NAM_Bulk_20260128_P080pct_T10s_A02_S01.csv | NAM | 80 | 10 | 2 | not stated | 18.12 | 15.1 | 848 | 185.5 | 2430 | 13500 | No | No | 408.4 | 384.77 | -26.23 | 18.16 | 280.32 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P080pct_T15s_A02_S01.csv | NAM | 80 | 15 | 2 | not stated | 23.28 | 16.71 | 1120.5 | 236 | 2760.5 | 18726 | No | No | 533.66 | 486.98 | -32.62 | 23.52 | 405.64 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P080pct_T20s_A02_S01.csv | NAM | 80 | 20 | 2 | not stated | 29.21 | 19.63 | 1552 | 361 | 3313.5 | 27033 | No | No | 717.26 | 608.49 | -41.32 | 29.34 | 591.47 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260128_P080pct_T25s_A02_S01.csv | NAM | 80 | 25 | 2 | not stated | 32.47 | 21.89 | 1787.5 | 418 | 3807.5 | 33683 | No | No | 817.08 | 667.82 | -45.27 | 32.69 | 732.19 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P050pct_T35s_A02_S01.csv | NAM | 50 | 35 | 2 | same spot | 21.27 | 23.11 | 1522.5 | 439.5 | 4740 | 22733 | No | No | 819.51 | 603.31 | -35.73 | 21.38 | 508.72 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P070pct_T30s_A02_S01.csv | NAM | 70 | 30 | 2 | different spot | 37.13 | 32.47 | 5711.5 | 1334.5 | 7378.5 | 32959 | No | No | 2999.63 | 3389.53 | -231.93 | 37.39 | 1247.67 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S01.csv | NAM | 75 | 30 | 2 | different spot | 29.53 | 25.31 | 3150.5 | 765.5 | 5225.5 | 31741 | No | No | 1653.78 | 1731.3 | -114.39 | 29.83 | 777.89 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S02.csv | NAM | 75 | 30 | 2 | different spot | 24.64 | 25.07 | 2550.5 | 684 | 4928.5 | 31608 | No | No | 1380.08 | 1277.66 | -81.47 | 24.7 | 635.93 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S03.csv | NAM | 75 | 30 | 2 | different spot | 24.54 | 29.61 | 5421.5 | 1409 | 7200.5 | 32607 | No | No | 3121.72 | 3584.13 | -246.56 | 24.7 | 762.69 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S04.csv | NAM | 75 | 30 | 2 | different spot | 37.04 | 30.4 | 5838 | 1333.5 | 7529.5 | 30358.5 | No | No | 3072.95 | 3591.7 | -251.26 | 37.23 | 1161.96 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S05.csv | NAM | 75 | 30 | 2 | different spot | 30.14 | 24.81 | 3641 | 1013 | 5583.5 | 32505 | No | No | 2029.27 | 1887.17 | -120.09 | 30.3 | 775.97 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S06.csv | NAM | 75 | 30 | 2 | different spot | 35.21 | 27.32 | 5581 | 1415 | 7523.5 | 29244.5 | No | No | 2844.3 | 2712.28 | -169.91 | 35.51 | 995.96 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S07.csv | NAM | 75 | 30 | 2 | different spot | 23.95 | 25.43 | 2913 | 929 | 6361.5 | 31359 | No | No | 1834.18 | 1377.85 | -72.43 | 24 | 633.92 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S08.csv | NAM | 75 | 30 | 2 | different spot | 21.71 | 26.86 | 3122 | 1027.5 | 6393 | 28505.5 | No | No | 1893.07 | 1588.9 | -90.29 | 21.72 | 613.6 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S09.csv | NAM | 75 | 30 | 2 | different spot | 23.8 | 31.22 | 3758.5 | 1076.5 | 6772 | 28498 | No | No | 2152.39 | 2071.03 | -132.37 | 24.08 | 773.04 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S10.csv | NAM | 75 | 30 | 2 | different spot | 34.75 | 28.17 | 4234 | 982 | 6898.5 | 28503 | No | No | 2159.15 | 2366.86 | -165.37 | 35.12 | 1010.14 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P075pct_T30s_A02_S11.csv | NAM | 75 | 30 | 2 | different spot | 24.11 | 25.71 | 3444.5 | 1415.5 | 5794.5 | 30848.5 | No | No | 2346.33 | 1568.68 | -70.71 | 28.16 | 657.88 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S01.csv | NAM | 80 | 30 | 2 | different spot | 24.86 | 27.13 | 3345 | 1357.5 | 5575 | 33083.5 | No | No | 2149.4 | 1217.47 | -36.85 | 28.97 | 697.83 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S02.csv | NAM | 80 | 30 | 2 | different spot | 19.51 | 24.97 | 3444.5 | 1240.5 | 5710.5 | 34213 | No | No | 2249.05 | 1881.93 | -107.41 | 20.59 | 517.57 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S03.csv | NAM | 80 | 30 | 2 | different spot | 25.31 | 23.74 | 3053.5 | 778 | 5242 | 31573.5 | No | No | 1675.44 | 1804.97 | -123.66 | 25.55 | 619.07 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S04.csv | NAM | 80 | 30 | 2 | different spot | 25.22 | 25.15 | 3078 | 917.5 | 5353 | 34314 | No | No | 1798.75 | 1629.4 | -99.66 | 25.28 | 662.3 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S05.csv | NAM | 80 | 30 | 2 | different spot | 38.41 | 26.58 | 4229.5 | 992 | 5911 | 32122 | No | No | 2145.73 | 2249.87 | -154.58 | 38.79 | 1050.8 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T30s_A02_S06.csv | NAM | 80 | 30 | 2 | different spot | 26.86 | 21.53 | 1970 | 445.5 | 4566 | 33034 | No | No | 975.73 | 858.28 | -56.8 | 26.92 | 594.51 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S01.csv | NAM | 80 | 35 | 2 | not stated | 27.62 | 26.62 | 3155.5 | 1044.5 | 5915 | 40189.5 | No | No | 1918.21 | 1428.4 | -81.14 | 27.96 | 761.29 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S02.csv | NAM | 80 | 35 | 2 | not stated | 28.55 | 26.13 | 2977 | 996.5 | 5722 | 41357 | No | No | 1773.75 | 1235.86 | -67.27 | 28.9 | 771.85 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S03.csv | NAM | 80 | 35 | 2 | not stated | 31.33 | 27.01 | 4134.5 | 1101.5 | 6289.5 | 40154 | No | No | 2255.57 | 2227.59 | -144.74 | 31.65 | 879.79 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S04.csv | NAM | 80 | 35 | 2 | not stated | 39.65 | 28.46 | 4145.5 | 961.5 | 6193 | 35991 | No | No | 2076.92 | 2103.35 | -144.41 | 39.97 | 1159.83 | 1 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S05.csv | NAM | 80 | 35 | 2 | not stated | 29.29 | 30.47 | 5393 | 2131 | 6969 | 37261 | No | No | 3008.37 | 2011.39 | 26.57 | 42.87 | 915.4 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S06.csv | NAM | 80 | 35 | 2 | not stated | 29.38 | 24.89 | 3076.5 | 965 | 5754.5 | 34361.5 | No | No | 1793.18 | 1383.19 | -82.7 | 29.75 | 759.7 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T35s_A02_S07.csv | NAM | 80 | 35 | 2 | same spot | 29.91 | 26.26 | 2320 | 666 | 5443.5 | 34102 | No | No | 1210.39 | 817.97 | -47.54 | 30.13 | 816.35 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S01.csv | NAM | 80 | 40 | 2 | not stated | 33.06 | 25.43 | 3337.5 | 1113 | 7796.5 | 43421.5 | No | No | 1989.33 | 1467.18 | -86.03 | 33.22 | 865.81 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S02.csv | NAM | 80 | 40 | 2 | not stated | 30.06 | 26.41 | 3537.5 | 943 | 7720.5 | 44708.5 | No | No | 1924.4 | 1839.61 | -122.34 | 30.15 | 824.64 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S03.csv | NAM | 80 | 40 | 2 | not stated | 28.89 | 29.1 | 3506 | 1042 | 7760 | 46215.5 | No | No | 2006.45 | 1627.54 | -97.03 | 29.07 | 873.52 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S04.csv | NAM | 80 | 40 | 2 | not stated | 32.11 | 30.26 | 3896 | 1125.5 | 8000.5 | 45726 | No | No | 2156.99 | 1897.35 | -119.28 | 32.39 | 997.01 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S05.csv | NAM | 80 | 40 | 2 | not stated | 46.07 | 30.5 | 5512.5 | 1678.5 | 7587.5 | 41911 | No | No | 3107.58 | 2547.01 | -153.24 | 46.45 | 1457.17 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S06.csv | NAM | 80 | 40 | 2 | not stated | 50.76 | 35.25 | 6759 | 1620.5 | 8600 | 41741.5 | No | No | 3369.74 | 3379.96 | -221.09 | 50.89 | 1848.25 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S07.csv | NAM | 80 | 40 | 2 | not stated | 33.65 | 26.89 | 3198.5 | 886.5 | 6023 | 41945.5 | No | No | 1672.97 | 1376.92 | -84.66 | 33.76 | 936.88 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P080pct_T40s_A02_S08.csv | NAM | 80 | 40 | 2 | not stated | 52.25 | 33.27 | 6448 | 1549 | 8289 | 45617.5 | No | No | 3240.49 | 3226.23 | -217.13 | 52.55 | 1788.3 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S01.csv | NAM | 85 | 30 | 2 | different spot | 23.67 | 26.38 | 2891 | 715.5 | 5124.5 | 36382.5 | No | No | 1551.09 | 1627.26 | -110.1 | 23.75 | 649.92 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S02.csv | NAM | 85 | 30 | 2 | different spot | 31.29 | 27.39 | 3769 | 948.5 | 5614.5 | 34413.5 | No | No | 2015.02 | 2042.5 | -134.58 | 31.36 | 884.15 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S03.csv | NAM | 85 | 30 | 2 | different spot | 38.76 | 30.44 | 4745.5 | 1177.5 | 6402.5 | 37693.5 | No | No | 2485.94 | 2484.87 | -167.1 | 38.93 | 1222.54 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S04.csv | NAM | 85 | 30 | 2 | different spot | 45.01 | 31.75 | 5423 | 1314.5 | 7080 | 38536.5 | No | No | 2748.82 | 2814.11 | -190.34 | 45.19 | 1471.62 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S05.csv | NAM | 85 | 30 | 2 | different spot | 41.63 | 30.2 | 3862 | 925 | 5622 | 38736 | No | No | 1894.9 | 1741.02 | -113.3 | 41.98 | 1285.91 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P085pct_T30s_A02_S06.csv | NAM | 85 | 30 | 2 | different spot | 24.3 | 24.64 | 1976 | 463 | 4515 | 33980 | No | No | 969.4 | 857.75 | -55.08 | 24.48 | 622.79 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P090pct_T30s_A02_S01.csv | NAM | 90 | 30 | 2 | different spot | 50.92 | 27.58 | 4010 | 945 | 5657.5 | 41124.5 | No | No | 1878.58 | 1630.18 | -104.82 | 51.38 | 1444.44 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P090pct_T30s_A02_S02.csv | NAM | 90 | 30 | 2 | different spot | 54.44 | 30.4 | 5228 | 1303.5 | 6867.5 | 42398 | No | No | 2617.96 | 2370.23 | -154.99 | 54.94 | 1703.01 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P090pct_T30s_A02_S03.csv | NAM | 90 | 30 | 2 | different spot | 45.31 | 27.08 | 3699 | 885.5 | 5479 | 40849 | No | No | 1781.57 | 1620.88 | -106.53 | 45.48 | 1263.47 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P090pct_T35s_A02_S01.csv | NAM | 90 | 35 | 2 | same spot | 31.93 | 27.28 | 2576.5 | 740 | 6471 | 39101 | No | No | 1287.96 | 857.56 | -46.64 | 31.97 | 896.82 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S01.csv | NAM | 95 | 30 | 2 | different spot | 52.13 | 27.73 | 3997 | 904 | 5636.5 | 42700 | No | No | 1811.4 | 1630.51 | -108.77 | 52.46 | 1487.56 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S02.csv | NAM | 95 | 30 | 2 | different spot | 46.75 | 26.55 | 4838 | 1231 | 6477.5 | 42762 | No | No | 2515.76 | 2472.75 | -167.25 | 47.39 | 1285.21 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S03.csv | NAM | 95 | 30 | 2 | different spot | 41.71 | 28.95 | 3984 | 989 | 5653.5 | 42004 | No | No | 2000.57 | 1813.88 | -119.61 | 42 | 1250.23 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S04.csv | NAM | 95 | 30 | 2 | different spot | 44.36 | 26.68 | 3629.5 | 880.5 | 5497 | 41241 | No | No | 1788.35 | 1562.08 | -100.97 | 44.71 | 1209.58 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P095pct_T30s_A02_S05.csv | NAM | 95 | 30 | 2 | different spot | 48.77 | 32.46 | 4478 | 1023.5 | 6126 | 44984 | No | No | 2098.8 | 1841 | -121.3 | 49.15 | 1628.57 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S01.csv | NAM | 100 | 30 | 2 | different spot | 57.91 | 27.49 | 4181.5 | 913.5 | 5813 | 44275 | No | No | 1904.26 | 1640.38 | -104.57 | 58.09 | 1635.24 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S02.csv | NAM | 100 | 30 | 2 | different spot | 52.46 | 30.76 | 4127 | 900 | 5746 | 45496.5 | No | No | 1830.89 | 1514.66 | -97.22 | 52.64 | 1656.36 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S03.csv | NAM | 100 | 30 | 2 | different spot | 53.48 | 28.98 | 4080.5 | 889.5 | 5712 | 44989.5 | No | No | 1807.96 | 1519.86 | -97.17 | 53.76 | 1593.32 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S04.csv | NAM | 100 | 30 | 2 | different spot | 33.04 | 22.93 | 2651.5 | 672 | 4952.5 | 41092 | No | No | 1369.29 | 1278.17 | -85.08 | 33.1 | 778.3 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P100pct_T30s_A02_S05.csv | NAM | 100 | 30 | 2 | different spot | 33.91 | 23.82 | 2765 | 772.5 | 4997.5 | 40053 | No | No | 1455.09 | 1160.18 | -70.48 | 34.21 | 830.7 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260403_P100pct_T35s_A02_S01.csv | NAM | 100 | 35 | 2 | same spot | 37.72 | 24.49 | 2590.5 | 727 | 6485 | 44393 | No | No | 1318.07 | 837.92 | -44.14 | 38.09 | 956.84 | 0 | 1 | Valid |  | No
NAM_Bulk_20260403_P100pct_T35s_A02_S02.csv | NAM | 100 | 35 | 2 | same spot | 35.53 | 26.1 | 2572 | 735 | 6466.5 | 44909.5 | No | No | 1284.66 | 771.33 | -40.61 | 35.93 | 966.02 | 0 | 1 | Valid |  | No
NAM_Bulk_20260529_P005pct_T05s_A01_S01.csv | NAM | 5 | 5 | 1 | not stated | 2.46 | 25.59 | 204 | 13 | 1862 | 2484 | No | No | 83.25 | 62.39 | -1.62 | 2.48 | 70.01 | 0 | 0.41 | Low SNR | SNR 2.5 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P005pct_T05s_A01_S02.csv | NAM | 5 | 5 | 1 | not stated | 3.11 | 23.43 | 286 | -33 | 1904 | 2524 | No | No | 123.08 | 179.09 | -11.81 | 3.13 | 82.59 | 0 | 0.46 | Low SNR | SNR 2.9 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P010pct_T05s_A01_S01.csv | NAM | 10 | 5 | 1 | not stated | 2.72 | 24.5 | 320 | 20 | 1918 | 2882 | No | No | 136.04 | 186.65 | -11.8 | 2.73 | 73.95 | 0 | 0.51 | Low SNR | SNR 2.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P010pct_T05s_A01_S02.csv | NAM | 10 | 5 | 1 | not stated | 2.76 | 26.88 | 372 | 45 | 1894 | 3011 | No | No | 184.34 | 215.99 | -14.76 | 2.77 | 80.21 | 0 | 0.58 | Low SNR | SNR 2.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P010pct_T05s_A01_S03.csv | NAM | 10 | 5 | 1 | not stated | 2.7 | 23.39 | 367 | 57 | 1892 | 3142 | No | No | 186.63 | 194.09 | -11.91 | 3.45 | 69.53 | 0 | 0.51 | Low SNR | SNR 2.7 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P010pct_T05s_A01_S04.csv | NAM | 10 | 5 | 1 | not stated | 2.75 | 24.12 | 337 | 31 | 1929 | 3001 | No | No | 158.32 | 177.95 | -10.74 | 2.83 | 75.59 | 0 | 0.38 | Low SNR | SNR 2.8 < 10 at low power / short integration; genuine NAM spectrum used only for parameter comparison. | No
NAM_Bulk_20260529_P070pct_T10s_A03_S01.csv | NAM | 70 | 10 | 3 | not stated | 13.29 | 16.37 | 819.67 | 302.67 | 2911.33 | 14006.67 | No | No | 499.69 | 270.31 | -11.43 | 13.29 | 226.55 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S02.csv | NAM | 70 | 10 | 3 | not stated | 18.91 | 14.65 | 1159.67 | 347.33 | 3061.33 | 16690 | No | No | 641.62 | 511.69 | -31.12 | 19.12 | 285.13 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S03.csv | NAM | 70 | 10 | 3 | not stated | 16.65 | 17.32 | 1276.33 | 282.67 | 3216.33 | 16771 | No | No | 711.9 | 745.24 | -51.64 | 16.86 | 312.12 | 0 | 0.93 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S04.csv | NAM | 70 | 10 | 3 | not stated | 12.32 | 15.12 | 777.33 | 187 | 2504 | 16412.33 | No | No | 427.41 | 432.9 | -30.09 | 12.48 | 195.21 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S05.csv | NAM | 70 | 10 | 3 | not stated | 12.07 | 18.83 | 847.67 | 176.67 | 2444 | 15882.67 | No | No | 372.52 | 403.38 | -30.6 | 12.2 | 243.1 | 0 | 0.86 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S06.csv | NAM | 70 | 10 | 3 | same spot | 17.84 | 13.95 | 687 | 221.67 | 2235 | 15450.67 | No | No | 363.98 | 230.36 | -3.91 | 18.05 | 258.95 | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S07.csv | NAM | 70 | 10 | 3 | same spot | 16.62 | 15.95 | 697.33 | 223.67 | 2211.33 | 16333 | No | No | 364.39 | 232.84 | -3.26 | 16.73 | 274.86 | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S08.csv | NAM | 70 | 10 | 3 | same spot | 18.45 | 13.56 | 695.33 | 210.33 | 2193.67 | 16162.33 | No | No | 348.94 | 232.06 | -2.44 | 18.6 | 261.06 | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S09.csv | NAM | 70 | 10 | 3 | same spot | 17.75 | 14.19 | 693 | 213 | 2200.33 | 17714.33 | No | No | 356.72 | 229.79 | -2.21 | 17.87 | 259.92 | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S10.csv | NAM | 70 | 10 | 3 | same spot | 19.52 | 13.57 | 705.33 | 227.67 | 2227.33 | 17965 | No | No | 362.21 | 234.95 | -1.68 | 19.65 | 274.03 | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S11.csv | NAM | 70 | 10 | 3 | same spot | 17.49 | 14.44 | 669.33 | 16.33 | 2059 | 17231.33 | No | No | 348.06 | 227.87 | -1.43 | 17.61 | 264.07 | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S12.csv | NAM | 70 | 10 | 3 | same spot | 16.64 | 14.44 | 637.67 | 211 | 2050 | 17076.33 | No | No | 339.67 | 224.01 | -1.74 | 16.74 | 258.6 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S13.csv | NAM | 70 | 10 | 3 | same spot | 17.19 | 13.67 | 628.33 | 201.33 | 2047 | 16072.33 | No | No | 333.13 | 218.28 | -1.6 | 17.24 | 245.91 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S14.csv | NAM | 70 | 10 | 3 | same spot | 16.96 | 14.71 | 647.33 | 217.33 | 2065.67 | 17056.33 | No | No | 341.86 | 221.43 | -1.49 | 17.07 | 261.98 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S15.csv | NAM | 70 | 10 | 3 | same spot | 16.81 | 14.05 | 628 | 210 | 2059 | 16633.33 | No | No | 335.47 | 212.25 | -1.64 | 17.01 | 246.7 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A03_S16.csv | NAM | 70 | 10 | 3 | same spot | 16.09 | 15.02 | 649.67 | 214 | 2049.67 | 16610.67 | No | No | 339.33 | 221.86 | -1.61 | 16.14 | 250.85 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S01.csv | NAM | 70 | 10 | 5 | not stated | 23.04 | 11.94 | 1043.6 | 333.8 | 3132.4 | 16030 | No | No | 600.98 | 411.23 | -23.19 | 23.32 | 285.66 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S02.csv | NAM | 70 | 10 | 5 | not stated | 17.37 | 13.54 | 888.2 | 245 | 3006.2 | 16257.8 | No | No | 475.31 | 386.74 | -24.46 | 17.58 | 244.31 | 0 | 0.96 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S03.csv | NAM | 70 | 10 | 5 | not stated | 18.38 | 12.85 | 963.6 | 337.8 | 3033.2 | 19107.4 | No | No | 586.02 | 407.37 | -21.99 | 18.54 | 242.45 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S04.csv | NAM | 70 | 10 | 5 | not stated | 21.07 | 12.24 | 909.4 | 278.4 | 2971.4 | 18602.2 | No | No | 516.32 | 381.16 | -22.17 | 21.24 | 266.43 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S05.csv | NAM | 70 | 10 | 5 | not stated | 17.24 | 12.59 | 829.6 | 314.4 | 2958.4 | 16510 | No | No | 528.68 | 308.31 | -14.45 | 17.3 | 229.78 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S06.csv | NAM | 70 | 10 | 5 | same spot | 26.62 | 10.17 | 687.6 | 230.4 | 2088.6 | 16970.2 | No | No | 360.7 | 229.18 | -2.48 | 26.91 | 277.37 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S07.csv | NAM | 70 | 10 | 5 | same spot | 21.83 | 11.62 | 664.4 | 226 | 2075.2 | 18062.2 | No | No | 355.39 | 223.93 | -1.55 | 22.03 | 261.25 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S08.csv | NAM | 70 | 10 | 5 | same spot | 21.96 | 11.59 | 653.2 | 217.6 | 2071.4 | 17374.4 | No | No | 349.35 | 214.39 | -1.44 | 22.2 | 262.98 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S09.csv | NAM | 70 | 10 | 5 | same spot | 20.92 | 11.91 | 634.2 | 211.2 | 2030.4 | 16183.2 | No | No | 335.28 | 214.34 | -1.64 | 21.13 | 257.95 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S10.csv | NAM | 70 | 10 | 5 | same spot | 21.3 | 11.52 | 651.2 | 202.4 | 2047.8 | 15970.4 | No | No | 335.64 | 215.99 | -1.63 | 21.5 | 252.41 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S11.csv | NAM | 70 | 10 | 5 | same spot | 20.2 | 12.32 | 735.2 | 214.6 | 2073.2 | 15597.2 | No | No | 337.69 | 210.36 | -1.42 | 20.42 | 256.45 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S12.csv | NAM | 70 | 10 | 5 | same spot | 21.69 | 11.52 | 702.2 | 224.2 | 2056.8 | 15366.8 | No | No | 336.72 | 214.68 | -0.89 | 21.85 | 259.86 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S13.csv | NAM | 70 | 10 | 5 | same spot | 21.56 | 11.74 | 687.8 | 211.4 | 2065.6 | 14760.4 | No | No | 333.32 | 210.22 | -1.5 | 21.79 | 263.69 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S14.csv | NAM | 70 | 10 | 5 | same spot | 21.85 | 11.35 | 699.2 | 223.6 | 2063.6 | 14644.4 | No | No | 333.55 | 212.78 | -1.09 | 22.04 | 256.42 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S15.csv | NAM | 70 | 10 | 5 | same spot | 19.94 | 12.49 | 641.4 | 213.8 | 2052.2 | 14902.4 | No | No | 338.02 | 220.83 | -0.92 | 20.01 | 255.93 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S16.csv | NAM | 70 | 10 | 5 | same spot | 21.2 | 11.54 | 636.4 | 209 | 2050.2 | 14691.8 | No | No | 334.56 | 218.62 | -1.11 | 21.43 | 254.39 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P070pct_T10s_A05_S17.txtr | NAM | 70 | 10 | 5 | same spot | 22.17 | 11.46 | 653.8 | 211.2 | 2057.2 | 14505 | No | No | 336.92 | 222.25 | -1.18 | 22.34 | 261.97 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S01.csv | NAM | 80 | 10 | 3 | not stated | 16.55 | 15.59 | 901.67 | 288 | 2349 | 14999 | No | No | 512.42 | 343 | -19.93 | 16.71 | 265.87 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S02.csv | NAM | 80 | 10 | 3 | not stated | 16.3 | 15.56 | 910.67 | 265.33 | 2377.33 | 17723.67 | No | No | 500.59 | 376.24 | -23.6 | 16.41 | 263 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S03.csv | NAM | 80 | 10 | 3 | not stated | 18.93 | 15.72 | 1132.67 | 446.33 | 2480.33 | 20839.33 | No | No | 682.18 | 395.05 | -7.02 | 19.11 | 305.84 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S04.csv | NAM | 80 | 10 | 3 | not stated | 18.49 | 15.41 | 1090.67 | 328 | 2559.67 | 17987.33 | No | No | 611.41 | 491.71 | -30.58 | 18.54 | 300.35 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S05.csv | NAM | 80 | 10 | 3 | not stated | 20.34 | 18.11 | 1866.67 | 518 | 3160.33 | 16874 | No | No | 1044.71 | 962.02 | -66.2 | 20.54 | 380.52 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S06.csv | NAM | 80 | 10 | 3 | not stated | 21.43 | 16.16 | 1593.67 | 625.67 | 3004.67 | 18849.33 | No | No | 1070.22 | 653.09 | -29.05 | 21.84 | 357.26 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A03_S07.csv | NAM | 80 | 10 | 3 | not stated | 16.72 | 15 | 984.67 | 370 | 2461.67 | 17151 | No | No | 619.34 | 372.6 | -17.08 | 16.88 | 264.92 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S01.txtr | NAM | 80 | 10 | 5 | not stated | 21.09 | 12.56 | 936.8 | 211.8 | 2387.6 | 17154.8 | No | No | 466.33 | 432.17 | -30.28 | 21.27 | 274.67 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S02.csv | NAM | 80 | 10 | 5 | not stated | 16.73 | 11.58 | 1102.2 | 470.8 | 2630.8 | 17562 | No | No | 770.24 | 472.71 | -20.88 | 19.98 | 203.37 | 0 | 0.94 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S03.csv | NAM | 80 | 10 | 5 | not stated | 18.98 | 11.87 | 896.6 | 334 | 2364.4 | 16116 | No | No | 552.36 | 318.25 | -8.7 | 19.18 | 232.1 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S04.csv | NAM | 80 | 10 | 5 | not stated | 19.56 | 11.57 | 961.8 | 360.4 | 2501.2 | 17593.6 | No | No | 641.65 | 438.12 | -22.44 | 19.71 | 237.78 | 0 | 0.97 | Valid |  | No
NAM_Bulk_20260529_P080pct_T10s_A05_S05.csv | NAM | 80 | 10 | 5 | not stated | 23.1 | 11.34 | 971 | 305.2 | 2454.4 | 17343.2 | No | No | 570.62 | 426.54 | -24.97 | 23.36 | 274.67 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S01.csv | NAM | 90 | 10 | 3 | not stated | 22.93 | 16.24 | 1126.33 | 315.67 | 2477 | 18741.33 | No | No | 586.49 | 416.9 | -24.16 | 23.12 | 384.79 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S02.csv | NAM | 90 | 10 | 3 | not stated | 19.94 | 15.72 | 1273 | 388.67 | 2479.67 | 21654.67 | No | No | 718.16 | 463.07 | 1.91 | 20.6 | 323.47 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S03.csv | NAM | 90 | 10 | 3 | not stated | 20.55 | 16.46 | 996.33 | 136.67 | 2379.67 | 19108.33 | No | No | 507.8 | 390.15 | -24.45 | 20.71 | 345.63 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S04.csv | NAM | 90 | 10 | 3 | not stated | 22.61 | 15.8 | 1191.33 | 237.67 | 2552.67 | 18035.67 | No | No | 617.42 | 546.81 | -36.72 | 22.86 | 370.78 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S05.csv | NAM | 90 | 10 | 3 | not stated | 22.68 | 14.67 | 1025.67 | 355.33 | 2361.67 | 19578.33 | No | No | 589.72 | 347.81 | -11.9 | 22.92 | 341.32 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S06.csv | NAM | 90 | 10 | 3 | not stated | 17.1 | 14.29 | 874 | 261 | 2329.67 | 19075 | No | No | 484.92 | 348.25 | -20.98 | 17.31 | 252.3 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S07.csv | NAM | 90 | 10 | 3 | same spot | 20.54 | 15.4 | 845.33 | 301.33 | 2299.67 | 21214 | No | No | 463.85 | 285.79 | -4.11 | 20.58 | 328.74 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S08.csv | NAM | 90 | 10 | 3 | same spot | 20.24 | 15.23 | 857.67 | 282.67 | 2311.67 | 19283.33 | No | No | 453.45 | 288.53 | -4.07 | 20.45 | 317.96 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S09.csv | NAM | 90 | 10 | 3 | same spot | 20.94 | 14.78 | 829 | 290.67 | 2293.67 | 19663 | No | No | 447.84 | 284.91 | -3.55 | 21.12 | 317.24 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S10.csv | NAM | 90 | 10 | 3 | same spot | 20.93 | 15.17 | 864 | 285.67 | 2306.33 | 21738.33 | No | No | 465 | 291.53 | -3.17 | 21.15 | 326.94 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S11.csv | NAM | 90 | 10 | 3 | same spot | 19.89 | 15.99 | 854 | 285.67 | 2332.67 | 22520 | No | No | 463.03 | 296.61 | -2.77 | 20.1 | 327.31 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S12.csv | NAM | 90 | 10 | 3 | same spot | 21.93 | 14.59 | 857.33 | 279 | 2305.67 | 22120 | No | No | 456.21 | 293.44 | -2.8 | 22.16 | 330.49 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S13.csv | NAM | 90 | 10 | 3 | same spot | 20.34 | 15.25 | 828.67 | 285.33 | 2325 | 20720 | No | No | 446.96 | 281.53 | -2.15 | 20.54 | 325.04 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S14.csv | NAM | 90 | 10 | 3 | same spot | 19.55 | 15.6 | 839 | 286.67 | 2290 | 19680.33 | No | No | 433.29 | 273.57 | -2.1 | 19.6 | 316.01 | 0 | 0.98 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S15.csv | NAM | 90 | 10 | 3 | same spot | 20.16 | 15.13 | 807 | 278.33 | 2291.33 | 20712.67 | No | No | 433.44 | 276.41 | -2.1 | 20.37 | 315.24 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S16.csv | NAM | 90 | 10 | 3 | same spot | 18.66 | 15.84 | 802 | 272 | 2282.33 | 19426 | No | No | 425.83 | 270.16 | -1.54 | 18.8 | 304.8 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S17.csv | NAM | 90 | 10 | 3 | same spot | 19.8 | 14.91 | 793 | 274 | 2281 | 18922.33 | No | No | 421.71 | 269.6 | -1.25 | 20.01 | 303.86 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A03_S18.csv | NAM | 90 | 10 | 3 | same spot | 19.3 | 15.28 | 800 | 279.67 | 2268 | 20246.67 | No | No | 429.71 | 273.23 | -1.45 | 19.52 | 305.52 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S01.csv | NAM | 90 | 10 | 5 | not stated | 23.72 | 11.35 | 866.6 | 270.4 | 2299.8 | 19423.8 | No | No | 475.43 | 310.11 | -18.18 | 23.79 | 275.97 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S02.csv | NAM | 90 | 10 | 5 | not stated | 26.14 | 12.81 | 1034 | 271.2 | 2390.2 | 19562.6 | No | No | 517.13 | 400.03 | -25.62 | 26.35 | 342.02 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S03.csv | NAM | 90 | 10 | 5 | not stated | 31.8 | 11.44 | 1146.2 | 320.6 | 2379.6 | 18134.8 | No | No | 608.72 | 458.72 | -28.74 | 32.16 | 375.51 | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S04.csv | NAM | 90 | 10 | 5 | not stated | 26.13 | 12.21 | 935.4 | 244.4 | 2219.2 | 18456.8 | No | No | 486.99 | 312.27 | -17.14 | 26.35 | 330.07 | 2 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S05.csv | NAM | 90 | 10 | 5 | not stated | 27.03 | 11.36 | 1002.8 | 313.8 | 2301.2 | 20913 | No | No | 573.86 | 317.95 | -12.87 | 27.26 | 319.79 | 2 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S06.csv | NAM | 90 | 10 | 5 | same spot | 27.6 | 11.46 | 953.8 | 331.6 | 2402 | 22019.4 | No | No | 540.14 | 313.71 | -10.19 | 27.91 | 324.66 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S07.csv | NAM | 90 | 10 | 5 | same spot | 24.75 | 12.36 | 906.4 | 304.6 | 2372.4 | 20380.8 | No | No | 504.65 | 299.62 | -7.91 | 25 | 315.49 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S08.csv | NAM | 90 | 10 | 5 | same spot | 24.36 | 13.01 | 902 | 307.6 | 2379.8 | 22224 | No | No | 503.85 | 302.13 | -6.91 | 24.65 | 326.6 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S09.csv | NAM | 90 | 10 | 5 | same spot | 25.43 | 12.25 | 889.6 | 319.8 | 2344.6 | 22095.6 | No | No | 494.73 | 298.21 | -5.69 | 25.76 | 319.64 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S10.csv | NAM | 90 | 10 | 5 | same spot | 23 | 13.49 | 863.6 | 300.6 | 2329.8 | 20870.2 | No | No | 482.22 | 292.1 | -5.34 | 23.08 | 322.24 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S11.csv | NAM | 90 | 10 | 5 | same spot | 21.71 | 14.03 | 849.2 | 298.4 | 2323 | 20235 | No | No | 472.63 | 289.48 | -4.36 | 21.92 | 312.36 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S12.csv | NAM | 90 | 10 | 5 | same spot | 25.29 | 12.1 | 851.4 | 289 | 2333.2 | 20155.4 | No | No | 463.16 | 288.23 | -3.7 | 25.53 | 316.91 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S13.csv | NAM | 90 | 10 | 5 | same spot | 25.24 | 11.88 | 833 | 293.8 | 2324.4 | 19840.4 | No | No | 454.68 | 283.96 | -3.45 | 25.5 | 306.77 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S14.csv | NAM | 90 | 10 | 5 | same spot | 25.95 | 11.57 | 829.2 | 286.8 | 2332.8 | 19930.6 | No | No | 453.43 | 278.29 | -3.25 | 26.02 | 307.14 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S15.csv | NAM | 90 | 10 | 5 | same spot | 23.24 | 12.94 | 853.6 | 293.2 | 2328 | 19715.8 | No | No | 450.87 | 281.21 | -2.91 | 23.31 | 308.32 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260529_P090pct_T10s_A05_S16.csv | NAM | 90 | 10 | 5 | same spot | 24.11 | 12.12 | 824.4 | 279.2 | 2302.4 | 19195 | No | No | 444.5 | 280.82 | -3.03 | 24.31 | 305.42 | 0 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S01.csv | NAM | 70 | 10 | 5 | different spot | 18.98 | 11.85 | 793 | 98.6 | 2315.2 | 13882.8 | No | No | 434.6 | 323.84 | -19.35 | 19.04 | 234.82 | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S02.csv | NAM | 70 | 10 | 5 | different spot | 21.7 | 12.66 | 840.4 | -40.8 | 2328.4 | 13323.6 | No | No | 418.94 | 352.99 | -22.87 | 21.82 | 283.05 | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S03.csv | NAM | 70 | 10 | 5 | different spot | 23.19 | 11.49 | 857.4 | -18.8 | 2342.6 | 15167 | No | No | 461.02 | 343.31 | -20.27 | 23.34 | 277.16 | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S04.csv | NAM | 70 | 10 | 5 | different spot | 18.6 | 11.25 | 736 | -100 | 2265.2 | 14680.4 | No | No | 398.85 | 282.69 | -16.36 | 18.85 | 220.82 | 1 | 0.97 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S05.csv | NAM | 70 | 10 | 5 | different spot | 23.56 | 12.07 | 913.4 | 17.2 | 2386.2 | 17422.4 | No | No | 481.62 | 374.01 | -22.54 | 23.64 | 294.36 | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S06.csv | NAM | 70 | 10 | 5 | different spot | 21.4 | 12.12 | 1177.6 | 177.8 | 2643.6 | 17829 | No | No | 617.83 | 665.81 | -46.12 | 21.59 | 266.15 | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S07.csv | NAM | 70 | 10 | 5 | different spot | 28.53 | 12.74 | 1366.6 | 259.8 | 2754.2 | 18385.2 | No | No | 679.26 | 747.37 | -52.47 | 28.61 | 374.66 | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S08.csv | NAM | 70 | 10 | 5 | different spot | 22.51 | 11.85 | 897.4 | 54.4 | 2728.4 | 16513 | No | No | 430.68 | 451.42 | -30.72 | 22.56 | 278.46 | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S09.csv | NAM | 70 | 10 | 5 | different spot | 16.98 | 12.23 | 1062 | 113.2 | 2856.8 | 14855.8 | No | No | 577.21 | 645.04 | -44.55 | 17.01 | 215.35 | 1 | 0.98 | Valid |  | No
NAM_Bulk_20260530_P070pct_T10s_A05_S10.csv | NAM | 70 | 10 | 5 | different spot | 23.76 | 11.67 | 889.8 | 35.4 | 2718.2 | 15195.4 | No | No | 410.82 | 410.1 | -27.79 | 23.85 | 290.6 | 1 | 0.99 | Valid |  | No
NAM_Bulk_20260530_P070pct_T30s_A05_S01.csv | NAM | 70 | 30 | 5 | not stated | 54 | 15.61 | 2322.2 | 511.4 | 4619.2 | 45379.2 | No | No | 1058.52 | 971.01 | -66.06 | 54.03 | 864.16 | 0 | 1 | Valid |  | No
NAM_Bulk_20260530_P070pct_T30s_A05_S02.csv | NAM | 70 | 30 | 5 | not stated | 40.31 | 19.09 | 3527.4 | 877 | 5854 | 49024.2 | No | No | 1865.06 | 1998.96 | -140.27 | 40.59 | 790.15 | 0 | 1 | Valid |  | No
NAM_Bulk_20260530_P070pct_T30s_A05_S03.csv | NAM | 70 | 30 | 5 | not stated | 45.87 | 17.44 | 2846.4 | 598 | 5926.4 | 47222.6 | No | No | 1201.82 | 1080.27 | -70.97 | 46.22 | 819.35 | 0 | 1 | Valid |  | No
NAM_Bulk_20260530_P070pct_T60s_A05_S01.csv | NAM | 70 | 60 | 5 | different spot | 68.43 | 20.63 | 4413.8 | 1009 | 8421.2 | 65535 | Yes | No | 2096.19 | 1973.13 | -136.25 | 68.81 | 1452.84 | 0 | 1 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 8421 counts (< 65 535), no negative counts, SNR 35.5, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | No
NAM_Bulk_20260530_P070pct_T60s_A05_S02.csv | NAM | 70 | 60 | 5 | different spot | 56.53 | 21.93 | 3290.8 | 751 | 7731.6 | 65523.2 | Yes | No | 1490.36 | 1241.71 | -84.22 | 56.77 | 1271.19 | 0 | 1 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 7732 counts (< 65 535), no negative counts, SNR 35.1, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | No
NAM_Bulk_20260530_P070pct_T60s_A05_S03.csv | NAM | 70 | 60 | 5 | different spot | 58.72 | 23.82 | 4237.2 | 1414 | 9146.6 | 65287.2 | Yes | No | 2364.24 | 1358.45 | -61.43 | 58.78 | 1440.85 | 0 | 1 | Valid – supporting (70%/60 s/5 acc) | Detector saturated only at the laser line (-2.2 to -2.2 cm-1, 1 pixel). In 400–1800 cm-1: max raw 9147 counts (< 65 535), no negative counts, SNR 33.9, r = 0.998 vs NAM reference. Retained as supporting NAM spectrum (separate acquisition group, not merged into reference). | No
NAM_Bulk_20260530_P090pct_T30s_A05_S01.csv | NAM | 90 | 30 | 5 | same spot | 52.12 | 18.16 | 2586.8 | 862 | 6083.8 | 45311 | No | No | 1405.07 | 868.09 | -26.23 | 52.2 | 979.88 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S02.csv | NAM | 90 | 30 | 5 | same spot | 53.14 | 17.74 | 2472.2 | 810 | 5982.8 | 46251.6 | No | No | 1312.42 | 835.8 | -19.58 | 53.24 | 967.96 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S03.csv | NAM | 90 | 30 | 5 | same spot | 51.93 | 17.66 | 2371 | 802.8 | 5910.2 | 51634 | No | No | 1258 | 810.98 | -14.48 | 52.02 | 945.88 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S04.csv | NAM | 90 | 30 | 5 | same spot | 52.21 | 17.95 | 2348.2 | 764.4 | 5871.2 | 51343.6 | No | No | 1231.49 | 804.12 | -11.7 | 52.24 | 970.21 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S05.csv | NAM | 90 | 30 | 5 | same spot | 57.79 | 16.03 | 2359.2 | 788.2 | 5858.6 | 51005.2 | No | No | 1223.95 | 799.89 | -10.7 | 57.89 | 949.63 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S06.csv | NAM | 90 | 30 | 5 | same spot | 48.54 | 18.67 | 2330.4 | 760.6 | 5886.2 | 45106.4 | No | No | 1208.12 | 797.67 | -11.34 | 48.7 | 936.52 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S07.csv | NAM | 90 | 30 | 5 | same spot | 53.98 | 17.24 | 2359.4 | 756.4 | 5869.8 | 44502 | No | No | 1210.54 | 794.91 | -10.73 | 54.27 | 954.14 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S08.csv | NAM | 90 | 30 | 5 | same spot | 54.27 | 17.16 | 2308.2 | 764.6 | 5839 | 49610.4 | No | No | 1184.04 | 789.37 | -8.6 | 54.59 | 955.16 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S09.csv | NAM | 90 | 30 | 5 | same spot | 51.58 | 18.25 | 2327.2 | 749.2 | 5859 | 52755.8 | No | No | 1189.23 | 791.21 | -7.38 | 51.76 | 960.27 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S10.csv | NAM | 90 | 30 | 5 | same spot | 53.52 | 17.76 | 2334 | 761.6 | 5865.4 | 45792.8 | No | No | 1196.03 | 797.87 | -6.92 | 53.54 | 977.73 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S11.csv | NAM | 90 | 30 | 5 | different spot | 56.57 | 16.69 | 2317.8 | 746.6 | 5867.4 | 43778.4 | No | No | 1177.35 | 779.63 | -8.47 | 56.73 | 966.45 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S12.csv | NAM | 90 | 30 | 5 | different spot | 47.22 | 17.76 | 3637.2 | 1323.8 | 6976.8 | 47144.2 | No | No | 2293.45 | 1581.97 | -81.94 | 47.49 | 863.22 | 0 | 0.99 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S13.csv | NAM | 90 | 30 | 5 | different spot | 48.54 | 17.19 | 2640.6 | 634.2 | 6090.6 | 51106 | No | No | 1311.14 | 1225.97 | -83.11 | 48.83 | 859.06 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S14.csv | NAM | 90 | 30 | 5 | different spot | 45.72 | 19.36 | 3058.6 | 704.2 | 6369.8 | 43173.8 | No | No | 1564.52 | 1540.53 | -104.72 | 45.83 | 914.28 | 0 | 1 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T30s_A05_S15.csv | NAM | 90 | 30 | 5 | different spot | 39.25 | 23.47 | 3095.6 | 705.8 | 6498 | 45705.8 | No | No | 1631.9 | 1584.44 | -105.5 | 39.26 | 971.91 | 0 | 0.98 | Valid |  | Yes
NAM_Bulk_20260530_P090pct_T60s_A05_S01.csv | NAM | 90 | 60 | 5 | different spot | 16.73 | 206.87 | 9798.4 | 1306.8 | 18624.8 | 65535 | Yes | No | 5109.08 | 5443.88 | -228.41 | 16.8 | 0 | 0 | 0.47 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 18625). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 310 counts (reference ≈ 26), r = 0.46. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S02.csv | NAM | 90 | 60 | 5 | different spot | 18.4 | 338.34 | 14041 | 1563 | 22234.8 | 65535 | Yes | No | 7564.04 | 8624.37 | -332.13 | 18.73 | 0 | 0 | 0.27 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 22235). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 554 counts (reference ≈ 26), r = 0.26. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S03.csv | NAM | 90 | 60 | 5 | different spot | 21.15 | 352.65 | 17015.2 | 2957.8 | 27829.4 | 65535 | Yes | No | 10416.47 | 10042.9 | -379.2 | 21.64 | 0 | 2 | 0.17 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 27829). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 622 counts (reference ≈ 26), r = 0.17. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S04.csv | NAM | 90 | 60 | 5 | different spot | 21.56 | 386.89 | 18123.6 | 2225 | 28506.4 | 65535 | Yes | No | 10383.5 | 11256.27 | -490.56 | 22.04 | 0 | 2 | 0.13 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 28506). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 694 counts (reference ≈ 26), r = 0.12. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S05.csv | NAM | 90 | 60 | 5 | different spot | 20.67 | 650.1 | 27315 | 2696 | 36617.4 | 65535 | Yes | No | 15183.25 | 17290.56 | -741.19 | 21.16 | 0 | 2 | 0.09 | Excluded – distorted (abnormal dark / ripple) | Laser-line saturation only (-4.2 to -2.2 cm-1); 400–1800 cm-1 not saturated (max raw 36617). Dark reference elevated (median 5910 vs 1 959) and strong high-frequency ripple: noise 1075 counts (reference ≈ 26), r = 0.09. Fingerprint unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S06.csv | NAM | 90 | 60 | 5 | same spot | 16.02 | 171.87 | 2642 | -2877.2 | 34426.4 | 65535 | Yes | No | -2009.74 | 1919.53 | 39.61 | 22.33 | 1761.4 | 0 | 0.37 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 88 % of 400–1800 cm-1 (min -2877); r = 0.36. Unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S07.csv | NAM | 90 | 60 | 5 | same spot | 14.94 | 266.96 | 1995.4 | -6048.6 | 30949.6 | 65535 | Yes | No | -4623.16 | 3894.19 | 185.6 | 23.09 | 2239.84 | 0 | 0.21 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 96 % of 400–1800 cm-1 (min -6049); r = 0.19. Unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S08.csv | NAM | 90 | 60 | 5 | same spot | 15.13 | 295.15 | 1736.4 | -8447.6 | 28975.6 | 65535 | Yes | No | -5549.24 | 4741.68 | 244.57 | 23.77 | 2089.12 | 0 | 0.18 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 97 % of 400–1800 cm-1 (min -8448); r = 0.17. Unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S09.csv | NAM | 90 | 60 | 5 | same spot | 14.27 | 349.46 | 1286.4 | -9894.6 | 27166.6 | 65535 | Yes | No | -6868.46 | 5652.01 | 311.88 | 23.02 | 2273.12 | 0 | 0.14 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 98 % of 400–1800 cm-1 (min -9895); r = 0.13. Unreliable. | No
NAM_Bulk_20260530_P090pct_T60s_A05_S10.csv | NAM | 90 | 60 | 5 | same spot | 14.29 | 396.93 | 973.8 | -11649.4 | 25199.6 | 65535 | Yes | No | -8423.82 | 6864.24 | 389.65 | 23.71 | 2516.65 | 0 | 0.1 | Excluded – distorted (negative counts) | Laser-line saturation only (-4.2 to -2.2 cm-1), but the stored dark reference is abnormally high (median 26587 vs 1 959 counts in the 70 % set), giving negative dark-subtracted counts in 99 % of 400–1800 cm-1 (min -11649); r = 0.09. Unreliable. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S01.csv | SUSPECT | 20 | 20 | 2 | different spot | 50.59 | 24.5 | 3694 | 699.5 | 4954 | 8546.5 | No | No | 1880.63 | 2728.27 | -185.89 | 50.62 | 1261.88 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S02.csv | SUSPECT | 20 | 20 | 2 | different spot | 87.92 | 23.46 | 4573 | 710.5 | 5852 | 10289.5 | No | No | 1735.03 | 2227.73 | -147.93 | 88.01 | 2088.61 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S03.csv | SUSPECT | 20 | 20 | 2 | different spot | 81.45 | 25.06 | 4385.5 | 677.5 | 5664.5 | 9935.5 | No | No | 1588.35 | 2000.17 | -131.58 | 81.51 | 2059.18 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S04.csv | SUSPECT | 20 | 20 | 2 | different spot | 69.88 | 22.55 | 4189.5 | 747 | 5468.5 | 9612.5 | No | No | 1969.89 | 2899.02 | -199.31 | 69.97 | 1599.97 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S05.csv | SUSPECT | 20 | 20 | 2 | different spot | 42.95 | 23.93 | 3594 | 688.5 | 4989.5 | 7964 | No | No | 1880.18 | 2796.37 | -195.02 | 42.98 | 1045.07 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S06.csv | SUSPECT | 20 | 20 | 2 | different spot | 58.35 | 24.74 | 4102.5 | 746.5 | 5353 | 9437.5 | No | No | 2041.29 | 3134.77 | -215.34 | 58.41 | 1472.58 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S07.csv | SUSPECT | 20 | 20 | 2 | different spot | 26.45 | 20.95 | 2392.5 | 442 | 3791 | 6216 | No | No | 1216.14 | 1885.05 | -130.29 | 26.46 | 568.35 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S08.csv | SUSPECT | 20 | 20 | 2 | different spot | 69.41 | 23.58 | 4088.5 | 676.5 | 5367.5 | 9514 | No | No | 1794.14 | 2650.3 | -179.11 | 69.48 | 1658.38 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S09.csv | SUSPECT | 20 | 20 | 2 | different spot | 83.93 | 26.07 | 5492 | 935.5 | 6771 | 11871 | No | No | 2416.57 | 3445.8 | -230.77 | 84.01 | 2217.58 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S10.csv | SUSPECT | 20 | 20 | 2 | different spot | 62.39 | 23.68 | 4078 | 739 | 5357 | 9434.5 | No | No | 2002.68 | 2988.72 | -204.34 | 62.49 | 1505.64 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S11.csv | SUSPECT | 20 | 20 | 2 | different spot | 41.56 | 23.7 | 3844.5 | 697 | 5199.5 | 8014 | No | No | 1961.05 | 2976.98 | -208.39 | 41.61 | 1002.27 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No
Unconfirmed_LabelledLTA_Bulk_20260209_P020pct_T20s_A02_S12.csv | SUSPECT | 20 | 20 | 2 | different spot | 96.6 | 28.82 | 6586.5 | 1015.5 | 7865.5 | 14278.5 | No | No | 2703.83 | 3956.31 | -267.87 | 96.71 | 2814.26 | 0 |  | Contaminated / misidentified | Polystyrene-like pattern (620, 1001, 1031, 1155, 1602 cm-1); labelled LTA but does not match LTA reference. Retained for review. | No

## 23_Preprocessing_Sensitivity
Sheet 23 — Sensitivity of marker heights to smoothing and baseline parameters (reference spectra)
SG 11/3 lowers the apex of the narrow NAM bands by ~16 % (positions within 1 pixel); LTA bands by ~3 %. ALS settings change NAM heights by < ±12 % but LTA heights by −40 % to +25 % (strong fluorescence). Absolute intensities are therefore pipeline-specific.
biomarker | Savitzky–Golay setting | Marker height relative to unsmoothed | Max peak-position shift (cm-1)
LTA | SG11/3 | 0.969 | 4.749
LTA | SG15/3 | 0.955 | 5.06
LTA | SG21/3 | 0.938 | 6.748
LTA | SG7/3 | 0.974 | 6.33
LTA | none | 1 | 0
NAM | SG11/3 | 0.844 | 1.679
NAM | SG15/3 | 0.741 | 1.679
NAM | SG21/3 | 0.613 | 4.711
NAM | SG7/3 | 0.942 | 1.679
NAM | none | 1 | 0
biomarker | ALS setting | Marker height / default (mean) | min | max
LTA | lambda=1e+04, p=0.001 | 1.046 | 1.024 | 1.058
LTA | lambda=1e+04, p=0.01 | 0.716 | 0.602 | 0.93
LTA | lambda=1e+04, p=0.05 | 0.603 | 0.494 | 0.764
LTA | lambda=1e+05, p=0.001 | 1.21 | 1.017 | 1.46
LTA | lambda=1e+05, p=0.01 | 1 | 1 | 1
LTA | lambda=1e+05, p=0.05 | 0.771 | 0.697 | 0.92
LTA | lambda=1e+06, p=0.001 | 1.253 | 0.993 | 1.73
LTA | lambda=1e+06, p=0.01 | 1.148 | 0.97 | 1.37
LTA | lambda=1e+06, p=0.05 | 0.991 | 0.947 | 1.054
NAM | lambda=1e+04, p=0.001 | 1.028 | 1.016 | 1.07
NAM | lambda=1e+04, p=0.01 | 0.942 | 0.833 | 0.99
NAM | lambda=1e+04, p=0.05 | 0.884 | 0.699 | 0.955
NAM | lambda=1e+05, p=0.001 | 1.044 | 1.028 | 1.088
NAM | lambda=1e+05, p=0.01 | 1 | 1 | 1
NAM | lambda=1e+05, p=0.05 | 0.92 | 0.846 | 0.957
NAM | lambda=1e+06, p=0.001 | 1.053 | 0.987 | 1.14
NAM | lambda=1e+06, p=0.01 | 1.012 | 0.965 | 1.053
NAM | lambda=1e+06, p=0.05 | 0.955 | 0.905 | 0.985
biomarker | SG window (order 3) | Residual noise ratio (smoothed/unsmoothed)
LTA | 7 | 0.39
LTA | 11 | 0.238
LTA | 15 | 0.168
LTA | 21 | 0.147
NAM | 7 | 0.361
NAM | 11 | 0.26
NAM | 15 | 0.215
NAM | 21 | 0.193

## 24_v3_Change_Log
Sheet 24 — Pipeline v3 corrections and their effect on reported values
Item | v2.1 | v3 | Reason
Spike removal | MAD z > 6 on first difference; 7-pt median replacement | Same criterion + width guard (≤ 2 px) + linear interpolation | v2.1 clipped the apex of sharp NAM bands (flagged 5.2 points/spectrum at 460–960 cm-1); marker heights were under-estimated by 11–12 % on average
Hot pixels | Not handled | Fixed CCD pixels 640 (1095.7 cm-1) and 700 (1189.5 cm-1) interpolated | Same pixel flagged in ≥ 59 spectra on 6 dates and visible in dark frames (instrument artefact)
Noise estimate | SD of (despiked − smoothed) | Robust SD (1.4826 × MAD) of the same residual | Sharp peaks inflated the SD; robust estimate reflects detector noise
NAM strongest-marker intensity | 820 ± 42 counts | 917 ± 37 counts | Result of the spike-removal fix
NAM SNR | 31.7 ± 1.4 | 51.1 ± 4.7 | Result of robust noise + spike fix
LTA strongest-marker intensity | 1103 ± 94 counts | 1110 ± 98 counts | Broad LTA bands barely affected
LTA SNR | 27.6 ± 0.7 | 28.5 ± 1.0
NAM marker set | 541, 829, 870, 929, 955, 1023, 1127 | unchanged | Re-verified with v3 criteria
NAM supporting peaks | 457, 487, 1249 | 457, 487, 1249, 1589 | 1589 meets the supporting rule (100 % ref; 91 % overall detection)

## 25_Acquisition_Log_Template
Sheet 25 — Acquisition log template for re-acquisition (fill one row per spectrum)
File name | Date | Biomarker | Power (%) | Integration (s) | Accumulations | Spot/replicate | Glass slide info | Spot position | Sample preparation | Objective | Operator | Saturation check | Notes
NAM_Bulk_YYYYMMDD_P090pct_T30s_A05_S01.csv | YYYY-MM-DD | NAM | 90 | 30 | 5 | S01 | slide ID / lot | spot position (x,y) or description | sample prep (e.g. powder pressed flat, amount) | objective (write it down) | operator | raw max 400–1800; saturation Y/N | notes
Glass_Blank_YYYYMMDD_P090pct_T30s_A05_S01.csv | YYYY-MM-DD | Glass blank | 90 | 30 | 5 | S01 | same slide type/lot | empty area | — | same objective |  |  | matched to NAM set

## 26_NAM_20260924_Inventory
library_file | category | power | original_folder | original_file | md5 | copy_verified | acq_time | note | identity_note
Glass_Blank_20260924_P060pct_T30s_A05_S01.csv | Glass_Blank | 60 | bulk-lps-24-09-26/60 power | glass-empty.csv | 82e29d5897437c23441899a46e5a4ddc | TRUE | 2026-09-24 03:10:51 |  | Empty glass slide blank
Glass_Blank_20260924_P060pct_T30s_A05_S02.csv | Glass_Blank | 60 | bulk-lps-24-09-26/60 power | glass-empty-2.csv | aaa31e432443372b9311207874b87dea | TRUE | 2026-09-24 03:13:57 |  | Empty glass slide blank
Glass_Blank_20260924_P060pct_T30s_A05_S03.csv | Glass_Blank | 60 | bulk-lps-24-09-26/60 power | glass-empty-3.csv | f0319da233ce3c49da7ee3bf6325cfb5 | TRUE | 2026-09-24 03:16:58 |  | Empty glass slide blank
Glass_Blank_20260924_P060pct_T30s_A05_S04.csv | Glass_Blank | 60 | bulk-lps-24-09-26/60 power | glass-empty-4.csv | 964f75ecc8bc1b63ff0bacb2df919a5e | TRUE | 2026-09-24 03:20:16 |  | Empty glass slide blank
NAM_Bulk_20260924_P060pct_T30s_A05_S01.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot1.csv | 069459b310fb4316aeb85991a78b8b9c | TRUE | 2026-09-24 03:28:47 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S02.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot2.csv | dcc76ab2ac9e10044705641f0fa3a52c | TRUE | 2026-09-24 03:32:21 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S03.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot3.csv | c7cce759cde8c7f89cbd47a5024087a7 | TRUE | 2026-09-24 03:43:32 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S04.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot4.csv | a892abae96012b7de732a861e1a80b2d | TRUE | 2026-09-24 03:48:35 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S05.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot5.csv | 2635788967c6c15bf58413cca07acccc | TRUE | 2026-09-24 03:51:54 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S06.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot6.csv | 4ac2638e7eb131e0650ec69f39a6990b | TRUE | 2026-09-24 03:55:56 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S07.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot7.csv | f5df20b3a0dcf08418f7e9d582885a20 | TRUE | 2026-09-24 03:59:27 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S08.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot8.csv | 6663b14b04d27e83e965b08b4f3e6660 | TRUE | 2026-09-24 04:02:50 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S09.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot9.csv | ce14922e55941c52ad9bc787dc40f392 | TRUE | 2026-09-24 04:05:55 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S10.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot10.csv | bb0264b39f87a4f4706448b8ff3c264b | TRUE | 2026-09-24 04:09:34 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P060pct_T30s_A05_S11.csv | NAM_Bulk | 60 | bulk-lps-24-09-26/60 power | spot11.csv | a2fecc7b017118727d8ee6b656e484cd | TRUE | 2026-09-24 04:12:59 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
Glass_Blank_20260924_P070pct_T30s_A05_S01.csv | Glass_Blank | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | glass1.csv | 68a582c05bf64896668a1f9b61eb1b7a | TRUE | 2026-09-24 04:28:49 |  | Empty glass slide blank
Glass_Blank_20260924_P070pct_T30s_A05_S02.csv | Glass_Blank | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | glass2.csv | 70ebf58c69991316cc4be6a67a0eea44 | TRUE | 2026-09-24 04:31:44 |  | Empty glass slide blank
Glass_Blank_20260924_P070pct_T30s_A05_S03.csv | Glass_Blank | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | glass3.csv | b96e9f567d92514e5de5951135e7d26e | TRUE | 2026-09-24 04:55:49 |  | Empty glass slide blank
NAM_Bulk_20260924_P070pct_T30s_A05_S01.csv | NAM_Bulk | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | spot1.csv | 7ba0afc3d088255bf6127e0e1ef6b2a5 | TRUE | 2026-09-24 05:09:27 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P070pct_T30s_A05_S02.csv | NAM_Bulk | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | spot2.csv | 65ef35aa6e5cf5e68fe17b964b73a955 | TRUE | 2026-09-24 05:14:00 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P070pct_T30s_A05_S03.csv | NAM_Bulk | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | spot3.csv | bd77fdc741002f15da62f2a3578253ef | TRUE | 2026-09-24 05:18:21 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P070pct_T30s_A05_S04.csv | NAM_Bulk | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | spot4.csv | 11c165db95b47bb6551511fd4bcd42b7 | TRUE | 2026-09-24 06:38:09 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P070pct_T30s_A05_S05.csv | NAM_Bulk | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | spot5.csv | aac39fa2997c93d535a797a0ea07682a | TRUE | 2026-09-24 06:47:00 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P070pct_T30s_A05_S06.csv | NAM_Bulk | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | spot6.csv | f1100bcf2bda972fdeeb93fa75c987fc | TRUE | 2026-09-24 08:05:36 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P070pct_T30s_A05_S07.csv | NAM_Bulk | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | spot7.csv | b2b41a8ec975973c77d57db5962a4c99 | TRUE | 2026-09-24 08:09:14 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P070pct_T30s_A05_S08.csv | NAM_Bulk | 70 | bulk-lps-24-09-26/lps-diffrent power-70 | spot8.csv | 1383d096aa00d5b9f2d278ec0e457fab | TRUE | 2026-09-24 08:23:16 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S01.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_1.csv | 3289bdf183b43d6d5c13f49bf233e7a2 | TRUE | 2026-09-24 10:38:25 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S02.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_2.csv | da446156e38feb82ec78f81c34c50cf6 | TRUE | 2026-09-24 10:45:06 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S03.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_3.csv | 06e9b375d21e108131a3a08af482c0fb | TRUE | 2026-09-24 10:48:52 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S04.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_4.csv | f894d24084f0c659743ddbedfcf54c38 | TRUE | 2026-09-24 10:53:54 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S05.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_5.csv | af43eb64c5921131d0cf326ea7ad2dc6 | TRUE | 2026-09-24 10:57:15 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S06.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_6.csv | 2eb627a5f351d4d925dae32e868e7894 | TRUE | 2026-09-24 11:01:11 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S07.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_7.csv | 3554587907ce0a6483c558877078636e | TRUE | 2026-09-24 11:07:24 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S08.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_8.csv | 3055dcd8a4d1ebdc63a7ddfa7c27dafb | TRUE | 2026-09-24 11:11:16 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P080pct_T30s_A05_S09.csv | NAM_Bulk | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | SP_9.csv | 861d8b0332ae06bc64131c33d5871d34 | TRUE | 2026-09-24 11:17:57 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
Glass_Blank_20260924_P080pct_T30s_A05_S01.csv | Glass_Blank | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | glass1.csv | 5ab20db1f455720e3dc464c55256cbdf | TRUE | 2026-09-24 11:34:12 |  | Empty glass slide blank
Glass_Blank_20260924_P080pct_T30s_A05_S02.csv | Glass_Blank | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | glass2.csv | c3c403c1de45cf5d1aa503595893a8dd | TRUE | 2026-09-24 11:37:02 |  | Empty glass slide blank
Glass_Blank_20260924_P080pct_T30s_A05_S03.csv | Glass_Blank | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | glass3.csv | 6b81148148a062190c0b154eb185c1ea | TRUE | 2026-09-24 11:40:03 |  | Empty glass slide blank
Glass_Blank_20260924_P080pct_T30s_A05_S04.csv | Glass_Blank | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | class4.csv | 0295cef39ad933ba33c633d7f45ecccf | TRUE | 2026-09-24 11:44:19 | Original name "class4.csv": classified as empty-glass blank from the data (saturated glass band, no sample bands); operator confirmation pending | Empty glass slide blank
Glass_Blank_20260924_P080pct_T30s_A05_S05.csv | Glass_Blank | 80 | bulk-lps-24-09-26/lps-diffrent-80 power | class5.csv | 4a1f7e833a73d98d5a436f8baf10c1ae | TRUE | 2026-09-24 11:47:12 | Original name "class5.csv": classified as empty-glass blank from the data (saturated glass band, no sample bands); operator confirmation pending | Empty glass slide blank
Glass_Blank_20260924_P090pct_T30s_A05_S01.csv | Glass_Blank | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | glass1.csv | 3015d2a1aa8169168bdf4f4afcae85ba | TRUE | 2026-09-24 11:58:00 |  | Empty glass slide blank
Glass_Blank_20260924_P090pct_T30s_A05_S02.csv | Glass_Blank | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | glass2.csv | f44a0f5bfe0d0bb64388f4743e7907ee | TRUE | 2026-09-24 12:04:49 |  | Empty glass slide blank
Glass_Blank_20260924_P090pct_T30s_A05_S03.csv | Glass_Blank | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | glass3.csv | 9dc61ac268faf3d34549d1ff0af650c6 | TRUE | 2026-09-24 12:08:05 |  | Empty glass slide blank
Glass_Blank_20260924_P090pct_T30s_A05_S04.csv | Glass_Blank | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | glass4.csv | b918896fb53ef81a2178bc4000abf171 | TRUE | 2026-09-24 12:11:44 |  | Empty glass slide blank
NAM_Bulk_20260924_P090pct_T30s_A05_S01.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot1.csv | 026186e4222d085c6e5a1b5c0349e04b | TRUE | 2026-09-24 12:15:13 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S02.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot2.csv | a22bf859f374a9d7867a023c1d2b7d23 | TRUE | 2026-09-24 12:18:35 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S03.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot3.csv | 4ad7729cfba49d6704999f229893de1c | TRUE | 2026-09-24 12:22:03 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S04.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot4.csv | 5a9d6cda5396eae15ba2b6818b5eebe9 | TRUE | 2026-09-24 12:25:16 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S05.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot5.csv | 401c5df29dd1ea2bfa956268ffcda876 | TRUE | 2026-09-24 12:28:42 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S06.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot6.csv | 684c97b0510fe5b76cf9d93ddb6315fd | TRUE | 2026-09-24 12:31:52 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S07.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot7.csv | 7fa3c58bb433d3cff2aa8da0400534ae | TRUE | 2026-09-24 12:35:25 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S08.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot8.csv | 50bfda9dce0895ce76ffa4035d018b95 | TRUE | 2026-09-24 12:38:49 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S09.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot9.csv | 42da349749b43e052ebec2a3c65a95de | TRUE | 2026-09-24 12:42:12 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S10.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot10.csv | ab3e0c6a75d5e90dce32839f3fcac930 | TRUE | 2026-09-24 12:50:42 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S11.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot11.csv | 9965f0044bb5da717ce306fd4199b13c | TRUE | 2026-09-24 12:54:32 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S12.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot12.csv | 1e9a28388b50cae9fc94e11c9644e832 | TRUE | 2026-09-24 12:57:43 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference
NAM_Bulk_20260924_P090pct_T30s_A05_S13.csv | NAM_Bulk | 90 | bulk-lps-24-09-26/lps-diffrent-90-power | spot13.csv | 9ce5ef39715057de2e15ff2d54093414 | TRUE | 2026-09-24 13:01:33 |  | Folder named "lps" in error; operator checked bottle 27 Sep 2026: NAM; spectra match NAM reference

## 27_NAM_20260924_QC
library_file | category | power | integration_s | accumulations | dark_median | ripple_ds | raw_max_400_1800 | n_sat_400_1800 | sat_from | sat_to | noise | h929_counts | snr929 | r_NAM_ref | QC | QC_reason
Glass_Blank_20260924_P060pct_T30s_A05_S01.csv | Glass_Blank | 60 | 30 | 5 | 1431 | 96.482 | 65535 | 90 | 1323.47 | 1454.672 | 41.928 | 914.541 | 21.812 | -0.009 | BLANK-SATURATED | Detector saturated 1323-1455 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P060pct_T30s_A05_S02.csv | Glass_Blank | 60 | 30 | 5 | 1431 | 98.513 | 65535 | 88 | 1323.47 | 1451.77 | 41.215 | 869.663 | 21.101 | 0.013 | BLANK-SATURATED | Detector saturated 1323-1452 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P060pct_T30s_A05_S03.csv | Glass_Blank | 60 | 30 | 5 | 1431 | 98.814 | 65535 | 82 | 1324.968 | 1444.505 | 40.394 | 808.516 | 20.016 | 0 | BLANK-SATURATED | Detector saturated 1325-1445 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P060pct_T30s_A05_S04.csv | Glass_Blank | 60 | 30 | 5 | 1431 | 95.128 | 65535 | 82 | 1324.968 | 1444.505 | 42.848 | 807.444 | 18.844 | -0.007 | BLANK-SATURATED | Detector saturated 1325-1445 cm-1 (glass band); usable only outside this range
NAM_Bulk_20260924_P060pct_T30s_A05_S01.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 15.729 | 5164.2 | 0 |  |  | 16.004 | 568.664 | 35.532 | 0.985 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S02.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 15.387 | 4894.8 | 0 |  |  | 16.754 | 567.737 | 33.887 | 0.99 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S03.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 13.918 | 4847.8 | 0 |  |  | 14.847 | 499.533 | 33.646 | 0.986 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S04.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 15.371 | 5040.8 | 0 |  |  | 16.763 | 499.253 | 29.782 | 0.982 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S05.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 15.691 | 4908.6 | 0 |  |  | 16.443 | 468.497 | 28.493 | 0.983 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S06.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 16.696 | 4939.8 | 0 |  |  | 15.477 | 501.133 | 32.379 | 0.984 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S07.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 16.167 | 4863 | 0 |  |  | 15.508 | 368.938 | 23.791 | 0.935 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S08.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 16.389 | 4908.2 | 0 |  |  | 15.548 | 440.293 | 28.318 | 0.974 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S09.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 16.704 | 5055.2 | 0 |  |  | 15.387 | 540.589 | 35.134 | 0.985 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S10.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 15.087 | 4760 | 0 |  |  | 14.42 | 352.949 | 24.477 | 0.967 | PASS
NAM_Bulk_20260924_P060pct_T30s_A05_S11.csv | NAM_Bulk | 60 | 30 | 5 | 1431 | 17.152 | 4831 | 0 |  |  | 17.121 | 399.068 | 23.308 | 0.967 | PASS
Glass_Blank_20260924_P070pct_T30s_A05_S01.csv | Glass_Blank | 70 | 30 | 5 | 1418.1 | 120.624 | 65535 | 123 | 1311.468 | 1490.774 | 44.325 | 944.929 | 21.318 | -0.058 | BLANK-SATURATED | Detector saturated 1311-1491 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P070pct_T30s_A05_S02.csv | Glass_Blank | 70 | 30 | 5 | 1418.1 | 112.992 | 65535 | 116 | 1312.97 | 1482.139 | 45.027 | 879.774 | 19.539 | -0.046 | BLANK-SATURATED | Detector saturated 1313-1482 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P070pct_T30s_A05_S03.csv | Glass_Blank | 70 | 30 | 5 | 1418.1 | 116.408 | 65535 | 128 | 1311.468 | 1497.956 | 46.565 | 964.734 | 20.718 | -0.063 | BLANK-SATURATED | Detector saturated 1311-1498 cm-1 (glass band); usable only outside this range
NAM_Bulk_20260924_P070pct_T30s_A05_S01.csv | NAM_Bulk | 70 | 30 | 5 | 1418.1 | 17.372 | 5411 | 0 |  |  | 17.079 | 499.435 | 29.243 | 0.976 | PASS
NAM_Bulk_20260924_P070pct_T30s_A05_S02.csv | NAM_Bulk | 70 | 30 | 5 | 1418.1 | 16.02 | 4737.2 | 0 |  |  | 16.022 | 571.247 | 35.654 | 0.991 | PASS
NAM_Bulk_20260924_P070pct_T30s_A05_S03.csv | NAM_Bulk | 70 | 30 | 5 | 1418.1 | 19.517 | 4506.6 | 0 |  |  | 17.694 | 644.314 | 36.413 | 0.993 | PASS
NAM_Bulk_20260924_P070pct_T30s_A05_S04.csv | NAM_Bulk | 70 | 30 | 5 | 1418.1 | 16.409 | 6003.2 | 0 |  |  | 16.1 | 520.23 | 32.312 | 0.99 | PASS
NAM_Bulk_20260924_P070pct_T30s_A05_S05.csv | NAM_Bulk | 70 | 30 | 5 | 1418.1 | 17.069 | 6145 | 0 |  |  | 16.551 | 608.073 | 36.739 | 0.988 | PASS
NAM_Bulk_20260924_P070pct_T30s_A05_S06.csv | NAM_Bulk | 70 | 30 | 5 | 1823.9 | 115.956 | 8458.8 | 0 |  |  | 47.335 | 834.426 | 17.628 | 0.767 | FAIL | Dark-reference mismatch ripple (residual 116 counts > 50)
NAM_Bulk_20260924_P070pct_T30s_A05_S07.csv | NAM_Bulk | 70 | 30 | 5 | 1823.9 | 160.153 | 8438.4 | 0 |  |  | 55.708 | 930.98 | 16.712 | 0.682 | FAIL | Dark-reference mismatch ripple (residual 160 counts > 50)
NAM_Bulk_20260924_P070pct_T30s_A05_S08.csv | NAM_Bulk | 70 | 30 | 5 | 1823.9 | 333.288 | 9031 | 0 |  |  | 89.799 | 981.506 | 10.93 | 0.293 | FAIL | Dark-reference mismatch ripple (residual 333 counts > 50)
NAM_Bulk_20260924_P080pct_T30s_A05_S01.csv | NAM_Bulk | 80 | 30 | 5 | 4043 | 40.794 | 10719.4 | 0 |  |  | 13.674 | 645.311 | 47.192 | 0.926 | PASS
NAM_Bulk_20260924_P080pct_T30s_A05_S02.csv | NAM_Bulk | 80 | 30 | 5 | 4043 | 196.672 | 9607 | 0 |  |  | 64.059 | 585.432 | 9.139 | 0.458 | FAIL | Dark-reference mismatch ripple (residual 197 counts > 50)
NAM_Bulk_20260924_P080pct_T30s_A05_S03.csv | NAM_Bulk | 80 | 30 | 5 | 4043 | 154.886 | 10627.8 | 0 |  |  | 58.123 | 658.63 | 11.332 | 0.621 | FAIL | Dark-reference mismatch ripple (residual 155 counts > 50)
NAM_Bulk_20260924_P080pct_T30s_A05_S04.csv | NAM_Bulk | 80 | 30 | 5 | 4043 | 169.298 | 10666 | 0 |  |  | 58.213 | 651.843 | 11.197 | 0.575 | FAIL | Dark-reference mismatch ripple (residual 169 counts > 50)
NAM_Bulk_20260924_P080pct_T30s_A05_S05.csv | NAM_Bulk | 80 | 30 | 5 | 4043 | 70.663 | 11571 | 0 |  |  | 45.177 | 861.748 | 19.075 | 0.977 | FAIL | Dark-reference mismatch ripple (residual 71 counts > 50)
NAM_Bulk_20260924_P080pct_T30s_A05_S06.csv | NAM_Bulk | 80 | 30 | 5 | 4426.9 | 24.971 | 13060 | 0 |  |  | 27.824 | 756.522 | 27.189 | 0.99 | PASS
NAM_Bulk_20260924_P080pct_T30s_A05_S07.csv | NAM_Bulk | 80 | 30 | 5 | 4426.9 | 117.24 | 11042.6 | 0 |  |  | 48.162 | 743.06 | 15.428 | 0.769 | FAIL | Dark-reference mismatch ripple (residual 117 counts > 50)
NAM_Bulk_20260924_P080pct_T30s_A05_S08.csv | NAM_Bulk | 80 | 30 | 5 | 3605.7 | 20.508 | 10264 | 0 |  |  | 23.327 | 665.762 | 28.54 | 0.991 | PASS
NAM_Bulk_20260924_P080pct_T30s_A05_S09.csv | NAM_Bulk | 80 | 30 | 5 | 4473.7 | 25.225 | 12981.8 | 0 |  |  | 26.353 | 803.361 | 30.485 | 0.994 | PASS
Glass_Blank_20260924_P080pct_T30s_A05_S01.csv | Glass_Blank | 80 | 30 | 5 | 1821.6 | 177.713 | 65535 | 449 | 1153.793 | 1798.853 | 70.027 | 1573.333 | 22.467 | 0.113 | BLANK-SATURATED | Detector saturated 1154-1799 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P080pct_T30s_A05_S02.csv | Glass_Blank | 80 | 30 | 5 | 1821.6 | 177.713 | 65535 | 450 | 1152.233 | 1798.853 | 67.994 | 1447.354 | 21.286 | 0.118 | BLANK-SATURATED | Detector saturated 1152-1799 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P080pct_T30s_A05_S03.csv | Glass_Blank | 80 | 30 | 5 | 1821.6 | 177.713 | 65535 | 450 | 1152.233 | 1798.853 | 70.925 | 1412.735 | 19.919 | 0.123 | BLANK-SATURATED | Detector saturated 1152-1799 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P080pct_T30s_A05_S04.csv | Glass_Blank | 80 | 30 | 5 | 1821.6 | 184.743 | 65535 | 248 | 1278.284 | 1637.521 | 53.273 | 1378.1 | 25.869 | -0.119 | BLANK-SATURATED | Detector saturated 1278-1638 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P080pct_T30s_A05_S05.csv | Glass_Blank | 80 | 30 | 5 | 1821.6 | 186.837 | 65535 | 250 | 1278.284 | 1638.906 | 56.574 | 1417.99 | 25.064 | -0.117 | BLANK-SATURATED | Detector saturated 1278-1639 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P090pct_T30s_A05_S01.csv | Glass_Blank | 90 | 30 | 5 | 1772.4 | 183.461 | 65535 | 273 | 1269.187 | 1662.368 | 55.147 | 1465.865 | 26.581 | -0.103 | BLANK-SATURATED | Detector saturated 1269-1662 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P090pct_T30s_A05_S02.csv | Glass_Blank | 90 | 30 | 5 | 1772.4 | 182.354 | 65535 | 265 | 1272.222 | 1654.104 | 53.064 | 1463.455 | 27.579 | -0.103 | BLANK-SATURATED | Detector saturated 1272-1654 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P090pct_T30s_A05_S03.csv | Glass_Blank | 90 | 30 | 5 | 1772.4 | 180.211 | 65535 | 253 | 1276.769 | 1641.674 | 57.539 | 1452.732 | 25.248 | -0.107 | BLANK-SATURATED | Detector saturated 1277-1642 cm-1 (glass band); usable only outside this range
Glass_Blank_20260924_P090pct_T30s_A05_S04.csv | Glass_Blank | 90 | 30 | 5 | 1772.4 | 186.549 | 65535 | 267 | 1272.222 | 1656.861 | 51.255 | 1430.523 | 27.91 | -0.104 | BLANK-SATURATED | Detector saturated 1272-1657 cm-1 (glass band); usable only outside this range
NAM_Bulk_20260924_P090pct_T30s_A05_S01.csv | NAM_Bulk | 90 | 30 | 5 | 1772.4 | 354.459 | 9592 | 0 |  |  | 93.993 | 1381.454 | 14.697 | 0.492 | FAIL | Dark-reference mismatch ripple (residual 354 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S02.csv | NAM_Bulk | 90 | 30 | 5 | 1772.4 | 397.131 | 10580.8 | 0 |  |  | 101.073 | 1300.701 | 12.869 | 0.38 | FAIL | Dark-reference mismatch ripple (residual 397 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S03.csv | NAM_Bulk | 90 | 30 | 5 | 1772.4 | 155.723 | 7889 | 0 |  |  | 51.953 | 890.223 | 17.135 | 0.685 | FAIL | Dark-reference mismatch ripple (residual 156 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S04.csv | NAM_Bulk | 90 | 30 | 5 | 1772.4 | 287.124 | 9322.8 | 0 |  |  | 73.541 | 1146.235 | 15.586 | 0.501 | FAIL | Dark-reference mismatch ripple (residual 287 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S05.csv | NAM_Bulk | 90 | 30 | 5 | 1772.4 | 307.952 | 9114 | 0 |  |  | 80.796 | 1142.122 | 14.136 | 0.45 | FAIL | Dark-reference mismatch ripple (residual 308 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S06.csv | NAM_Bulk | 90 | 30 | 5 | 1772.4 | 376.2 | 9895.2 | 0 |  |  | 92.443 | 1284.153 | 13.891 | 0.408 | FAIL | Dark-reference mismatch ripple (residual 376 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S07.csv | NAM_Bulk | 90 | 30 | 5 | 1772.4 | 279.198 | 8889.2 | 0 |  |  | 73.954 | 1140.348 | 15.42 | 0.504 | FAIL | Dark-reference mismatch ripple (residual 279 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S08.csv | NAM_Bulk | 90 | 30 | 5 | 1772.4 | 290.493 | 8994.8 | 0 |  |  | 72.426 | 988.217 | 13.644 | 0.414 | FAIL | Dark-reference mismatch ripple (residual 290 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S09.csv | NAM_Bulk | 90 | 30 | 5 | 3814.3 | 21.829 | 10458.8 | 0 |  |  | 25.788 | 708.644 | 27.479 | 0.994 | PASS
NAM_Bulk_20260924_P090pct_T30s_A05_S10.csv | NAM_Bulk | 90 | 30 | 5 | 3814.3 | 176.726 | 8949 | 0 |  |  | 58.672 | 998.446 | 17.017 | 0.7 | FAIL | Dark-reference mismatch ripple (residual 177 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S11.csv | NAM_Bulk | 90 | 30 | 5 | 3814.3 | 60.323 | 10271.6 | 0 |  |  | 34.132 | 720.057 | 21.096 | 0.916 | FAIL | Dark-reference mismatch ripple (residual 60 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S12.csv | NAM_Bulk | 90 | 30 | 5 | 3814.3 | 151.926 | 8814 | 0 |  |  | 51.672 | 789.138 | 15.272 | 0.673 | FAIL | Dark-reference mismatch ripple (residual 152 counts > 50)
NAM_Bulk_20260924_P090pct_T30s_A05_S13.csv | NAM_Bulk | 90 | 30 | 5 | 3814.3 | 307.689 | 8700.2 | 0 |  |  | 80.895 | 916.527 | 11.33 | 0.433 | FAIL | Dark-reference mismatch ripple (residual 308 counts > 50)

## 28_NAM_20260924_SetStats
set | n | r_mean_vs_REFmean | pairwise_r_mean | pairwise_r_min | r_mean_median | frac_within_1SD | noise_mean
REF_90_30May_all15 | 15 | 1 | 0.9943 | 0.9751 | 0.9992 | 1 | 18.0717
REF_90_30May_indep5 | 5 | 0.9985 | 0.9893 | 0.9751 | 0.9903 | 0.9989 | 18.8938
NEW_60_24Sep | 11 | 0.9824 | 0.9866 | 0.9509 | 0.9956 | 1 | 15.8429
NEW_70_24Sep | 5 | 0.991 | 0.9915 | 0.983 | 0.9989 | 1 | 16.6892
NEW_80_24Sep | 4 | 0.9904 | 0.9584 | 0.9226 | 0.9969 | 1 | 22.7946
NEW_90_24Sep | 1 | 0.994 |  |  |  |  | 25.7883
NEW_all_pass_24Sep | 21 | 0.9915 | 0.9755 | 0.8545 | 0.9912 | 1 | 17.8421

## 29_NAM_20260924_Old_vs_New
set | role | ref_pos | glass_zone | ref_pos_nativeaxis | new_pos | dpos | new_detected | ref_rel929 | new_rel929 | rel_change_pct
NEW_60_24Sep | Primary marker | 541 | FALSE | 541.022 | 539.236 | -1.786 | 11/11 | 0.778 | 0.765 | -1.674
NEW_60_24Sep | Primary marker | 829 | FALSE | 828.481 | 828.481 | 0 | 11/11 | 0.682 | 0.675 | -1.055
NEW_60_24Sep | Primary marker | 870.53 | FALSE | 870.262 | 870.262 | 0 | 11/11 | 0.957 | 0.967 | 1.102
NEW_60_24Sep | Primary marker | 929 | FALSE | 929.317 | 928.113 | -1.204 | 11/11 | 1 | 1 | 0
NEW_60_24Sep | Primary marker | 955 | FALSE | 954.53 | 954.313 | -0.218 | 11/11 | 0.939 | 0.894 | -4.866
NEW_60_24Sep | Primary marker | 1022.73 | FALSE | 1022.358 | 1022.066 | -0.292 | 11/11 | 0.822 | 0.752 | -8.539
NEW_60_24Sep | Primary marker | 1127 | FALSE | 1127.196 | 1125.626 | -1.57 | 11/11 | 0.633 | 0.567 | -10.339
NEW_60_24Sep | Supporting peak | 457 | FALSE | 456.473 | 456.352 | -0.121 | 11/11 | 0.598 | 0.56 | -6.219
NEW_60_24Sep | Supporting peak | 487 | FALSE | 487.149 | 486 | -1.149 | 11/11 | 0.434 | 0.402 | -7.328
NEW_60_24Sep | Supporting peak | 1248.67 | FALSE | 1248.902 | 1248.024 | -0.878 | 11/11 | 0.316 | 0.332 | 5.119
NEW_60_24Sep | Uncertain | 583.8 | FALSE | 583.687 | 583.687 | 0 | 11/11 | 0.302 | 0.302 | 0.035
NEW_60_24Sep | Uncertain | 1058.93 | FALSE | 1059.074 | 1058.021 | -1.053 | 11/11 | 0.31 | 0.288 | -7.048
NEW_60_24Sep | Uncertain | 1155 | FALSE | 1155.352 | 1154.927 | -0.425 | 11/11 | 0.399 | 0.333 | -16.578
NEW_60_24Sep | Supporting peak | 1589.47 | FALSE | 1589.401 | 1588.874 | -0.527 | 11/11 | 0.286 | 0.303 | 5.948
NEW_60_24Sep | Uncertain | 1344.87 | TRUE | 1344.687 | 1345.338 | 0.65 | 11/11 | 0.681 | 0.852 | 25.085
NEW_60_24Sep | Uncertain | 1392.47 | TRUE | 1392.497 | 1392.479 | -0.018 | 11/11 | 0.66 | 0.802 | 21.527
NEW_60_24Sep | Uncertain | 1450 | TRUE | 1450.124 | 1449.011 | -1.114 | 10/11 | 0.543 | 0.555 | 2.254
NEW_70_24Sep | Primary marker | 541 | FALSE | 541.022 | 539.236 | -1.786 | 5/5 | 0.778 | 0.768 | -1.25
NEW_70_24Sep | Primary marker | 829 | FALSE | 828.481 | 828.481 | 0 | 5/5 | 0.682 | 0.69 | 1.185
NEW_70_24Sep | Primary marker | 870.53 | FALSE | 870.262 | 870.262 | 0 | 5/5 | 0.957 | 0.966 | 0.957
NEW_70_24Sep | Primary marker | 929 | FALSE | 929.317 | 928.113 | -1.204 | 5/5 | 1 | 1 | 0
NEW_70_24Sep | Primary marker | 955 | FALSE | 954.53 | 954.313 | -0.218 | 5/5 | 0.939 | 0.939 | -0.088
NEW_70_24Sep | Primary marker | 1022.73 | FALSE | 1022.358 | 1022.358 | 0 | 5/5 | 0.822 | 0.784 | -4.653
NEW_70_24Sep | Primary marker | 1127 | FALSE | 1127.196 | 1125.626 | -1.57 | 5/5 | 0.633 | 0.607 | -4.095
NEW_70_24Sep | Supporting peak | 457 | FALSE | 456.473 | 456.352 | -0.121 | 5/5 | 0.598 | 0.583 | -2.354
NEW_70_24Sep | Supporting peak | 487 | FALSE | 487.149 | 486.788 | -0.361 | 5/5 | 0.434 | 0.423 | -2.582
NEW_70_24Sep | Supporting peak | 1248.67 | FALSE | 1248.902 | 1248.19 | -0.712 | 5/5 | 0.316 | 0.331 | 4.594
NEW_70_24Sep | Uncertain | 583.8 | FALSE | 583.687 | 583.687 | 0 | 5/5 | 0.302 | 0.299 | -0.87
NEW_70_24Sep | Uncertain | 1058.93 | FALSE | 1059.074 | 1058.224 | -0.85 | 5/5 | 0.31 | 0.29 | -6.577
NEW_70_24Sep | Uncertain | 1155 | FALSE | 1155.352 | 1154.729 | -0.624 | 5/5 | 0.399 | 0.377 | -5.527
NEW_70_24Sep | Supporting peak | 1589.47 | FALSE | 1589.401 | 1589.027 | -0.374 | 5/5 | 0.286 | 0.295 | 3.26
NEW_70_24Sep | Uncertain | 1344.87 | TRUE | 1344.687 | 1345.582 | 0.894 | 5/5 | 0.681 | 0.759 | 11.372
NEW_70_24Sep | Uncertain | 1392.47 | TRUE | 1392.497 | 1390.925 | -1.572 | 5/5 | 0.66 | 0.706 | 7.048
NEW_70_24Sep | Uncertain | 1450 | TRUE | 1450.124 | 1450.318 | 0.194 | 5/5 | 0.543 | 0.568 | 4.512
NEW_80_24Sep | Primary marker | 541 | FALSE | 541.022 | 540.575 | -0.447 | 4/4 | 0.778 | 0.795 | 2.123
NEW_80_24Sep | Primary marker | 829 | FALSE | 828.481 | 828.9 | 0.42 | 4/4 | 0.682 | 0.687 | 0.657
NEW_80_24Sep | Primary marker | 870.53 | FALSE | 870.262 | 870.678 | 0.416 | 4/4 | 0.957 | 0.971 | 1.551
NEW_80_24Sep | Primary marker | 929 | FALSE | 929.317 | 929.345 | 0.027 | 4/4 | 1 | 1 | 0
NEW_80_24Sep | Primary marker | 955 | FALSE | 954.53 | 955.129 | 0.599 | 4/4 | 0.939 | 0.915 | -2.616
NEW_80_24Sep | Primary marker | 1022.73 | FALSE | 1022.358 | 1022.76 | 0.402 | 4/4 | 0.822 | 0.85 | 3.454
NEW_80_24Sep | Primary marker | 1127 | FALSE | 1127.196 | 1126.804 | -0.392 | 4/4 | 0.633 | 0.619 | -2.156
NEW_80_24Sep | Supporting peak | 457 | FALSE | 456.473 | 457.26 | 0.787 | 4/4 | 0.598 | 0.623 | 4.323
NEW_80_24Sep | Supporting peak | 487 | FALSE | 487.149 | 487.149 | 0 | 4/4 | 0.434 | 0.515 | 18.69
NEW_80_24Sep | Supporting peak | 1248.67 | FALSE | 1248.902 | 1248.266 | -0.635 | 4/4 | 0.316 | 0.342 | 8.287
NEW_80_24Sep | Uncertain | 583.8 | FALSE | 583.687 | 584.129 | 0.442 | 4/4 | 0.302 | 0.404 | 33.886
NEW_80_24Sep | Uncertain | 1058.93 | FALSE | 1059.074 | 1058.782 | -0.292 | 4/4 | 0.31 | 0.322 | 3.767
NEW_80_24Sep | Uncertain | 1155 | FALSE | 1155.352 | 1155.352 | 0 | 4/4 | 0.399 | 0.384 | -3.89
NEW_80_24Sep | Supporting peak | 1589.47 | FALSE | 1589.401 | 1589.798 | 0.397 | 4/4 | 0.286 | 0.348 | 21.69
NEW_80_24Sep | Uncertain | 1344.87 | TRUE | 1344.687 | 1343.644 | -1.043 | 4/4 | 0.681 | 0.602 | -11.548
NEW_80_24Sep | Uncertain | 1392.47 | TRUE | 1392.497 | 1392.914 | 0.417 | 4/4 | 0.66 | 0.56 | -15.153
NEW_80_24Sep | Uncertain | 1450 | TRUE | 1450.124 | 1449.228 | -0.896 | 4/4 | 0.543 | 0.504 | -7.251
NEW_90_24Sep | Primary marker | 541 | FALSE | 541.022 | 541.022 | 0 | 1/1 | 0.778 | 0.802 | 3.041
NEW_90_24Sep | Primary marker | 829 | FALSE | 828.481 | 830.159 | 1.679 | 1/1 | 0.682 | 0.72 | 5.582
NEW_90_24Sep | Primary marker | 870.53 | FALSE | 870.262 | 870.262 | 0 | 1/1 | 0.957 | 0.977 | 2.142
NEW_90_24Sep | Primary marker | 929 | FALSE | 929.317 | 929.755 | 0.438 | 1/1 | 1 | 1 | 0
NEW_90_24Sep | Primary marker | 955 | FALSE | 954.53 | 955.945 | 1.415 | 1/1 | 0.939 | 0.941 | 0.215
NEW_90_24Sep | Primary marker | 1022.73 | FALSE | 1022.358 | 1022.358 | 0 | 1/1 | 0.822 | 0.825 | 0.362
NEW_90_24Sep | Primary marker | 1127 | FALSE | 1127.196 | 1127.196 | 0 | 1/1 | 0.633 | 0.659 | 4.087
NEW_90_24Sep | Supporting peak | 457 | FALSE | 456.473 | 456.352 | -0.121 | 1/1 | 0.598 | 0.558 | -6.659
NEW_90_24Sep | Supporting peak | 487 | FALSE | 487.149 | 487.149 | 0 | 1/1 | 0.434 | 0.436 | 0.419
NEW_90_24Sep | Supporting peak | 1248.67 | FALSE | 1248.902 | 1249.41 | 0.508 | 1/1 | 0.316 | 0.327 | 3.408
NEW_90_24Sep | Uncertain | 583.8 | FALSE | 583.687 | 583.687 | 0 | 1/1 | 0.302 | 0.284 | -5.978
NEW_90_24Sep | Uncertain | 1058.93 | FALSE | 1059.074 | 1059.18 | 0.106 | 1/1 | 0.31 | 0.314 | 1.359
NEW_90_24Sep | Uncertain | 1155 | FALSE | 1155.352 | 1155.352 | 0 | 1/1 | 0.399 | 0.423 | 6.053
NEW_90_24Sep | Supporting peak | 1589.47 | FALSE | 1589.401 | 1588.747 | -0.654 | 1/1 | 0.286 | 0.301 | 5.424
NEW_90_24Sep | Uncertain | 1344.87 | TRUE | 1344.687 | 1344.389 | -0.298 | 1/1 | 0.681 | 0.645 | -5.321
NEW_90_24Sep | Uncertain | 1392.47 | TRUE | 1392.497 | 1393.282 | 0.786 | 1/1 | 0.66 | 0.634 | -3.94
NEW_90_24Sep | Uncertain | 1450 | TRUE | 1450.124 | 1450.318 | 0.194 | 1/1 | 0.543 | 0.568 | 4.538
NEW_all_pass_24Sep | Primary marker | 541 | FALSE | 541.022 | 539.576 | -1.446 | 21/21 | 0.778 | 0.773 | -0.625
NEW_all_pass_24Sep | Primary marker | 829 | FALSE | 828.481 | 828.641 | 0.16 | 21/21 | 0.682 | 0.683 | 0.12
NEW_all_pass_24Sep | Primary marker | 870.53 | FALSE | 870.262 | 870.341 | 0.079 | 21/21 | 0.957 | 0.968 | 1.202
NEW_all_pass_24Sep | Primary marker | 929 | FALSE | 929.317 | 928.426 | -0.891 | 21/21 | 1 | 1 | 0
NEW_all_pass_24Sep | Primary marker | 955 | FALSE | 954.53 | 954.546 | 0.016 | 21/21 | 0.939 | 0.911 | -3.058
NEW_all_pass_24Sep | Primary marker | 1022.73 | FALSE | 1022.358 | 1022.282 | -0.077 | 21/21 | 0.822 | 0.782 | -4.905
NEW_all_pass_24Sep | Primary marker | 1127 | FALSE | 1127.196 | 1125.925 | -1.271 | 21/21 | 0.633 | 0.591 | -6.606
NEW_all_pass_24Sep | Supporting peak | 457 | FALSE | 456.473 | 456.525 | 0.052 | 21/21 | 0.598 | 0.578 | -3.312
NEW_all_pass_24Sep | Supporting peak | 487 | FALSE | 487.149 | 486.461 | -0.688 | 21/21 | 0.434 | 0.43 | -0.873
NEW_all_pass_24Sep | Supporting peak | 1248.67 | FALSE | 1248.902 | 1248.176 | -0.726 | 21/21 | 0.316 | 0.333 | 5.516
NEW_all_pass_24Sep | Uncertain | 583.8 | FALSE | 583.687 | 583.771 | 0.084 | 21/21 | 0.302 | 0.32 | 5.981
NEW_all_pass_24Sep | Uncertain | 1058.93 | FALSE | 1059.074 | 1058.269 | -0.805 | 21/21 | 0.31 | 0.296 | -4.475
NEW_all_pass_24Sep | Uncertain | 1155 | FALSE | 1155.352 | 1154.981 | -0.371 | 21/21 | 0.399 | 0.357 | -10.452
NEW_all_pass_24Sep | Supporting peak | 1589.47 | FALSE | 1589.401 | 1589.08 | -0.321 | 21/21 | 0.286 | 0.31 | 8.281
NEW_all_pass_24Sep | Uncertain | 1344.87 | TRUE | 1344.687 | 1345.028 | 0.341 | 21/21 | 0.681 | 0.772 | 13.395
NEW_all_pass_24Sep | Uncertain | 1392.47 | TRUE | 1392.497 | 1392.23 | -0.267 | 21/21 | 0.66 | 0.725 | 9.881
NEW_all_pass_24Sep | Uncertain | 1450 | TRUE | 1450.124 | 1449.446 | -0.678 | 20/21 | 0.543 | 0.549 | 1.032

## 30_NAM_20260924_Bands
set | role | ref_pos | glass_zone | detected | det_frac | pos_mean | pos_sd | height_mean | height_cv | snr_mean | rel929_mean | rel929_sd
REF_90_30May_all15 | Primary marker | 541 | FALSE | 15/15 | 1 | 541.022 | 0 | 715.425 | 4.551 | 39.873 | 0.778 | 0.012
REF_90_30May_all15 | Primary marker | 829 | FALSE | 15/15 | 1 | 828.481 | 0 | 627.325 | 4.856 | 34.982 | 0.682 | 0.011
REF_90_30May_all15 | Primary marker | 870.53 | FALSE | 15/15 | 1 | 870.262 | 0 | 879.633 | 4.278 | 49.016 | 0.957 | 0.011
REF_90_30May_all15 | Primary marker | 929 | FALSE | 15/15 | 1 | 929.317 | 0.752 | 919.439 | 3.917 | 51.24 | 1 | 0
REF_90_30May_all15 | Primary marker | 955 | FALSE | 15/15 | 1 | 954.53 | 0.574 | 864.027 | 5.128 | 48.182 | 0.939 | 0.017
REF_90_30May_all15 | Primary marker | 1022.73 | FALSE | 15/15 | 1 | 1022.358 | 0 | 756.306 | 6.314 | 42.203 | 0.822 | 0.026
REF_90_30May_all15 | Primary marker | 1127 | FALSE | 15/15 | 1 | 1127.196 | 0 | 582.096 | 5.125 | 32.46 | 0.633 | 0.014
REF_90_30May_all15 | Supporting peak | 457 | FALSE | 15/15 | 1 | 456.473 | 0.469 | 549.7 | 5.609 | 30.634 | 0.598 | 0.013
REF_90_30May_all15 | Supporting peak | 487 | FALSE | 15/15 | 1 | 487.149 | 0 | 399.754 | 8.286 | 22.321 | 0.434 | 0.023
REF_90_30May_all15 | Supporting peak | 1248.67 | FALSE | 15/15 | 1 | 1248.902 | 0.744 | 290.471 | 3.744 | 16.174 | 0.316 | 0.007
REF_90_30May_all15 | Uncertain | 583.8 | FALSE | 14/15 | 0.933 | 583.687 | 0 | 277.871 | 11.387 | 15.76 | 0.302 | 0.027
REF_90_30May_all15 | Uncertain | 1058.93 | FALSE | 15/15 | 1 | 1059.074 | 0.412 | 285.236 | 6.334 | 15.871 | 0.31 | 0.013
REF_90_30May_all15 | Uncertain | 1155 | FALSE | 15/15 | 1 | 1155.352 | 0 | 367.244 | 5.86 | 20.454 | 0.399 | 0.013
REF_90_30May_all15 | Supporting peak | 1589.47 | FALSE | 15/15 | 1 | 1589.401 | 0.724 | 263.092 | 8.866 | 14.7 | 0.286 | 0.02
REF_90_30May_all15 | Uncertain | 1344.87 | TRUE | 15/15 | 1 | 1344.687 | 1.403 | 627.231 | 17.119 | 35.202 | 0.681 | 0.111
REF_90_30May_all15 | Uncertain | 1392.47 | TRUE | 15/15 | 1 | 1392.497 | 0.761 | 607.975 | 18.201 | 34.132 | 0.66 | 0.112
REF_90_30May_all15 | Uncertain | 1450 | TRUE | 15/15 | 1 | 1450.124 | 0.511 | 499.804 | 10.015 | 27.971 | 0.543 | 0.046
REF_90_30May_indep5 | Primary marker | 541 | FALSE | 5/5 | 1 | 541.022 | 0 | 684.065 | 5.714 | 36.699 | 0.771 | 0.006
REF_90_30May_indep5 | Primary marker | 829 | FALSE | 5/5 | 1 | 828.481 | 0 | 596.551 | 5.837 | 32.044 | 0.672 | 0.009
REF_90_30May_indep5 | Primary marker | 870.53 | FALSE | 5/5 | 1 | 870.262 | 0 | 842.488 | 5.212 | 45.176 | 0.949 | 0.01
REF_90_30May_indep5 | Primary marker | 929 | FALSE | 5/5 | 1 | 929.755 | 0 | 887.636 | 5.341 | 47.627 | 1 | 0
REF_90_30May_indep5 | Primary marker | 955 | FALSE | 5/5 | 1 | 954.313 | 0 | 817.589 | 6.163 | 43.893 | 0.921 | 0.011
REF_90_30May_indep5 | Primary marker | 1022.73 | FALSE | 5/5 | 1 | 1022.358 | 0 | 704.642 | 7.375 | 37.873 | 0.793 | 0.025
REF_90_30May_indep5 | Primary marker | 1127 | FALSE | 5/5 | 1 | 1127.196 | 0 | 548.661 | 5.611 | 29.451 | 0.618 | 0.014
REF_90_30May_indep5 | Supporting peak | 457 | FALSE | 5/5 | 1 | 456.352 | 0 | 524.553 | 8.25 | 28.137 | 0.59 | 0.02
REF_90_30May_indep5 | Supporting peak | 487 | FALSE | 5/5 | 1 | 487.149 | 0 | 363.826 | 9.937 | 19.599 | 0.409 | 0.024
REF_90_30May_indep5 | Supporting peak | 1248.67 | FALSE | 5/5 | 1 | 1249.41 | 0 | 283.797 | 5.184 | 15.194 | 0.32 | 0.01
REF_90_30May_indep5 | Uncertain | 583.8 | FALSE | 4/5 | 0.8 | 583.687 | 0 | 238.222 | 14.353 | 13.535 | 0.271 | 0.033
REF_90_30May_indep5 | Uncertain | 1058.93 | FALSE | 5/5 | 1 | 1059.18 | 0 | 275.6 | 10.27 | 14.704 | 0.31 | 0.022
REF_90_30May_indep5 | Uncertain | 1155 | FALSE | 5/5 | 1 | 1155.352 | 0 | 350.06 | 9.15 | 18.741 | 0.394 | 0.021
REF_90_30May_indep5 | Supporting peak | 1589.47 | FALSE | 5/5 | 1 | 1589.307 | 0.768 | 237.599 | 9.993 | 12.815 | 0.268 | 0.024
REF_90_30May_indep5 | Uncertain | 1344.87 | TRUE | 5/5 | 1 | 1343.495 | 1.7 | 532.05 | 28.508 | 29.061 | 0.6 | 0.174
REF_90_30May_indep5 | Uncertain | 1392.47 | TRUE | 5/5 | 1 | 1392.988 | 0.659 | 501.287 | 29.117 | 27.387 | 0.565 | 0.164
REF_90_30May_indep5 | Uncertain | 1450 | TRUE | 5/5 | 1 | 1449.737 | 0.796 | 446.963 | 13.193 | 24.22 | 0.504 | 0.067
NEW_60_24Sep | Primary marker | 541 | FALSE | 11/11 | 1 | 539.236 | 0 | 362.359 | 16.445 | 22.881 | 0.765 | 0.018
NEW_60_24Sep | Primary marker | 829 | FALSE | 11/11 | 1 | 828.481 | 0 | 319.76 | 16.781 | 20.197 | 0.675 | 0.02
NEW_60_24Sep | Primary marker | 870.53 | FALSE | 11/11 | 1 | 870.262 | 0 | 456.331 | 13.954 | 28.834 | 0.967 | 0.027
NEW_60_24Sep | Primary marker | 929 | FALSE | 11/11 | 1 | 928.113 | 0 | 473.332 | 15.905 | 29.886 | 1 | 0
NEW_60_24Sep | Primary marker | 955 | FALSE | 11/11 | 1 | 954.313 | 0 | 424.273 | 18.138 | 26.802 | 0.894 | 0.031
NEW_60_24Sep | Primary marker | 1022.73 | FALSE | 11/11 | 1 | 1022.066 | 0.65 | 358.057 | 20.005 | 22.606 | 0.752 | 0.036
NEW_60_24Sep | Primary marker | 1127 | FALSE | 11/11 | 1 | 1125.626 | 0 | 270.425 | 20.766 | 17.074 | 0.567 | 0.034
NEW_60_24Sep | Supporting peak | 457 | FALSE | 11/11 | 1 | 456.352 | 0 | 265.897 | 17.737 | 16.787 | 0.56 | 0.016
NEW_60_24Sep | Supporting peak | 487 | FALSE | 11/11 | 1 | 486 | 0.911 | 191.129 | 18.534 | 12.078 | 0.402 | 0.02
NEW_60_24Sep | Supporting peak | 1248.67 | FALSE | 11/11 | 1 | 1248.024 | 0.823 | 156.995 | 16.015 | 9.91 | 0.332 | 0.017
NEW_60_24Sep | Uncertain | 583.8 | FALSE | 11/11 | 1 | 583.687 | 0 | 142.779 | 16.536 | 9.026 | 0.302 | 0.017
NEW_60_24Sep | Uncertain | 1058.93 | FALSE | 11/11 | 1 | 1058.021 | 0.745 | 137.628 | 21.912 | 8.684 | 0.288 | 0.026
NEW_60_24Sep | Uncertain | 1155 | FALSE | 11/11 | 1 | 1154.927 | 1.008 | 158.325 | 19.443 | 9.994 | 0.333 | 0.018
NEW_60_24Sep | Supporting peak | 1589.47 | FALSE | 11/11 | 1 | 1588.874 | 1.324 | 142.666 | 14.22 | 9.015 | 0.303 | 0.023
NEW_60_24Sep | Uncertain | 1344.87 | TRUE | 11/11 | 1 | 1345.338 | 1.205 | 383.897 | 27.432 | 24.261 | 0.852 | 0.365
NEW_60_24Sep | Uncertain | 1392.47 | TRUE | 11/11 | 1 | 1392.479 | 0.769 | 360.442 | 29.637 | 22.77 | 0.802 | 0.363
NEW_60_24Sep | Uncertain | 1450 | TRUE | 10/11 | 0.909 | 1449.011 | 0.459 | 265.747 | 8.385 | 16.748 | 0.555 | 0.055
NEW_70_24Sep | Primary marker | 541 | FALSE | 5/5 | 1 | 539.236 | 0 | 436.47 | 9.567 | 26.148 | 0.768 | 0.012
NEW_70_24Sep | Primary marker | 829 | FALSE | 5/5 | 1 | 828.481 | 0 | 392.385 | 10.798 | 23.493 | 0.69 | 0.02
NEW_70_24Sep | Primary marker | 870.53 | FALSE | 5/5 | 1 | 870.262 | 0 | 548.699 | 9.772 | 32.863 | 0.966 | 0.022
NEW_70_24Sep | Primary marker | 929 | FALSE | 5/5 | 1 | 928.113 | 0 | 568.66 | 10.559 | 34.072 | 1 | 0
NEW_70_24Sep | Primary marker | 955 | FALSE | 5/5 | 1 | 954.313 | 0 | 534.417 | 12.179 | 32.009 | 0.939 | 0.016
NEW_70_24Sep | Primary marker | 1022.73 | FALSE | 5/5 | 1 | 1022.358 | 0 | 446.865 | 13.933 | 26.761 | 0.784 | 0.03
NEW_70_24Sep | Primary marker | 1127 | FALSE | 5/5 | 1 | 1125.626 | 0 | 346.01 | 13.58 | 20.728 | 0.607 | 0.022
NEW_70_24Sep | Supporting peak | 457 | FALSE | 5/5 | 1 | 456.352 | 0 | 331.549 | 10.006 | 19.857 | 0.583 | 0.012
NEW_70_24Sep | Supporting peak | 487 | FALSE | 5/5 | 1 | 486.788 | 0.808 | 241.128 | 14.465 | 14.428 | 0.423 | 0.022
NEW_70_24Sep | Supporting peak | 1248.67 | FALSE | 5/5 | 1 | 1248.19 | 0.682 | 187.553 | 8.737 | 11.235 | 0.331 | 0.014
NEW_70_24Sep | Uncertain | 583.8 | FALSE | 5/5 | 1 | 583.687 | 0 | 170.694 | 15.203 | 10.217 | 0.299 | 0.014
NEW_70_24Sep | Uncertain | 1058.93 | FALSE | 5/5 | 1 | 1058.224 | 0.873 | 165.011 | 13.637 | 9.882 | 0.29 | 0.017
NEW_70_24Sep | Uncertain | 1155 | FALSE | 5/5 | 1 | 1154.729 | 0.854 | 215.184 | 14.466 | 12.893 | 0.377 | 0.017
NEW_70_24Sep | Supporting peak | 1589.47 | FALSE | 5/5 | 1 | 1589.027 | 0.627 | 168.251 | 13.38 | 10.076 | 0.295 | 0.01
NEW_70_24Sep | Uncertain | 1344.87 | TRUE | 5/5 | 1 | 1345.582 | 0.666 | 430.633 | 22.362 | 25.807 | 0.759 | 0.15
NEW_70_24Sep | Uncertain | 1392.47 | TRUE | 5/5 | 1 | 1390.925 | 0.807 | 400.891 | 23.334 | 24.026 | 0.706 | 0.147
NEW_70_24Sep | Uncertain | 1450 | TRUE | 5/5 | 1 | 1450.318 | 0 | 322.747 | 13.167 | 19.334 | 0.568 | 0.045
NEW_80_24Sep | Primary marker | 541 | FALSE | 4/4 | 1 | 540.575 | 0.893 | 567.773 | 6.22 | 26.819 | 0.795 | 0.055
NEW_80_24Sep | Primary marker | 829 | FALSE | 4/4 | 1 | 828.9 | 0.839 | 491.427 | 7.281 | 22.982 | 0.687 | 0.029
NEW_80_24Sep | Primary marker | 870.53 | FALSE | 4/4 | 1 | 870.678 | 0.832 | 695.865 | 7.942 | 32.534 | 0.971 | 0.026
NEW_80_24Sep | Primary marker | 929 | FALSE | 4/4 | 1 | 929.345 | 0.821 | 717.739 | 10.421 | 33.352 | 1 | 0
NEW_80_24Sep | Primary marker | 955 | FALSE | 4/4 | 1 | 955.129 | 0.942 | 655.004 | 8.632 | 30.837 | 0.915 | 0.049
NEW_80_24Sep | Primary marker | 1022.73 | FALSE | 4/4 | 1 | 1022.76 | 0.804 | 607.729 | 8.212 | 28.875 | 0.85 | 0.077
NEW_80_24Sep | Primary marker | 1127 | FALSE | 4/4 | 1 | 1126.804 | 0.785 | 444.033 | 10.9 | 20.883 | 0.619 | 0.037
NEW_80_24Sep | Supporting peak | 457 | FALSE | 4/4 | 1 | 457.26 | 1.049 | 443.66 | 9.04 | 21.407 | 0.623 | 0.091
NEW_80_24Sep | Supporting peak | 487 | FALSE | 4/4 | 1 | 487.149 | 0 | 366.044 | 17.697 | 18.017 | 0.515 | 0.121
NEW_80_24Sep | Supporting peak | 1248.67 | FALSE | 4/4 | 1 | 1248.266 | 0.763 | 244.161 | 4.343 | 11.574 | 0.342 | 0.029
NEW_80_24Sep | Uncertain | 583.8 | FALSE | 4/4 | 1 | 584.129 | 0.885 | 282.985 | 36.667 | 14.638 | 0.404 | 0.181
NEW_80_24Sep | Uncertain | 1058.93 | FALSE | 4/4 | 1 | 1058.782 | 0.797 | 230.043 | 6.385 | 10.863 | 0.322 | 0.021
NEW_80_24Sep | Uncertain | 1155 | FALSE | 4/4 | 1 | 1155.352 | 0 | 273.842 | 5.729 | 13.023 | 0.384 | 0.034
NEW_80_24Sep | Supporting peak | 1589.47 | FALSE | 4/4 | 1 | 1589.798 | 1.342 | 247.168 | 17.452 | 12.091 | 0.348 | 0.079
NEW_80_24Sep | Uncertain | 1344.87 | TRUE | 4/4 | 1 | 1343.644 | 1.491 | 430.638 | 9.878 | 20.484 | 0.602 | 0.063
NEW_80_24Sep | Uncertain | 1392.47 | TRUE | 4/4 | 1 | 1392.914 | 0.737 | 400.297 | 10.202 | 19.049 | 0.56 | 0.058
NEW_80_24Sep | Uncertain | 1450 | TRUE | 4/4 | 1 | 1449.228 | 0.726 | 361.54 | 11.139 | 16.844 | 0.504 | 0.022
NEW_90_24Sep | Primary marker | 541 | FALSE | 1/1 | 1 | 541.022 |  | 568.084 |  | 22.029 | 0.802
NEW_90_24Sep | Primary marker | 829 | FALSE | 1/1 | 1 | 830.159 |  | 510.353 |  | 19.79 | 0.72
NEW_90_24Sep | Primary marker | 870.53 | FALSE | 1/1 | 1 | 870.262 |  | 692.427 |  | 26.85 | 0.977
NEW_90_24Sep | Primary marker | 929 | FALSE | 1/1 | 1 | 929.755 |  | 708.644 |  | 27.479 | 1
NEW_90_24Sep | Primary marker | 955 | FALSE | 1/1 | 1 | 955.945 |  | 667.103 |  | 25.868 | 0.941
NEW_90_24Sep | Primary marker | 1022.73 | FALSE | 1/1 | 1 | 1022.358 |  | 584.594 |  | 22.669 | 0.825
NEW_90_24Sep | Primary marker | 1127 | FALSE | 1/1 | 1 | 1127.196 |  | 466.833 |  | 18.103 | 0.659
NEW_90_24Sep | Supporting peak | 457 | FALSE | 1/1 | 1 | 456.352 |  | 395.231 |  | 15.326 | 0.558
NEW_90_24Sep | Supporting peak | 487 | FALSE | 1/1 | 1 | 487.149 |  | 308.992 |  | 11.982 | 0.436
NEW_90_24Sep | Supporting peak | 1248.67 | FALSE | 1/1 | 1 | 1249.41 |  | 231.581 |  | 8.98 | 0.327
NEW_90_24Sep | Uncertain | 583.8 | FALSE | 1/1 | 1 | 583.687 |  | 201 |  | 7.794 | 0.284
NEW_90_24Sep | Uncertain | 1058.93 | FALSE | 1/1 | 1 | 1059.18 |  | 222.748 |  | 8.638 | 0.314
NEW_90_24Sep | Uncertain | 1155 | FALSE | 1/1 | 1 | 1155.352 |  | 300.027 |  | 11.634 | 0.423
NEW_90_24Sep | Supporting peak | 1589.47 | FALSE | 1/1 | 1 | 1588.747 |  | 213.592 |  | 8.283 | 0.301
NEW_90_24Sep | Uncertain | 1344.87 | TRUE | 1/1 | 1 | 1344.389 |  | 456.99 |  | 17.721 | 0.645
NEW_90_24Sep | Uncertain | 1392.47 | TRUE | 1/1 | 1 | 1393.282 |  | 449.104 |  | 17.415 | 0.634
NEW_90_24Sep | Uncertain | 1450 | TRUE | 1/1 | 1 | 1450.318 |  | 402.374 |  | 15.603 | 0.568
NEW_all_pass_24Sep | Primary marker | 541 | FALSE | 21/21 | 1 | 539.576 | 0.719 | 428.928 | 22.827 | 24.368 | 0.773 | 0.029
NEW_all_pass_24Sep | Primary marker | 829 | FALSE | 21/21 | 1 | 828.641 | 0.505 | 378.826 | 22.641 | 21.493 | 0.683 | 0.023
NEW_all_pass_24Sep | Primary marker | 870.53 | FALSE | 21/21 | 1 | 870.341 | 0.363 | 535.192 | 21.243 | 30.404 | 0.968 | 0.024
NEW_all_pass_24Sep | Primary marker | 929 | FALSE | 21/21 | 1 | 928.426 | 0.661 | 553.788 | 21.807 | 31.428 | 1 | 0
NEW_all_pass_24Sep | Primary marker | 955 | FALSE | 21/21 | 1 | 954.546 | 0.585 | 506.01 | 23.301 | 28.766 | 0.911 | 0.036
NEW_all_pass_24Sep | Primary marker | 1022.73 | FALSE | 21/21 | 1 | 1022.282 | 0.618 | 437.546 | 27.124 | 24.792 | 0.782 | 0.057
NEW_all_pass_24Sep | Primary marker | 1127 | FALSE | 21/21 | 1 | 1125.925 | 0.632 | 330.842 | 26.903 | 18.718 | 0.591 | 0.04
NEW_all_pass_24Sep | Supporting peak | 457 | FALSE | 21/21 | 1 | 456.525 | 0.547 | 321.547 | 25.181 | 18.328 | 0.578 | 0.045
NEW_all_pass_24Sep | Supporting peak | 487 | FALSE | 21/21 | 1 | 486.461 | 0.899 | 241.963 | 32.609 | 13.764 | 0.43 | 0.066
NEW_all_pass_24Sep | Supporting peak | 1248.67 | FALSE | 21/21 | 1 | 1248.176 | 0.781 | 184.425 | 21.876 | 10.498 | 0.333 | 0.018
NEW_all_pass_24Sep | Uncertain | 583.8 | FALSE | 21/21 | 1 | 583.771 | 0.386 | 178.904 | 39.342 | 10.32 | 0.32 | 0.083
NEW_all_pass_24Sep | Uncertain | 1058.93 | FALSE | 21/21 | 1 | 1058.269 | 0.809 | 165.804 | 27.057 | 9.382 | 0.296 | 0.026
NEW_all_pass_24Sep | Uncertain | 1155 | FALSE | 21/21 | 1 | 1154.981 | 0.84 | 200.614 | 28.669 | 11.34 | 0.357 | 0.034
NEW_all_pass_24Sep | Supporting peak | 1589.47 | FALSE | 21/21 | 1 | 1589.08 | 1.165 | 172.04 | 27.752 | 9.819 | 0.31 | 0.04
NEW_all_pass_24Sep | Uncertain | 1344.87 | TRUE | 21/21 | 1 | 1345.028 | 1.297 | 407.409 | 22.417 | 23.598 | 0.772 | 0.286
NEW_all_pass_24Sep | Uncertain | 1392.47 | TRUE | 21/21 | 1 | 1392.23 | 1.057 | 381.886 | 23.918 | 22.105 | 0.725 | 0.283
NEW_all_pass_24Sep | Uncertain | 1450 | TRUE | 20/21 | 0.952 | 1449.446 | 0.73 | 305.987 | 17.709 | 17.356 | 0.549 | 0.05

## 31_NAM_Bulk_Intensity_Power
set | band_ref | pos_mean | pos_sd | height_mean_counts | height_cv_pct | snr_mean | detected
REF_90_30May_all15 | 541 | 541.02 | 0 | 715.42 | 4.55 | 39.87 | 15/15
REF_90_30May_all15 | 829 | 828.48 | 0 | 627.33 | 4.86 | 34.98 | 15/15
REF_90_30May_all15 | 870.53 | 870.26 | 0 | 879.63 | 4.28 | 49.02 | 15/15
REF_90_30May_all15 | 929 | 929.32 | 0.75 | 919.44 | 3.92 | 51.24 | 15/15
REF_90_30May_all15 | 955 | 954.53 | 0.57 | 864.03 | 5.13 | 48.18 | 15/15
REF_90_30May_all15 | 1022.73 | 1022.36 | 0 | 756.31 | 6.31 | 42.2 | 15/15
REF_90_30May_all15 | 1127 | 1127.2 | 0 | 582.1 | 5.12 | 32.46 | 15/15
REF_90_30May_indep5 | 541 | 541.02 | 0 | 684.06 | 5.71 | 36.7 | 5/5
REF_90_30May_indep5 | 829 | 828.48 | 0 | 596.55 | 5.84 | 32.04 | 5/5
REF_90_30May_indep5 | 870.53 | 870.26 | 0 | 842.49 | 5.21 | 45.18 | 5/5
REF_90_30May_indep5 | 929 | 929.76 | 0 | 887.64 | 5.34 | 47.63 | 5/5
REF_90_30May_indep5 | 955 | 954.31 | 0 | 817.59 | 6.16 | 43.89 | 5/5
REF_90_30May_indep5 | 1022.73 | 1022.36 | 0 | 704.64 | 7.38 | 37.87 | 5/5
REF_90_30May_indep5 | 1127 | 1127.2 | 0 | 548.66 | 5.61 | 29.45 | 5/5
NEW_60_24Sep | 541 | 539.24 | 0 | 362.36 | 16.44 | 22.88 | 11/11
NEW_60_24Sep | 829 | 828.48 | 0 | 319.76 | 16.78 | 20.2 | 11/11
NEW_60_24Sep | 870.53 | 870.26 | 0 | 456.33 | 13.95 | 28.83 | 11/11
NEW_60_24Sep | 929 | 928.11 | 0 | 473.33 | 15.9 | 29.89 | 11/11
NEW_60_24Sep | 955 | 954.31 | 0 | 424.27 | 18.14 | 26.8 | 11/11
NEW_60_24Sep | 1022.73 | 1022.07 | 0.65 | 358.06 | 20 | 22.61 | 11/11
NEW_60_24Sep | 1127 | 1125.63 | 0 | 270.42 | 20.77 | 17.07 | 11/11
NEW_70_24Sep | 541 | 539.24 | 0 | 436.47 | 9.57 | 26.15 | 5/5
NEW_70_24Sep | 829 | 828.48 | 0 | 392.38 | 10.8 | 23.49 | 5/5
NEW_70_24Sep | 870.53 | 870.26 | 0 | 548.7 | 9.77 | 32.86 | 5/5
NEW_70_24Sep | 929 | 928.11 | 0 | 568.66 | 10.56 | 34.07 | 5/5
NEW_70_24Sep | 955 | 954.31 | 0 | 534.42 | 12.18 | 32.01 | 5/5
NEW_70_24Sep | 1022.73 | 1022.36 | 0 | 446.86 | 13.93 | 26.76 | 5/5
NEW_70_24Sep | 1127 | 1125.63 | 0 | 346.01 | 13.58 | 20.73 | 5/5
NEW_80_24Sep | 541 | 540.58 | 0.89 | 567.77 | 6.22 | 26.82 | 4/4
NEW_80_24Sep | 829 | 828.9 | 0.84 | 491.43 | 7.28 | 22.98 | 4/4
NEW_80_24Sep | 870.53 | 870.68 | 0.83 | 695.86 | 7.94 | 32.53 | 4/4
NEW_80_24Sep | 929 | 929.34 | 0.82 | 717.74 | 10.42 | 33.35 | 4/4
NEW_80_24Sep | 955 | 955.13 | 0.94 | 655 | 8.63 | 30.84 | 4/4
NEW_80_24Sep | 1022.73 | 1022.76 | 0.8 | 607.73 | 8.21 | 28.88 | 4/4
NEW_80_24Sep | 1127 | 1126.8 | 0.78 | 444.03 | 10.9 | 20.88 | 4/4
NEW_90_24Sep | 541 | 541.02 |  | 568.08 |  | 22.03 | 1/1
NEW_90_24Sep | 829 | 830.16 |  | 510.35 |  | 19.79 | 1/1
NEW_90_24Sep | 870.53 | 870.26 |  | 692.43 |  | 26.85 | 1/1
NEW_90_24Sep | 929 | 929.76 |  | 708.64 |  | 27.48 | 1/1
NEW_90_24Sep | 955 | 955.94 |  | 667.1 |  | 25.87 | 1/1
NEW_90_24Sep | 1022.73 | 1022.36 |  | 584.59 |  | 22.67 | 1/1
NEW_90_24Sep | 1127 | 1127.2 |  | 466.83 |  | 18.1 | 1/1

## 32_NAM_20260924_Decision
Item | Finding | Decision
Sample identity | Operator confirmed on 27 Sep 2026 that the powder measured on 24 Sep was NAM (folder named "lps" in error). Spectra match the NAM reference (r = 0.98–0.99 for set means). | Files renamed NAM_Bulk_20260924_*; the LPS screening outputs were withdrawn (moved to _to_delete/LPS_mislabel_20260924_superseded).
New data vs existing reference | 21 QC-passed independent spots (60 %: 11, 70 %: 5, 80 %: 4, 90 %: 1). All 7 primary and 4 supporting markers detected in 21/21; position shift ≤ 1.5 cm-1 (< 1 pixel); relative intensity (I/I929) within ±7 % for primary markers. | SUPPORTS the existing NAM reference (band positions, relative intensities, marker hierarchy).
Reference condition | Absolute intensity differs between sessions (929 band at 90 %: 919 counts on 30 May vs 709 counts, n = 1, on 24 Sep) — not a reason to change the condition. | Reference condition unchanged: 90 % / 30 s / 5 acc.
Top-up at 90 % | Only 1 of 13 new spectra at 90 % passed QC (dark-reference ripple). | Reference remains INTERIM (6 independent spots at 90 % incl. the new S09). Collect ≥ 4 more clean independent spots at 90/30/5 (target 10) with a fresh dark before every spectrum.
Glass blanks | 16 matched blanks, all saturated across the glass band. | Glass-zone bands (1345, 1392, 1450) stay "uncertain – inside glass zone"; add one labelled short-integration blank next session.
LPS | No LPS data in this session. | LPS screening (70/80/90 % × 30 s × 5 acc + blanks) still required; the earlier "70 % provisional" LPS recommendation is withdrawn because it was based on NAM data.

## 33_LPS_20260928_QC
library_file | category | power | int_s | acc | orig | md5 | copy_ok | acq | dark_med | ripple | raw_max | nsat | sat_from | sat_to | ds_med | noise | strong_pos | strong_h | snr | rNAM | rLTA | rOLD | rBlank | r_to_others_mean | in_core | QC | QC_reason | note
Glass_Blank_20260928_P070pct_T30s_A05_S01.csv | Glass_Blank | 70 | 30 | 5 | 30s-70p-5ac/glass1.csv | eae1069dd19a16f512a335f834350f58 | TRUE | 2026-09-28 06:44:02 | 1405.8 | 175.035 | 65535 | 270 | 1273.738 | 1662.368 | 38153 | 48.982 | 1246.359 | 9455.483 | 193.038 | -0.077 | -0.196 | -0.267 | 1 |  | FALSE | BLANK-SATURATED
Glass_Blank_20260928_P070pct_T30s_A05_S02.csv | Glass_Blank | 70 | 30 | 5 | 30s-70p-5ac/glass2.csv | 63effa1b1df19688746067b6a9db2652 | TRUE | 2026-09-28 06:52:56 | 1405.8 | 172.802 | 65535 | 281 | 1267.669 | 1671.988 | 38637.4 | 51.427 | 1247.885 | 10047.928 | 195.382 | -0.071 | -0.198 | -0.268 | 0.999 |  | FALSE | BLANK-SATURATED
Glass_Blank_20260928_P070pct_T30s_A05_S03.csv | Glass_Blank | 70 | 30 | 5 | 30s-70p-5ac/glass3.csv | e680afa7fb18d5619cf154ca776e8d4d | TRUE | 2026-09-28 06:55:44 | 1405.8 | 170.92 | 65535 | 276 | 1270.705 | 1667.868 | 37470.6 | 51.545 | 1264.632 | 9363.149 | 181.649 | -0.083 | -0.189 | -0.259 | 0.999 |  | FALSE | BLANK-SATURATED
LPS_Bulk_20260928_P070pct_T30s_A05_S01.csv | LPS_Bulk | 70 | 30 | 5 | 30s-70p-5ac/powder-bulk-lps-30s-70p-5ac-1.csv | cde72e7e2c6fd68c6bcf3df67f3f44d6 | TRUE | 2026-09-28 03:26:57 | 1364 | 14.451 | 6148 | 0 |  |  | 1859.2 | 13.714 | 425.358 | 388.957 | 28.362 | 0.158 | 0.514 | 0.77 | -0.055 | 0.959 | TRUE | PASS
LPS_Bulk_20260928_P070pct_T30s_A05_S02.csv | LPS_Bulk | 70 | 30 | 5 | 30s-70p-5ac/powder-bulk-lps-30s-70p-5ac-2.csv | 2c50634ccfb001c89103f919e04e01e6 | TRUE | 2026-09-28 03:33:46 | 1364 | 18.661 | 8457.6 | 0 |  |  | 2661.8 | 17.01 | 1456.122 | 485.299 | 28.53 | 0.065 | 0.324 | 0.547 | 0.376 | 0.672 | FALSE | REVIEW | Valid but glass-dominated (r to other spots = 0.67 < 0.85; strong glass background) - excluded from mean
LPS_Bulk_20260928_P070pct_T30s_A05_S03.csv | LPS_Bulk | 70 | 30 | 5 | 30s-70p-5ac/powder-bulk-lps-30s-70p-5ac-3.csv | 06989184cc36ed7a8b0de684eaeb2d7d | TRUE | 2026-09-28 03:37:51 | 1364 | 15.235 | 5865.4 | 0 |  |  | 1641.8 | 13.787 | 425.358 | 316.258 | 22.939 | 0.156 | 0.48 | 0.767 | -0.04 | 0.969 | TRUE | PASS
LPS_Bulk_20260928_P070pct_T30s_A05_S04.csv | LPS_Bulk | 70 | 30 | 5 | 30s-70p-5ac/powder-bulk-lps-30s-70p-5ac-4.csv | 9b05d8e0568540843b6b9b6c888eaf3c | TRUE | 2026-09-28 03:51:59 | 1364 | 13.604 | 6096.4 | 0 |  |  | 1851.6 | 13.645 | 425.358 | 368.865 | 27.033 | 0.172 | 0.559 | 0.766 | -0.086 | 0.94 | TRUE | PASS
LPS_Bulk_20260928_P070pct_T30s_A05_S05.csv | LPS_Bulk | 70 | 30 | 5 | 30s-70p-5ac/powder-bulk-lps-30s-70p-5ac-5.csv | e2860350d3baf8f4d92b0858aaa6ea6e | TRUE | 2026-09-28 04:16:50 | 1405.8 | 14.931 | 5832.6 | 0 |  |  | 1558.6 | 13.871 | 425.358 | 305.631 | 22.034 | 0.185 | 0.516 | 0.777 | -0.02 | 0.961 | TRUE | PASS
Glass_Blank_20260928_P080pct_T30s_A05_S01.csv | Glass_Blank | 80 | 30 | 5 | 30s-80p-5a/glass1.csv | f9f5b38edea323403ae4ec97a049167e | TRUE | 2026-09-28 07:09:39 | 1412.4 | 177.382 | 65535 | 352 | 1220.337 | 1727.848 | 45648.8 | 57.138 | 1217.265 | 12255.101 | 214.483 | -0.07 | -0.226 | -0.244 | 0.62 |  | FALSE | BLANK-SATURATED
Glass_Blank_20260928_P080pct_T30s_A05_S02.csv | Glass_Blank | 80 | 30 | 5 | 30s-80p-5a/glass2.csv | 968856843bb00148c50ec10e6062e14c | TRUE | 2026-09-28 07:13:31 | 1412.4 | 138.83 | 65535 | 440 | 1167.807 | 1798.853 | 61349 | 61.697 | 1164.697 | 14935.508 | 242.08 | 0.064 | -0.1 | -0.048 | 0.993 |  | FALSE | BLANK-SATURATED
Glass_Blank_20260928_P080pct_T30s_A05_S03.csv | Glass_Blank | 80 | 30 | 5 | 30s-80p-5a/glass3.csv | 5f375aa22d483d355167bf3caf792647 | TRUE | 2026-09-28 07:17:20 | 1412.4 | 138.83 | 65535 | 439 | 1169.361 | 1798.853 | 60076.4 | 59.417 | 1167.807 | 15055.306 | 253.382 | 0.052 | -0.113 | -0.068 | 0.998 |  | FALSE | BLANK-SATURATED
Glass_Blank_20260928_P080pct_T30s_A05_S04.csv | Glass_Blank | 80 | 30 | 5 | 30s-80p-5a/glass4.csv | 57ba463f1be91bd5b3a32e2fbc287c5f | TRUE | 2026-09-28 07:38:11 | 1412.4 | 138.83 | 65535 | 445 | 1160.027 | 1798.853 | 63030.8 | 63.326 | 1158.47 | 15130.985 | 238.938 | 0.087 | -0.08 | -0.023 | 0.982 |  | FALSE | BLANK-SATURATED
LPS_Bulk_20260928_P080pct_T30s_A05_S01.csv | LPS_Bulk | 80 | 30 | 5 | 30s-80p-5a/spot1.csv | a4dca6f466039d56a24646bfeaa26c0c | TRUE | 2026-09-28 07:59:39 | 1412.4 | 203.681 | 10440.6 | 0 |  |  | 5993.8 | 68.952 | 549.942 | 1252.433 | 18.164 | 0.111 | 0.27 | 0.527 | 0.232 |  | FALSE | FAIL | Ripple (fixed-pattern residual 204 > 50 counts)
LPS_Bulk_20260928_P080pct_T30s_A05_S02.csv | LPS_Bulk | 80 | 30 | 5 | 30s-80p-5a/spot2.csv | 3ecaba23f0ec22688bfb049cb1d60d5e | TRUE | 2026-09-28 08:37:25 | 1412.4 | 478.169 | 10480.2 | 0 |  |  | 5582 | 132.764 | 549.942 | 2801.285 | 21.1 | 0.028 | 0.159 | 0.384 | 0.176 |  | FALSE | FAIL | Ripple (fixed-pattern residual 478 > 50 counts)
LPS_Bulk_20260928_P080pct_T30s_A05_S03.csv | LPS_Bulk | 80 | 30 | 5 | 30s-80p-5a/spot3.csv | 49f6c40e6261ab664de199c3a4a90f11 | TRUE | 2026-09-28 08:52:08 | 1412.4 | 474.002 | 9841.8 | 0 |  |  | 5311.8 | 130.312 | 549.942 | 2759.681 | 21.178 | 0.035 | 0.15 | 0.392 | 0.153 |  | FALSE | FAIL | Ripple (fixed-pattern residual 474 > 50 counts)
LPS_Bulk_20260928_P080pct_T30s_A05_S04.csv | LPS_Bulk | 80 | 30 | 5 | 30s-80p-5a/spot4.csv | 09c4a404f8f351a168296c27c007ab75 | TRUE | 2026-09-28 09:00:21 | 1412.4 | 580.816 | 11386.2 | 0 |  |  | 6860.2 | 150.979 | 549.942 | 3324.179 | 22.017 | 0.039 | 0.163 | 0.407 | 0.165 |  | FALSE | FAIL | Ripple (fixed-pattern residual 581 > 50 counts)
LPS_Bulk_20260928_P080pct_T30s_A05_S05.csv | LPS_Bulk | 80 | 30 | 5 | 30s-80p-5a/spot5.csv | ff7e206ad0117741f5a4a4f68830549f | TRUE | 2026-09-28 09:08:28 | 3354.7 | 128.495 | 10198.6 | 0 |  |  | 680.2 | 43.662 | 583.687 | 792.96 | 18.161 | -0.016 | -0.07 | 0.159 | -0.29 |  | FALSE | FAIL | Ripple (fixed-pattern residual 128 > 50 counts)
Glass_Blank_20260928_P060pct_T30s_A05_S01.csv | Glass_Blank | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/glass1.csv | bbac44486ba280273585327428fc6dbd | TRUE | 2026-09-27 21:25:53 | 1429.5 | 104.641 | 65535 | 108 | 1318.973 | 1476.372 | 23060.2 | 45.295 | 1306.959 | 8818.673 | 194.694 | -0.036 | 0.125 | 0.07 | 0.993 |  | FALSE | BLANK-SATURATED
Glass_Blank_20260928_P060pct_T30s_A05_S02.csv | Glass_Blank | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/glass2.csv | 630fffa99ebabf50c87a8053cbb5b81b | TRUE | 2026-09-27 21:28:37 | 1429.5 | 109.396 | 65535 | 101 | 1317.474 | 1466.26 | 23629.8 | 45.352 | 1306.959 | 8633.706 | 190.372 | -0.011 | 0.166 | 0.117 | 0.969 |  | FALSE | BLANK-SATURATED
Glass_Blank_20260928_P060pct_T30s_A05_S03.csv | Glass_Blank | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/glass3.csv | f53bf245af94b4788f9466f9240ebd2f | TRUE | 2026-09-27 21:31:50 | 1429.5 | 108.451 | 65535 | 114 | 1312.97 | 1480.698 | 23600.6 | 44.951 | 1306.959 | 10172.445 | 226.302 | -0.036 | 0.121 | 0.042 | 0.985 |  | FALSE | BLANK-SATURATED
LPS_Bulk_20260928_P060pct_T30s_A05_S01.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-1.csv | 07f62fb99314ff3b1d94dbb29b064d0c | TRUE | 2026-09-27 21:37:32 | 1429.5 | 13.478 | 5914.6 | 0 |  |  | 1565.4 | 13.544 | 425.358 | 251.724 | 18.586 | 0.158 | 0.545 | 0.785 | 0.233 | 0.973 | TRUE | PASS
LPS_Bulk_20260928_P060pct_T30s_A05_S10.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-10.csv | c88d3cf5a5751c99226c115052807170 | TRUE | 2026-09-28 03:19:01 | 1381.7 | 23.162 | 12308 | 0 |  |  | 3255.8 | 18.283 | 1456.122 | 773.397 | 42.301 | 0.073 | 0.224 | 0.358 | 0.416 | 0.405 | FALSE | REVIEW | Valid but glass-dominated (r to other spots = 0.41 < 0.85; strong glass background) - excluded from mean
LPS_Bulk_20260928_P060pct_T30s_A05_S02.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-2.csv | 442d18ee55d009d12d5ad3f8d2212d4f | TRUE | 2026-09-27 21:41:50 | 1441.6 | 13.688 | 5469 | 0 |  |  | 1112 | 12.787 | 425.358 | 194.245 | 15.191 | 0.175 | 0.537 | 0.746 | 0.218 | 0.944 | TRUE | PASS
LPS_Bulk_20260928_P060pct_T30s_A05_S03.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-3.csv | 44525d10009f36c4889d22502c5a7ced | TRUE | 2026-09-27 21:57:34 | 1441.6 | 11.94 | 5326.8 | 0 |  |  | 963 | 12.654 | 425.358 | 207.686 | 16.413 | 0.168 | 0.531 | 0.751 | 0.178 | 0.94 | TRUE | PASS
LPS_Bulk_20260928_P060pct_T30s_A05_S04.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-4.csv | 5ad89a53a0aa82716d7fbd7054e21084 | TRUE | 2026-09-27 22:00:33 | 1441.6 | 12.904 | 5589 | 0 |  |  | 1164.8 | 12.303 | 425.358 | 225.715 | 18.346 | 0.185 | 0.538 | 0.767 | 0.152 | 0.953 | TRUE | PASS
LPS_Bulk_20260928_P060pct_T30s_A05_S06.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-6.csv | 0715a47675fa5b07cb1998cee63ae448 | TRUE | 2026-09-28 03:02:55 | 1416.6 | 14.715 | 6037.8 | 0 |  |  | 1719.344 | 15.243 | 425.358 | 307.64 | 20.182 | 0.163 | 0.54 | 0.816 | 0.186 | 0.976 | TRUE | PASS |  | Instrument software applied Savitzky-Golay smoothing to the stored dark-subtracted column (header Data Pretreat_0); dark-subtracted recomputed as Raw - Dark (same definition as all other files)
LPS_Bulk_20260928_P060pct_T30s_A05_S07.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-7.csv | a51fe3f1f7490c014fd927d22ba6c1b7 | TRUE | 2026-09-28 03:05:46 | 1416.6 | 14.89 | 6188.8 | 0 |  |  | 1918.123 | 14.755 | 423.528 | 366.022 | 24.807 | 0.141 | 0.498 | 0.804 | 0.207 | 0.97 | TRUE | PASS |  | Instrument software applied Savitzky-Golay smoothing to the stored dark-subtracted column (header Data Pretreat_0); dark-subtracted recomputed as Raw - Dark (same definition as all other files)
LPS_Bulk_20260928_P060pct_T30s_A05_S08.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-8.csv | a49ea476800a35b4d76d4977f5a93221 | TRUE | 2026-09-28 03:09:26 | 1416.6 | 15.096 | 6300 | 0 |  |  | 1954 | 14.174 | 423.528 | 366.581 | 25.864 | 0.167 | 0.512 | 0.814 | 0.104 | 0.956 | TRUE | PASS
LPS_Bulk_20260928_P060pct_T30s_A05_S09.csv | LPS_Bulk | 60 | 30 | 5 | raw glass slide and 30s-60p-5a/powder-bulk-lps-30s-60p-5ac-9.csv | 1c8156d84ccdbc76d10e6bfbe2204694 | TRUE | 2026-09-28 03:13:16 | 1381.7 | 13.16 | 5984.6 | 0 |  |  | 1662.2 | 14.127 | 425.358 | 336.749 | 23.837 | 0.185 | 0.539 | 0.777 | 0.123 | 0.949 | TRUE | PASS

## 34_LPS_20260928_Power
power | n_spectra | n_pass | n_review | n_fail | n_core | n_blanks | pairwise_r | r_mean_median | h425_mean | h425_cv | snr425 | r_mean_vs_April_LPS | r_mean_vs_NAM
60 | 9 | 8 | 1 | 0 | 8 | 3 | 0.938 | 0.999 | 282.045 | 25.134 | 20.403 | 0.804 | 0.173
70 | 5 | 4 | 1 | 0 | 4 | 3 | 0.963 | 1 | 344.928 | 11.69 | 25.092 | 0.781 | 0.17
80 | 5 | 0 | 0 | 5 | 0 | 4

## 35_LPS_20260928_Bands
band | 60_det | 60_frac | 60_pos | 60_pos_sd | 60_h | 60_cv | 60_snr | 70_det | 70_frac | 70_pos | 70_pos_sd | 70_h | 70_cv | 70_snr | fwhm | glass_zone | blank | LPS_lit_within_10 | near_NAM_marker | class
425.4 | 8/8 | 1 | 424.9 | 0.85 | 282.05 | 25.13 | 20.4 | 4/4 | 1 | 425.36 | 0 | 344.93 | 11.69 | 25.09 | 14.68 | FALSE | not in blank | none | FALSE | C1 - candidate primary band
474.5 | 3/8 | 0.38 | 474.49 | 0 | 49.07 | 25.78 | 3.57 | 2/4 | 0.5 | 476.3 | 2.56 | 48.08 | 8.83 | 3.51 | 13.09 | FALSE | not in blank | none | FALSE | C4 - uncertain / weak
539.2 | 5/8 | 0.62 | 539.95 | 2.04 | 53.46 | 19.85 | 3.83 | 0/4 | 0 | [5 empty] | 0 | FALSE | not in blank | 543 (saccharide and lipid; L1) | TRUE | C4 - uncertain / weak
659.2 | 4/8 | 0.5 | 659.62 | 2.19 | 46.86 | 12 | 3.37 | 1/4 | 0.25 | 662.66 |  | 53.93 |  | 3.93 | 22.61 | FALSE | not in blank | none | FALSE | C4 - uncertain / weak
765.1 | 8/8 | 1 | 764.65 | 3.25 | 96.21 | 22.06 | 7 | 4/4 | 1 | 765.51 | 0.85 | 124.94 | 23.78 | 9.09 | 19.56 | FALSE | not in blank | none | FALSE | C2 - candidate supporting band (weak)
860.3 | 4/8 | 0.5 | 860.69 | 0.83 | 114.19 | 6.87 | 8.12 | 1/4 | 0.25 | 865.27 |  | 115.35 |  | 8.32 | 8.25 | FALSE | not in blank | none | FALSE | C4 - uncertain / weak
896.8 | 2/8 | 0.25 | 896.8 | 0 | 124.87 | 9.72 | 8.82 | 3/4 | 0.75 | 899 | 2.52 | 130.33 | 10.68 | 9.51 | 101.65 | FALSE | not in blank | 894 (common LPS peak; L2) | FALSE | C4 - uncertain / weak
996.6 | 7/8 | 0.88 | 996.33 | 1.46 | 56.25 | 20.77 | 4.04 | 4/4 | 1 | 995.75 | 2.09 | 53.24 | 21.57 | 3.87 | 25.7 | FALSE | not in blank | 988 (β(CH) carbohydrates; L1); 1003 (common LPS peak; L2) | FALSE | C2 - candidate supporting band (weak)
1081.4 | 8/8 | 1 | 1081.04 | 3.03 | 110.88 | 24.71 | 8.04 | 4/4 | 1 | 1083.03 | 2.89 | 132.93 | 18.86 | 9.67 | 47.67 | FALSE | not in blank | none | FALSE | C2 - candidate supporting band (weak)
1125.6 | 4/8 | 0.5 | 1125.23 | 1.5 | 97.18 | 19.16 | 6.86 | 4/4 | 1 | 1125.63 | 0 | 98.31 | 19.69 | 7.16 | 15.74 | FALSE | not in blank | 1131 (ν(C-C) fatty acid; L1) | TRUE | C4 - uncertain / weak
1249.4 | 3/8 | 0.38 | 1247.38 | 2.33 | 39.13 | 13.97 | 2.78 | 4/4 | 1 | 1248.27 | 2.29 | 46.34 | 10.37 | 3.37 | 42.38 | FALSE | 60%: blank feature 1246 (FWHM 16) - same width; 70%: blank feature 1248 (FWHM 117) - same width | none | FALSE | C3 - background (in blank)
1348.9 | 5/8 | 0.62 | 1347.07 | 1.94 | 308.96 | 25.46 | 22.42 | 4/4 | 1 | 1350.35 | 2.72 | 446.15 | 1.8 | 32.44 | 125.27 | TRUE | 60%: blank saturated; 70%: blank saturated | none | FALSE | C3 - glass zone (glass-dominated)

## 36_LPS_20260928_Decision
Item | Finding | Decision
Sample identity | Spectra do not match NAM (r 0.14-0.19) and match the April 2026 LPS data (r 0.75-0.82): same pattern with a sharp band at 425 cm-1. | Accepted as LPS (as labelled).
Best power | 70 %: highest reproducibility (pairwise r 0.963), 425 band CV 12 %, SNR 25. 60 %: more spots but CV 25 %. 80 %: 0/5 pass (fixed-pattern ripple), as for NAM at 80-90 %. | Use 70 %. Do not use 80 or 90 %.
Signal strength | Only one strong band (425 cm-1, SNR 20-25). Other bands (765, 997, 1081, 1126, 860-900) have SNR 4-10 per spectrum; glass band dominates 1308-1455. | Increase SNR: 70 % / 30 s / 10 acc (same frame, more averaging, SNR x1.4); thicker packed powder.
Reference status | Core set at 70 % n = 4 (+1 glass-dominated); 60 % n = 8. | Not enough for a final LPS reference: collect 10-15 spots at 70 % / 30 s / 10 acc + 3 matched blanks + 1 short blank.
Sample identity (label) | LPS from E. coli O111:B4, Sigma-Aldrich L2630-10MG, Lot 0000442391; phenol-extracted; smooth LPS (lipid A + core + O-antigen); <=3 % protein (supplier spec); store 2-8 C | Recorded; confirm source/CoA number. Use the same lot for the final 70 % acquisition.

## 37_NAM_Final_Status
Item | Final value
Status | FINAL (28 Sep 2026) – operator decision: existing data sufficient, no further NAM acquisition
Reference condition | 785 nm, 90 % laser power (instrument setting), 30 s, 5 accumulations; bulk powder on glass slide
Reference set | n = 15 (30 May 2026): 5 independent spots + 10 repeats of one spot; files unchanged (03_Reference_Spectra/NAM)
Representative spectrum | Mean (median agrees, r = 0.9992)
Bulk intensity (929 cm-1) | 919 ± 36 counts (CV 3.9 %), SNR 51
Primary markers (cm-1) | 541, 829, 870.5, 929, 955, 1023, 1127
Supporting peaks (cm-1) | 457, 487, 1249, 1589
Uncertain | 584, 1059, 1155; 1345, 1392, 1450 (inside glass zone)
Independent validation | 24 Sep 2026: 21 independent spots (60 %: 11, 70 %: 5, 80 %: 4, 90 %: 1) – all markers 21/21, positions ≤ 1.5 cm-1, primary relative intensities ±7 %; earlier 70 %/10 s/5 acc: 10 spots, all markers found
Limitation | Only 6 independent spots at the reference condition (target 10). Absolute intensity varies between sessions (90 %: 919 counts on 30 May vs 709 counts, n = 1, on 24 Sep); band positions and relative intensities are stable.
Use in the library | Band positions and relative intensities: robust. Absolute counts: indicative for this instrument, pipeline and condition only.

## 38_LPS_Ref_Inventory
original | category | library_file | role | md5 | copy_ok
40-60-5-diffrentposition-/SP_10.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S10.csv | final set | de6d97a39ea27e6094902d3aee3533ac | TRUE
40-60-5-diffrentposition-/SP_11.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S11.csv | final set | fa05ee5b26f3cf471175415a78a64f77 | TRUE
40-60-5-diffrentposition-/SP_12.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S12.csv | final set | 4eb95155a325c8a4eb98eb6912c4c8ca | TRUE
40-60-5-diffrentposition-/SP_13.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S13.csv | final set | 7b2ae137a9d5df0a115fcf6b8833e2f8 | TRUE
40-60-5-diffrentposition-/SP_14.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S14.csv | final set | 56645d8203e1873eece5f69ac413b60e | TRUE
40-60-5-diffrentposition-/SP_15.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S15.csv | final set | d64ecbc5adac788832f69f51e5b3011a | TRUE
40-60-5-diffrentposition-/SP_16.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S16.csv | final set | 87a95c82ab1aa9c2a67fac43e7eaace9 | TRUE
40-60-5-diffrentposition-/SP_17.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S17.csv | final set | f2486710aa33fa5192659f24fa927291 | TRUE
40-60-5-diffrentposition-/SP_18.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S18.csv | final set | f839818f355558594186d72b9de4060b | TRUE
40-60-5-diffrentposition-/SP_19.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S19.csv | final set | 9db76bc2b0585733c45af102848b5218 | TRUE
40-60-5-diffrentposition-/SP_20.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S20.csv | final set | 71ea623219e31f87772362850d1ef65f | TRUE
blank+short/blank.csv | Glass_Blank | Glass_Blank_20260928_P060pct_T40s_A05_S01.csv | matched blank | 5353d7a191ee37e5d2e0e91750c379a1 | TRUE
blank+short/blank2.csv | Glass_Blank | Glass_Blank_20260928_P060pct_T40s_A05_S02.csv | matched blank | a26775e1dbef45e27eb24242e08dfc88 | TRUE
blank+short/blank3.csv | Glass_Blank | Glass_Blank_20260928_P060pct_T40s_A05_S03.csv | matched blank | 0bfab89d5a84a751fe2ed7be897880b5 | TRUE
blank+short/short blank at 60p- 5s-1 acc-1.csv | Glass_Blank | Glass_Blank_20260928_P060pct_T05s_A01_S01_short.csv | short (unsaturated) blank | e8c0eb0b67833477046a8a9c000ecb6a | TRUE
blank+short/short blank at 60p- 5s-1 acc-2.csv | Glass_Blank | Glass_Blank_20260928_P060pct_T05s_A01_S02_short.csv | short (unsaturated) blank | f6dd7237ffdd3d57bb6aa5622d07045a | TRUE
blank+short/shutter-closed check frame at 60p-40 s- 5 acc.csv | Instrument_Checks | Check_20260928_P060pct_T40s_A05_ShutterClosedFrame_S02.csv | shutter-closed check frame (end of session) | 245e45c8fc79eabf2697de08138779a5 | TRUE
unwanted/test-machine/SP_8.csv | Instrument_Checks | Check_20260928_P060pct_T40s_A05_ShutterClosedFrame_S01.csv | shutter-closed check frame (orig SP_8) | 9629cda796ca8fd75777405d3b0c13c3 | TRUE
unwanted/test-machine/dark.csv | Instrument_Checks | Check_20260928_P060pct_T40s_A05_ShutterClosedDark_S01.csv | dark acquired with shutter closed by hand (orig dark.csv) | 7baa5e8bc521bcacaec699fce7922710 | TRUE
unwanted/laseroff test.csv | Instrument_Checks | Check_20260928_P070pct_T30s_A05_LaserOff_S01.csv | laser-off test (orig "laseroff test") | 4d7749df59afd1358c65b2a2e3380d92 | TRUE
unwanted/test-machine/laseroff-manualllyclosed shutter.csv | Instrument_Checks | Check_20260928_P060pct_T30s_A05_ShutterClosedFrame_S01.csv | laser off, shutter closed by hand | e6cde71a9b8ae37a78cd19d57acb4ed7 | TRUE
unwanted/SP_3.csv | LPS_Bulk | LPS_Bulk_20260928_P070pct_T30s_A05_S06.csv | troubleshooting (orig unwanted/SP_3) - ripple | 3faafee9121a09efb7d130441925e861 | TRUE
unwanted/SP_4.csv | LPS_Bulk | LPS_Bulk_20260928_P070pct_T30s_A05_S07.csv | troubleshooting (orig unwanted/SP_4) - ripple | 2a6872f3ef8fffd4300b91160d7359c6 | TRUE
unwanted/SP_5.csv | LPS_Bulk | LPS_Bulk_20260928_P070pct_T30s_A05_S08.csv | troubleshooting (orig unwanted/SP_5) - ripple | 9c9ed3dfa81782196f7bde68ef7212dc | TRUE
unwanted/test-machine/30s-60p-5ac-sp1.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T30s_A05_S11.csv | troubleshooting (orig sp1) - dark contains sample light | a1386c043bba63207a8f035fca9a47e2 | TRUE
unwanted/test-machine/30s-60p-5ac-sp2.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T30s_A05_S12.csv | troubleshooting (orig sp2) - dark contains sample light | 092e4d89ae880d492eb4a82222f9dc08 | TRUE
unwanted/test-machine/SP_6-diffrent.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S06.csv | supporting (orig SP_6-diffrent) - dark contains sample light | 857946c2e9c456b3156e58079c3a42f7 | TRUE
unwanted/test-machine/SP_7-diffrent.csv | LPS_Bulk | LPS_Bulk_20260928_P060pct_T40s_A05_S07.csv | supporting (orig SP_7-diffrent) - dark contains sample light | 1008dfec97750c80c7271b90b07a91d4 | TRUE
unwanted/test-machine/SP_4.txtr | Instrument_Checks | Check_20260928_SP_4_duplicate.txtr | TXTR export of orig SP_4 (duplicate format) | 163fc44c4d6fdc4bc486c31430405c3f | TRUE
unwanted/test-machine/SP_5.txtr | Instrument_Checks | Check_20260928_SP_5_duplicate.txtr | TXTR export of orig SP_5 (duplicate format) | 0d217fae08f9deddf20fc693885a3ae4 | TRUE

## 39_LPS_Ref_QC
library_file | category | role | acq | power | int_s | acc | dark_med | raw_med | raw_max | nsat | sat_from | sat_to | ripple | noise | h425 | snr425 | nspikes | rNAM | rLTA | rLPSapr | rLPS28 | glass_ratio | dark_excess | QC | QC_reason | original | in_reference
LPS_Bulk_20260928_P060pct_T40s_A05_S10.csv | LPS_Bulk | final set | 2026-09-28 15:12:00 | 60 | 40 | 5 | 2089.8 | 3739.8 | 6241.8 | 0 |  |  | 22.427 | 17.138 | 366.75 | 21.4 | 0 | 0.163 | 0.5 | 0.749 | 0.942 | 3.717 | 587.4 | PASS |  | 40-60-5-diffrentposition-/SP_10.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S11.csv | LPS_Bulk | final set | 2026-09-28 15:25:49 | 60 | 40 | 5 | 2055 | 4195.2 | 7936.6 | 0 |  |  | 16.269 | 15.549 | 404.961 | 26.044 | 0 | 0.135 | 0.517 | 0.768 | 0.968 | 1.913 | 552.6 | PASS |  | 40-60-5-diffrentposition-/SP_11.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S12.csv | LPS_Bulk | final set | 2026-09-28 15:33:27 | 60 | 40 | 5 | 1967.8 | 3672.6 | 7610.2 | 0 |  |  | 15.029 | 16.032 | 312.045 | 19.464 | 0 | 0.19 | 0.511 | 0.757 | 0.965 | 2.048 | 465.4 | PASS |  | 40-60-5-diffrentposition-/SP_12.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S13.csv | LPS_Bulk | final set | 2026-09-28 15:41:25 | 60 | 40 | 5 | 1927.2 | 5226.2 | 11550.4 | 0 |  |  | 22.127 | 20.736 | 348.992 | 16.83 | 0 | 0.089 | 0.296 | 0.487 | 0.565 | 12.792 | 424.8 | EXCLUDE | glass-dominated spot (glass-zone/800-1000 area ratio 12.8; r to LPS screening mean 0.57) | 40-60-5-diffrentposition-/SP_13.csv | FALSE
LPS_Bulk_20260928_P060pct_T40s_A05_S14.csv | LPS_Bulk | final set | 2026-09-28 15:49:41 | 60 | 40 | 5 | 2052.4 | 3719.8 | 7679.6 | 0 |  |  | 15.715 | 15.142 | 306.096 | 20.215 | 0 | 0.176 | 0.544 | 0.756 | 0.963 | 1.721 | 550 | PASS |  | 40-60-5-diffrentposition-/SP_14.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S15.csv | LPS_Bulk | final set | 2026-09-28 15:57:05 | 60 | 40 | 5 | 1937.2 | 6398.2 | 10703 | 0 |  |  | 17.996 | 19.303 | 970.642 | 50.285 | 0 | 0.156 | 0.49 | 0.776 | 0.938 | 1.842 | 434.8 | PASS |  | 40-60-5-diffrentposition-/SP_15.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S16.csv | LPS_Bulk | final set | 2026-09-28 16:08:31 | 60 | 40 | 5 | 1903.8 | 8419.4 | 14817.2 | 0 |  |  | 22.651 | 25.002 | 1540.925 | 61.631 | 0 | 0.16 | 0.47 | 0.769 | 0.918 | 1.594 | 401.4 | PASS |  | 40-60-5-diffrentposition-/SP_16.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S17.csv | LPS_Bulk | final set | 2026-09-28 16:16:24 | 60 | 40 | 5 | 1922 | 9224 | 16213.8 | 0 |  |  | 23.368 | 25.913 | 1630.66 | 62.929 | 0 | 0.158 | 0.484 | 0.78 | 0.927 | 1.586 | 419.6 | PASS |  | 40-60-5-diffrentposition-/SP_17.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S18.csv | LPS_Bulk | final set | 2026-09-28 16:26:00 | 60 | 40 | 5 | 1861.8 | 7598.2 | 13131 | 0 |  |  | 18.488 | 23.13 | 1314.424 | 56.828 | 1 | 0.172 | 0.483 | 0.764 | 0.924 | 1.471 | 359.4 | PASS |  | 40-60-5-diffrentposition-/SP_18.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S19.csv | LPS_Bulk | final set | 2026-09-28 16:35:58 | 60 | 40 | 5 | 1844.8 | 5975.4 | 10229.8 | 0 |  |  | 17.734 | 17.406 | 954.555 | 54.84 | 0 | 0.186 | 0.473 | 0.774 | 0.937 | 1.795 | 342.4 | PASS |  | 40-60-5-diffrentposition-/SP_19.csv | TRUE
LPS_Bulk_20260928_P060pct_T40s_A05_S20.csv | LPS_Bulk | final set | 2026-09-28 16:46:52 | 60 | 40 | 5 | 1844.8 | 3719.2 | 7455.4 | 0 |  |  | 18.226 | 15.586 | 302.279 | 19.394 | 0 | 0.153 | 0.49 | 0.754 | 0.927 | 4.751 | 342.4 | PASS |  | 40-60-5-diffrentposition-/SP_20.csv | TRUE
Glass_Blank_20260928_P060pct_T40s_A05_S01.csv | Glass_Blank | matched blank | 2026-09-28 17:04:13 | 60 | 40 | 5 | 1481.6 | 35807.4 | 65535 | 247 | 1278.284 | 1637.521 | 193.541 | 58.651 | 1638.289 | 27.933 | 2 | -0.111 | -0.153 | -0.246 | -0.074 |  | -20.8 | BLANK-SATURATED |  | blank+short/blank.csv | FALSE
Glass_Blank_20260928_P060pct_T40s_A05_S02.csv | Glass_Blank | matched blank | 2026-09-28 17:08:33 | 60 | 40 | 5 | 1481.6 | 22363 | 65535 | 72 | 1329.458 | 1434.312 | 99.548 | 43.065 | 805.637 | 18.707 | 0 | 0.017 | 0.205 | 0.187 | 0.263 |  | -20.8 | BLANK-SATURATED |  | blank+short/blank2.csv | FALSE
Glass_Blank_20260928_P060pct_T40s_A05_S03.csv | Glass_Blank | matched blank | 2026-09-28 17:12:34 | 60 | 40 | 5 | 1545.4 | 14892.8 | 47939.4 | 0 |  |  | 73.689 | 31.764 | 911.747 | 28.704 | 0 | -0.019 | 0.094 | -0.058 | 0.122 |  | 43 | BLANK-OK |  | blank+short/blank3.csv | FALSE
Glass_Blank_20260928_P060pct_T05s_A01_S01_short.csv | Glass_Blank | short (unsaturated) blank | 2026-09-28 17:18:27 | 60 | 5 | 1 | 1086 | 2801 | 7015 | 0 |  |  | 35.86 | 28.124 | 123.958 | 4.407 | 0 | -0.048 | 0.08 | -0.037 | 0.113 |  |  | BLANK-OK |  | blank+short/short blank at 60p- 5s-1 acc-1.csv | FALSE
Glass_Blank_20260928_P060pct_T05s_A01_S02_short.csv | Glass_Blank | short (unsaturated) blank | 2026-09-28 17:18:44 | 60 | 5 | 1 | 1086 | 2780 | 7128 | 0 |  |  | 36.595 | 27.831 | 121.228 | 4.356 | 0 | -0.017 | 0.096 | -0.081 | 0.073 |  |  | BLANK-OK |  | blank+short/short blank at 60p- 5s-1 acc-2.csv | FALSE
Check_20260928_P060pct_T40s_A05_ShutterClosedFrame_S02.csv | Instrument_Checks | shutter-closed check frame (end of session) | 2026-09-28 16:56:13 | 60 | 40 | 5 | 1481.6 | 1479.6 | 5360.6 | 0 |  |  | 14.72 | 12.452 | 18.154 | 1.458 | 0 | 0.056 | 0.008 | 0.106 | 0.043 |  | -20.8 | CHECK |  | blank+short/shutter-closed check frame at 60p-40 s- 5 acc.csv | FALSE
Check_20260928_P060pct_T40s_A05_ShutterClosedFrame_S01.csv | Instrument_Checks | shutter-closed check frame (orig SP_8) | 2026-09-28 14:41:05 | 60 | 40 | 5 | 2344.8 | 1525.2 | 5388.4 | 0 |  |  | 91.481 | 28.889 | 59.593 | 2.063 | 0 | -0.075 | -0.25 | -0.064 | -0.248 |  | 842.4 | CHECK |  | unwanted/test-machine/SP_8.csv | FALSE
Check_20260928_P060pct_T40s_A05_ShutterClosedDark_S01.csv | Instrument_Checks | dark acquired with shutter closed by hand (orig dark.csv) | 2026-09-28 14:58:19 | 60 | 40 | 5 | 1525.4 | 1526 | 5451 | 0 |  |  | 13.152 | 12.868 | 3.512 | 0.273 | 0 | 0.094 | 0.012 | -0.027 | -0.095 |  | 23 | CHECK |  | unwanted/test-machine/dark.csv | FALSE
Check_20260928_P070pct_T30s_A05_LaserOff_S01.csv | Instrument_Checks | laser-off test (orig "laseroff test") | 2026-09-28 12:12:33 | 70 | 30 | 5 | 3309 | 1404.8 | 4326.2 | 0 |  |  | 228.306 | 59.54 | 173.83 | 2.92 | 0 | -0.07 | -0.237 | -0.044 | -0.235 |  |  | CHECK |  | unwanted/laseroff test.csv | FALSE
Check_20260928_P060pct_T30s_A05_ShutterClosedFrame_S01.csv | Instrument_Checks | laser off, shutter closed by hand | 2026-09-28 13:56:37 | 60 | 30 | 5 | 1938.8 | 1412.8 | 4327.8 | 0 |  |  | 59.793 | 20.909 | 47.862 | 2.289 | 0 | -0.08 | -0.246 | -0.047 | -0.237 |  |  | CHECK |  | unwanted/test-machine/laseroff-manualllyclosed shutter.csv | FALSE
LPS_Bulk_20260928_P070pct_T30s_A05_S06.csv | LPS_Bulk | troubleshooting (orig unwanted/SP_3) - ripple | 2026-09-28 11:57:02 | 70 | 30 | 5 | 2987.4 | 6463.8 | 9572.4 | 0 |  |  | 192.82 | 52.393 | 829.764 | 15.837 | 0 | 0.099 | 0.258 | 0.505 | 0.379 | 0.755 |  | FAIL | ripple > 50 (troubleshooting spectrum) | unwanted/SP_3.csv | FALSE
LPS_Bulk_20260928_P070pct_T30s_A05_S07.csv | LPS_Bulk | troubleshooting (orig unwanted/SP_4) - ripple | 2026-09-28 12:01:22 | 70 | 30 | 5 | 3309 | 6665.6 | 9689.8 | 0 |  |  | 114.678 | 36.601 | 711.018 | 19.426 | 0 | 0.145 | 0.35 | 0.647 | 0.558 | 1.581 |  | FAIL | ripple > 50 (troubleshooting spectrum) | unwanted/SP_4.csv | FALSE
LPS_Bulk_20260928_P070pct_T30s_A05_S08.csv | LPS_Bulk | troubleshooting (orig unwanted/SP_5) - ripple | 2026-09-28 12:09:24 | 70 | 30 | 5 | 3309 | 6070.6 | 9124.4 | 0 |  |  | 99.756 | 32.196 | 645.994 | 20.064 | 0 | 0.134 | 0.347 | 0.666 | 0.576 | 1.356 |  | FAIL | ripple > 50 (troubleshooting spectrum) | unwanted/SP_5.csv | FALSE
LPS_Bulk_20260928_P060pct_T30s_A05_S11.csv | LPS_Bulk | troubleshooting (orig sp1) - dark contains sample light | 2026-09-28 13:48:08 | 60 | 30 | 5 | 1938.8 | 3499.4 | 6249.6 | 0 |  |  | 17.538 | 15.079 | 233.928 | 15.514 | 0 | 0.118 | 0.442 | 0.735 | 0.882 | 7.512 |  | SUPPORTING | valid spectrum but dark contained sample light / different condition - not used in reference | unwanted/test-machine/30s-60p-5ac-sp1.csv | FALSE
LPS_Bulk_20260928_P060pct_T30s_A05_S12.csv | LPS_Bulk | troubleshooting (orig sp2) - dark contains sample light | 2026-09-28 13:52:02 | 60 | 30 | 5 | 1938.8 | 3400 | 6352.4 | 0 |  |  | 13.951 | 13.802 | 192.315 | 13.934 | 0 | 0.173 | 0.553 | 0.779 | 0.931 | 1.928 |  | SUPPORTING | valid spectrum but dark contained sample light / different condition - not used in reference | unwanted/test-machine/30s-60p-5ac-sp2.csv | FALSE
LPS_Bulk_20260928_P060pct_T40s_A05_S06.csv | LPS_Bulk | supporting (orig SP_6-diffrent) - dark contains sample light | 2026-09-28 14:24:53 | 60 | 40 | 5 | 2344.8 | 3848.8 | 7712.2 | 0 |  |  | 18.552 | 16.962 | 326.124 | 19.227 | 0 | 0.109 | 0.445 | 0.748 | 0.902 | 2.814 | 842.4 | SUPPORTING | valid spectrum but dark contained sample light / different condition - not used in reference | unwanted/test-machine/SP_6-diffrent.csv | FALSE
LPS_Bulk_20260928_P060pct_T40s_A05_S07.csv | LPS_Bulk | supporting (orig SP_7-diffrent) - dark contains sample light | 2026-09-28 14:32:40 | 60 | 40 | 5 | 2344.8 | 3689.8 | 7663.8 | 0 |  |  | 15.074 | 15.008 | 274.599 | 18.296 | 0 | 0.221 | 0.524 | 0.746 | 0.953 | 1.621 | 842.4 | SUPPORTING | valid spectrum but dark contained sample light / different condition - not used in reference | unwanted/test-machine/SP_7-diffrent.csv | FALSE

## 40_LPS_Ref_Bands
class | band | detected | pos_mean | pos_sd | height_median | height_cv | snr_median | snr_min | rel425_mean | rel425_cv | fwhm | glass_share_pct | snr_after_glass_removal | blank_features | literature | general | note
C1 - primary marker | 425.4 | 10/10 | 424.81 | 0.88 | 679.76 | 66.7 | 38.16 | 19.39 | 1 | 0 | 14.8 | 3.52 | 36.88 | none | none within ±12 | carbohydrate skeletal δ(C-C-O)/δ(C-O-C) region (GEN-1 region I); O-P-O bending region (general) — no LPS literature band | Strongest LPS band; also in April 2026 and 28 Sep screening data; absent from blanks and NAM; no literature assignment
C4 - uncertain / weak | 478.1 | 8/10 | 478.79 | 2.15 | 115.42 | 47.34 | 5.96 | 2.45 | 0.13 | 28.28 | 14.14 | 27.93 | 3.59 | none | none within ±12 | —
C4 - uncertain / weak | 541 | 6/10 | 540.13 | 1.49 | 160.2 | 35.09 | 8.35 | 4.88 | 0.16 | 26.94 | 34.04 | 2.87 | 6.25 | none | 550 β(CH2) in ring [LPS-1 SERS]; 543 δ(C-O-C) glycosidic ring [LPS-1 SERS] | —
C4 - uncertain / weak | 631.2 | 4/10 | 631.23 | 0 | 173.3 | 26.45 | 7.03 | 6.08 | 0.12 | 14.17 | 43.97 | 8.46 | 4.3 | none | none within ±12 | —
C4 - uncertain / weak | 655.7 | 5/10 | 656.74 | 2.34 | 59.37 | 43.68 | 3.77 | 3.39 | 0.13 | 29.44 | 13.33 | 13.48 | 3.94 | none | none within ±12 | —
C2 - supporting peak | 765.9 | 8/10 | 765.51 | 1.2 | 305.47 | 52.2 | 16.7 | 7.02 | 0.34 | 12.4 | 22.27 | 7.24 | 11.26 | none | none within ±12 | —
C2 - supporting peak | 861.1 | 8/10 | 861.1 | 0.89 | 235.32 | 39.77 | 12.73 | 7.65 | 0.31 | 27.36 | 96.69 | 3.54 | 10.54 | none | 850 ν(C-O-C) 1,4-glycosidic [LPS-1 SERS]; 855 δ(C-O-C) [LPS-1 SERS] | anomeric/glycosidic region; β-anomer C1-H ~898 (GEN-2)
C2 - supporting peak | 996.6 | 9/10 | 995.84 | 0.85 | 167.34 | 59.5 | 8.67 | 3.23 | 0.17 | 12.44 | 19.83 | 1.54 | 6.58 | none | 988 β(CH) carbohydrates [LPS-1 SERS]; 1003 common LPS peak [LPS-2 SERS] | —
C1 - primary marker | 1080.6 | 10/10 | 1080.49 | 1.11 | 268.36 | 72.18 | 15.29 | 6.36 | 0.4 | 11.18 | 78.7 | 13.86 | 13.21 | none | 1075 O-antigen [LPS-3 normal Raman]; 1087 ν(C-O) lipid [LPS-1 SERS] | ν(C-O)/ν(C-C) sugar (GEN-1 region II); ν_s(PO2-) ~1080-1090 (GEN-5, GEN-6) | Broad band; glass Si-O feature at ~1091 contributes ~14 %; best literature match 1075 (LPS-3 normal Raman)
C4 - uncertain / weak | 1124.1 | 5/10 | 1123.74 | 1.31 | 433.93 | 26.5 | 18.76 | 15.1 | 0.28 | 18.6 | 13.68 | 10.23 | 9.51 | none | 1131 ν(C-C) fatty acid [LPS-1 SERS] | acyl-chain ν(C-C) 1060-1130 (GEN-5)
C4 - uncertain / weak | 1247.9 | 4/10 | 1247.89 | 1.25 | 149.2 | 48.54 | 6.91 | 3.23 | 0.12 | 17.99 | 33.24 | 29.68 | 3.55 | none | none within ±12 | —
C3 - background / glass | 1257 | 5/10 | 1256.72 | 0.68 | 54.23 | 57.62 | 3.49 | 2.79 | 0.12 | 19.16 | 3.5 | 51.01 | 2.64 | none | none within ±12 | —
C3 - glass zone (glass-dominated) | 1333.9 | 7/10 | 1333.09 | 2.71 | 488 | 38.58 | 25.28 | 13.42 | 0.73 | 57.89 | 111.5 | 77.04 | 5.1 | none | 1330 δ(CH) phospholipid [LPS-1 SERS]; 1333 common LPS peak (C-H def.) [LPS-2 SERS] | —
C4 - uncertain / weak | 1457.6 | 5/10 | 1456.41 | 1.89 | 337.65 | 21.89 | 14.91 | 14.6 | 0.35 | 36.81 | 15.49 | 12.02 | 11.55 | none | 1450 CH2/CH3 deformation [LPS-1 SERS] | — | Exploratory: appears after subtracting the short-blank glass shape (8/10); consistent with CH2/CH3 deformation 1450 (LPS-1) - tentative
C3 - background / glass | 1656.9 | 5/10 | 1656.58 | 2.99 | 143.04 | 20.35 | 6.96 | 5.89 | 0.15 | 28.85 | 133.67 | 53.86 | 2.47 | 1658 (FWHM 95), 1653 (FWHM 93) | none within ±12 | —

## 41_LPS_Ref_Members
library_file | category | role | acq | power | int_s | acc | dark_med | raw_med | raw_max | nsat | sat_from | sat_to | ripple | noise | h425 | snr425 | nspikes | rNAM | rLTA | rLPSapr | rLPS28 | glass_ratio | dark_excess | QC | QC_reason | r_to_others | dark_light_fraction | intensity_factor
LPS_Bulk_20260928_P060pct_T40s_A05_S10.csv | LPS_Bulk | final set | 2026-09-28 15:12:00 | 60 | 40 | 5 | 2089.8 | 3739.8 | 6241.8 | 0 |  |  | 22.427 | 17.138 | 366.75 | 21.4 | 0 | 0.163 | 0.5 | 0.749 | 0.942 | 3.717 | 587.4 | PASS |  | 0.9 | 0.253 | 1.339
LPS_Bulk_20260928_P060pct_T40s_A05_S11.csv | LPS_Bulk | final set | 2026-09-28 15:25:49 | 60 | 40 | 5 | 2055 | 4195.2 | 7936.6 | 0 |  |  | 16.269 | 15.549 | 404.961 | 26.044 | 0 | 0.135 | 0.517 | 0.768 | 0.968 | 1.913 | 552.6 | PASS |  | 0.924 | 0.143 | 1.167
LPS_Bulk_20260928_P060pct_T40s_A05_S12.csv | LPS_Bulk | final set | 2026-09-28 15:33:27 | 60 | 40 | 5 | 1967.8 | 3672.6 | 7610.2 | 0 |  |  | 15.029 | 16.032 | 312.045 | 19.464 | 0 | 0.19 | 0.511 | 0.757 | 0.965 | 2.048 | 465.4 | PASS |  | 0.913 | 0.164 | 1.196
LPS_Bulk_20260928_P060pct_T40s_A05_S14.csv | LPS_Bulk | final set | 2026-09-28 15:49:41 | 60 | 40 | 5 | 2052.4 | 3719.8 | 7679.6 | 0 |  |  | 15.715 | 15.142 | 306.096 | 20.215 | 0 | 0.176 | 0.544 | 0.756 | 0.963 | 1.721 | 550 | PASS |  | 0.911 | 0.186 | 1.229
LPS_Bulk_20260928_P060pct_T40s_A05_S15.csv | LPS_Bulk | final set | 2026-09-28 15:57:05 | 60 | 40 | 5 | 1937.2 | 6398.2 | 10703 | 0 |  |  | 17.996 | 19.303 | 970.642 | 50.285 | 0 | 0.156 | 0.49 | 0.776 | 0.938 | 1.842 | 434.8 | PASS |  | 0.939 | 0.049 | 1.051
LPS_Bulk_20260928_P060pct_T40s_A05_S16.csv | LPS_Bulk | final set | 2026-09-28 16:08:31 | 60 | 40 | 5 | 1903.8 | 8419.4 | 14817.2 | 0 |  |  | 22.651 | 25.002 | 1540.925 | 61.631 | 0 | 0.16 | 0.47 | 0.769 | 0.918 | 1.594 | 401.4 | PASS |  | 0.928 | 0.029 | 1.029
LPS_Bulk_20260928_P060pct_T40s_A05_S17.csv | LPS_Bulk | final set | 2026-09-28 16:16:24 | 60 | 40 | 5 | 1922 | 9224 | 16213.8 | 0 |  |  | 23.368 | 25.913 | 1630.66 | 62.929 | 0 | 0.158 | 0.484 | 0.78 | 0.927 | 1.586 | 419.6 | PASS |  | 0.933 | 0.027 | 1.027
LPS_Bulk_20260928_P060pct_T40s_A05_S18.csv | LPS_Bulk | final set | 2026-09-28 16:26:00 | 60 | 40 | 5 | 1861.8 | 7598.2 | 13131 | 0 |  |  | 18.488 | 23.13 | 1314.424 | 56.828 | 1 | 0.172 | 0.483 | 0.764 | 0.924 | 1.471 | 359.4 | PASS |  | 0.93 | 0.032 | 1.033
LPS_Bulk_20260928_P060pct_T40s_A05_S19.csv | LPS_Bulk | final set | 2026-09-28 16:35:58 | 60 | 40 | 5 | 1844.8 | 5975.4 | 10229.8 | 0 |  |  | 17.734 | 17.406 | 954.555 | 54.84 | 0 | 0.186 | 0.473 | 0.774 | 0.937 | 1.795 | 342.4 | PASS |  | 0.939 | 0.047 | 1.049
LPS_Bulk_20260928_P060pct_T40s_A05_S20.csv | LPS_Bulk | final set | 2026-09-28 16:46:52 | 60 | 40 | 5 | 1844.8 | 3719.2 | 7455.4 | 0 |  |  | 18.226 | 15.586 | 302.279 | 19.394 | 0 | 0.153 | 0.49 | 0.754 | 0.927 | 4.751 | 342.4 | PASS |  | 0.846 | 0.146 | 1.171

## 42_LPS_Final_Status
Item | Value
Status | ESTABLISHED (28 Sep 2026)
Sample | LPS from E. coli O111:B4, Sigma-Aldrich L2630, lot 0000442391 (phenol-extracted; smooth LPS: lipid A + core + O-antigen)
Reference condition | 785 nm, 60 % laser power (instrument setting), 40 s, 5 accumulations; bulk powder directly on glass slide
n | 10 independent spots (S10-S12, S14-S20); S13 excluded (glass-dominated)
Reproducibility | pairwise r (glass zone excluded) mean 0.916, min 0.789; mean vs median r 0.9816, ¦mean − median¦ < 1 SD at all points
Primary markers (C1) | 425 cm-1 (424.8 ± 0.9; SNR median 38; unassigned) and 1080 cm-1 (1080.5 ± 1.1; broad; SNR 15; ~14 % glass contribution)
Supporting (C2) | 766, 861, 997 cm-1
Uncertain (C4) | 478, 541, 631, 656, 1124, 1248, 1458 cm-1
Glass/background (C3) | 1257, 1334 (glass zone), 1657 cm-1
Bulk intensity (425 cm-1) | median 680 counts (range 302-1631), 3-5x spot-to-spot spread from focus/packing
Identity check | r = 0.98 vs LPS screening 28 Sep; 0.80 vs April 2026 LPS; 0.52 vs LTA; 0.17 vs NAM
Limitations | Dark frames contained some sample light (dark 1845-2090 vs true ~1480-1525): intensities under-estimated by 3-34 %; weak bands SNR 4-17; glass zone 1308-1455 cm-1 not separable (exploratory glass removal only); no pure-compound normal-Raman literature for O111:B4
Instrument findings | Fixed-pattern ripple at >= 70-90 % laser power (in the sample light, not removed by dark); software dark step leaks light unless the shutter is closed by hand

## 43_LTA_Ref_Inventory
original | category | library_file | role | md5 | copy_ok
bulk-lta/glass blank/SP_1.csv | Glass_Blank | Glass_Blank_20261028_P080pct_T45s_A02_S01.csv | matched glass blank (80 % / 45 s / 2 acc) | 3efce312bb838105b551ae66bd775afa | TRUE
bulk-lta/glass blank/SP_2.csv | Glass_Blank | Glass_Blank_20261028_P080pct_T45s_A02_S02.csv | matched glass blank (80 % / 45 s / 2 acc) | 368929834cd87ef70a71c446a1319a4a | TRUE
bulk-lta/glass blank/SP_3.csv | Glass_Blank | Glass_Blank_20261028_P080pct_T45s_A02_S03.csv | matched glass blank (80 % / 45 s / 2 acc) | 9266adc245f23977ac8201c0b89a0588 | TRUE
bulk-lta/glass blank/SP_4.csv | Glass_Blank | Glass_Blank_20261028_P080pct_T45s_A02_S04.csv | matched glass blank (80 % / 45 s / 2 acc) | c05f91120411f08405757c14c5c3ed3f | TRUE
bulk-lta/glass blank/SP_5.csv | Glass_Blank | Glass_Blank_20261028_P080pct_T45s_A02_S05.csv | matched glass blank (80 % / 45 s / 2 acc) | bdbefa3c53294f425004cba534665ead | TRUE
bulk-lta/glass blank/SP_6.csv | Glass_Blank | Glass_Blank_20261028_P080pct_T45s_A02_S06.csv | matched glass blank (80 % / 45 s / 2 acc) | 7f533d7c58f05a7554477ea5d95bd89b | TRUE
bulk-lta/glass blank/SP_7.csv | Glass_Blank | Glass_Blank_20261028_P080pct_T45s_A02_S07.csv | matched glass blank (80 % / 45 s / 2 acc) | d19b127b6096f1740844939d73b597e6 | TRUE
bulk-lta/glass blank/SP_8.csv | Glass_Blank | Glass_Blank_20261028_P080pct_T45s_A02_S08.csv | matched glass blank (80 % / 45 s / 2 acc) | 249a2f05543957271893e9dea539f31d | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_9.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S09.csv | LTA spot | bd3520621ecb48ad0e7d0bcadb3ec45f | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_10.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S10.csv | LTA spot | be82021cf1dbd656d24fcb49adb64f65 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_11.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S11.csv | LTA spot | 82b49a419a590a74b3349e2425c656a0 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_12.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S12.csv | LTA spot | 279ff0b6a49ade22a9dd7c44917e9fc6 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_13.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S13.csv | LTA spot | 37d0980cb8af888d93f70b5e5ab3711c | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_14.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S14.csv | LTA spot | 2166a4a9ba89d179544fe37addd4c8b1 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_15.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S15.csv | LTA spot | bba8e27c5cd522e386fa9acd91a896e5 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_16.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S16.csv | LTA spot | 9bfd69bc7554a936411b97506452334f | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_17.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S17.csv | LTA spot | 83775a7a0551b2de666d3f02f483ad58 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_18.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S18.csv | LTA spot | 023bbd58cc32e26f2ad7de733ac693eb | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_19.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S19.csv | LTA spot | b57c795fd1b1aef9ea4d78fbcc881bbe | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_20.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S20.csv | LTA spot | 9c97b9cc7a260ac94fa148ac387ae255 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_21.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S21.csv | LTA spot | 11d16839bc2686f187759bcbb618166e | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_22.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S22.csv | LTA spot | 5fd5cfa8c870320c80d3c2ebbed92f9e | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_23.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S23.csv | LTA spot | d50f5402ed3f8ae71c556b8e8164ce8c | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_24.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S24.csv | LTA spot | 25de733e7f2994c677d94076da972ef6 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_25.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S25.csv | LTA spot | 3695df920ef220d4c0c86fed44f10137 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_26.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S26.csv | LTA spot | 070e346ec7b3ed96b75b0ec62b6ef18c | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_27.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S27.csv | LTA spot | 2c5dc1c563fe1a8e4a25003a3f11814f | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_28.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S28.csv | LTA spot | 95cf164d7baaf3f89a69b751ea01fc72 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_29.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S29.csv | LTA spot | 4f6010f7b7e108fd7e242cee430be760 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_30.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S30.csv | LTA spot | b77b93eccd4e48376102b23c75fb88d6 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_31.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S31.csv | LTA spot | bfba6ddeb2473413c5d6d6c8229f0855 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_32.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S32.csv | LTA spot | 20f819f7a0d21dd7e79e78adb4071631 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_33.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S33.csv | LTA spot | 90ff08ed4b1808d959acdee7cbd4a6e4 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_34.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S34.csv | LTA spot | 22514c90c27015f4768cd87a609ed6fd | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_35.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S35.csv | LTA spot | 1ee6dd481e1c0a73a6e17734ab9f6395 | TRUE
bulk-lta/45sec-80p-2acc-diffrentspot/SP_36.csv | LTA_Bulk | LTA_Bulk_20261028_P080pct_T45s_A02_S36.csv | LTA spot | 19164ee2d6f46afcd07d9ca4caea534a | TRUE

## 44_LTA_Ref_QC
library_file | category | role | acq | power | int_s | acc | dark_med | raw_med | raw_max | nsat | ripple | noise | h1092 | snr1092 | glass_ratio | r_core | rLTA_April | rNAM | rLPS | QC | QC_reason | original | in_reference
Glass_Blank_20261028_P080pct_T45s_A02_S01.csv | Glass_Blank | matched glass blank (80 % / 45 s / 2 acc) | 2026-10-28 02:11:25 | 80 | 45 | 2 | 1504.5 | 3565 | 9889.5 | 0 | 96.27 | 25.31 | [7 empty] | BLANK-SUSPECT | SUSPECT: sharp bands at 619, 1000, 1030, 1154, 1583 and 1602 cm-1 (strongest at 1000, polystyrene-like); r to the 28 Sep glass blanks = 0.07. Not used as the glass template; no saturation (raw max 9890) | bulk-lta/glass blank/SP_1.csv | FALSE
Glass_Blank_20261028_P080pct_T45s_A02_S02.csv | Glass_Blank | matched glass blank (80 % / 45 s / 2 acc) | 2026-10-28 02:13:47 | 80 | 45 | 2 | 1504.5 | 3524 | 9675.5 | 0 | 94.26 | 23.66 | [7 empty] | BLANK-SUSPECT | SUSPECT: sharp bands at 619, 1000, 1030, 1154, 1583 and 1602 cm-1 (strongest at 1000, polystyrene-like); r to the 28 Sep glass blanks = 0.07. Not used as the glass template; no saturation (raw max 9676) | bulk-lta/glass blank/SP_2.csv | FALSE
Glass_Blank_20261028_P080pct_T45s_A02_S03.csv | Glass_Blank | matched glass blank (80 % / 45 s / 2 acc) | 2026-10-28 02:15:33 | 80 | 45 | 2 | 1504.5 | 3496 | 9329.5 | 0 | 91.81 | 25.54 | [7 empty] | BLANK-SUSPECT | SUSPECT: sharp bands at 619, 1000, 1030, 1154, 1583 and 1602 cm-1 (strongest at 1000, polystyrene-like); r to the 28 Sep glass blanks = 0.07. Not used as the glass template; no saturation (raw max 9330) | bulk-lta/glass blank/SP_3.csv | FALSE
Glass_Blank_20261028_P080pct_T45s_A02_S04.csv | Glass_Blank | matched glass blank (80 % / 45 s / 2 acc) | 2026-10-28 02:17:53 | 80 | 45 | 2 | 1504.5 | 3546 | 9263.5 | 0 | 86.27 | 25.06 | [7 empty] | BLANK-SUSPECT | SUSPECT: sharp bands at 619, 1000, 1030, 1154, 1583 and 1602 cm-1 (strongest at 1000, polystyrene-like); r to the 28 Sep glass blanks = 0.07. Not used as the glass template; no saturation (raw max 9264) | bulk-lta/glass blank/SP_4.csv | FALSE
Glass_Blank_20261028_P080pct_T45s_A02_S05.csv | Glass_Blank | matched glass blank (80 % / 45 s / 2 acc) | 2026-10-28 02:19:36 | 80 | 45 | 2 | 1504.5 | 3432 | 9382 | 0 | 93.34 | 22.54 | [7 empty] | BLANK-SUSPECT | SUSPECT: sharp bands at 619, 1000, 1030, 1154, 1583 and 1602 cm-1 (strongest at 1000, polystyrene-like); r to the 28 Sep glass blanks = 0.07. Not used as the glass template; no saturation (raw max 9382) | bulk-lta/glass blank/SP_5.csv | FALSE
Glass_Blank_20261028_P080pct_T45s_A02_S06.csv | Glass_Blank | matched glass blank (80 % / 45 s / 2 acc) | 2026-10-28 02:21:26 | 80 | 45 | 2 | 1504.5 | 3492 | 9791 | 0 | 94.68 | 25.2 | [7 empty] | BLANK-SUSPECT | SUSPECT: sharp bands at 619, 1000, 1030, 1154, 1583 and 1602 cm-1 (strongest at 1000, polystyrene-like); r to the 28 Sep glass blanks = 0.07. Not used as the glass template; no saturation (raw max 9791) | bulk-lta/glass blank/SP_6.csv | FALSE
Glass_Blank_20261028_P080pct_T45s_A02_S07.csv | Glass_Blank | matched glass blank (80 % / 45 s / 2 acc) | 2026-10-28 02:23:13 | 80 | 45 | 2 | 1504.5 | 3504 | 9802 | 0 | 92.85 | 25.17 | [7 empty] | BLANK-SUSPECT | SUSPECT: sharp bands at 619, 1000, 1030, 1154, 1583 and 1602 cm-1 (strongest at 1000, polystyrene-like); r to the 28 Sep glass blanks = 0.07. Not used as the glass template; no saturation (raw max 9802) | bulk-lta/glass blank/SP_7.csv | FALSE
Glass_Blank_20261028_P080pct_T45s_A02_S08.csv | Glass_Blank | matched glass blank (80 % / 45 s / 2 acc) | 2026-10-28 02:25:08 | 80 | 45 | 2 | 1504.5 | 3564 | 10167.5 | 0 | 102.17 | 23.01 | [7 empty] | BLANK-SUSPECT | SUSPECT: sharp bands at 619, 1000, 1030, 1154, 1583 and 1602 cm-1 (strongest at 1000, polystyrene-like); r to the 28 Sep glass blanks = 0.07. Not used as the glass template; no saturation (raw max 10168) | bulk-lta/glass blank/SP_8.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S09.csv | LTA_Bulk | LTA spot | 2026-10-28 02:29:56 | 80 | 45 | 2 | 1504.5 | 4092 | 7232 | 0 | 29.23 | 23.22 | 170.6 | 7.35 | 0.94 | 0.409 | 0.345 | 0.09 | 0.263 | FAIL | SNR of the 1092 band 7.3 (< 10): weak signal, spot probably off the powder | bulk-lta/45sec-80p-2acc-diffrentspot/SP_9.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S10.csv | LTA_Bulk | LTA spot | 2026-10-28 02:31:56 | 80 | 45 | 2 | 1509.5 | 3256 | 5902 | 0 | 26.45 | 21.62 | 96.6 | 4.47 | 0.96 | 0.341 | 0.272 | 0.031 | 0.218 | FAIL | SNR of the 1092 band 4.5 (< 10): weak signal, spot probably off the powder | bulk-lta/45sec-80p-2acc-diffrentspot/SP_10.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S11.csv | LTA_Bulk | LTA spot | 2026-10-28 02:36:44 | 80 | 45 | 2 | 1509.5 | 13112 | 22959.5 | 0 | 38.29 | 40.16 | 1117.8 | 27.83 | 0.85 | 0.992 | 0.929 | 0.153 | 0.454 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_11.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S12.csv | LTA_Bulk | LTA spot | 2026-10-28 02:39:56 | 80 | 45 | 2 | 1509.5 | 34422 | 62381.5 | 0 | 56.28 | 75.88 | 2018.7 | 26.6 | 0.61 | 0.76 | 0.689 | 0.132 | 0.233 | EXCLUDE | atypical shape (r to the core mean 0.76 < 0.80; r to April LTA 0.69): off-powder or glass-influenced spot | bulk-lta/45sec-80p-2acc-diffrentspot/SP_12.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S13.csv | LTA_Bulk | LTA spot | 2026-10-28 02:42:01 | 80 | 45 | 2 | 1509.5 | 10560 | 14655 | 0 | 41.69 | 40.41 | 933.6 | 23.1 | 3.12 | 0.621 | 0.414 | 0.031 | 0.182 | EXCLUDE | atypical shape (r to the core mean 0.62 < 0.80; r to April LTA 0.41): off-powder or glass-influenced spot | bulk-lta/45sec-80p-2acc-diffrentspot/SP_13.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S14.csv | LTA_Bulk | LTA spot | 2026-10-28 02:43:58 | 80 | 45 | 2 | 1509.5 | 36064 | 62424 | 0 | 57.98 | 74.46 | 2302.9 | 30.93 | 0.57 | 0.824 | 0.772 | 0.145 | 0.281 | REVIEW | ripple 58.0 counts (limit 50): marginal; brightest spots, ripple tracks noise | bulk-lta/45sec-80p-2acc-diffrentspot/SP_14.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S15.csv | LTA_Bulk | LTA spot | 2026-10-28 02:47:16 | 80 | 45 | 2 | 1509.5 | 28864 | 53279 | 0 | 55.55 | 65.53 | 2014.2 | 30.74 | 0.53 | 0.94 | 0.911 | 0.176 | 0.489 | REVIEW | ripple 55.5 counts (limit 50): marginal; brightest spots, ripple tracks noise | bulk-lta/45sec-80p-2acc-diffrentspot/SP_15.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S16.csv | LTA_Bulk | LTA spot | 2026-10-28 02:50:41 | 80 | 45 | 2 | 1509.5 | 25753 | 46778.5 | 0 | 52.17 | 63.51 | 1880.7 | 29.61 | 0.51 | 0.949 | 0.926 | 0.169 | 0.473 | REVIEW | ripple 52.2 counts (limit 50): marginal; brightest spots, ripple tracks noise | bulk-lta/45sec-80p-2acc-diffrentspot/SP_16.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S17.csv | LTA_Bulk | LTA spot | 2026-10-28 02:55:11 | 80 | 45 | 2 | 1512.5 | 59702 | 65535 | 320 | 84.18 | 96.85 | 1520.3 | 15.7 | 1.18 | 0.245 | 0.261 | 0.136 | 0.089 | FAIL | saturated: 320 pixels at 65535 in 400-1800 cm-1 | bulk-lta/45sec-80p-2acc-diffrentspot/SP_17.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S18.csv | LTA_Bulk | LTA spot | 2026-10-28 03:06:21 | 80 | 45 | 2 | 1512.5 | 44960 | 65535 | 151 | 74.68 | 89.63 | 2986.3 | 33.32 | 0.58 | 0.833 | 0.874 | 0.217 | 0.369 | FAIL | saturated: 151 pixels at 65535 in 400-1800 cm-1 | bulk-lta/45sec-80p-2acc-diffrentspot/SP_18.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S19.csv | LTA_Bulk | LTA spot | 2026-10-28 03:08:36 | 80 | 45 | 2 | 1512.5 | 39690 | 65535 | 83 | 70.27 | 86.67 | 2763.3 | 31.88 | 0.58 | 0.73 | 0.704 | 0.183 | 0.18 | FAIL | saturated: 83 pixels at 65535 in 400-1800 cm-1 | bulk-lta/45sec-80p-2acc-diffrentspot/SP_19.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S20.csv | LTA_Bulk | LTA spot | 2026-10-28 03:11:01 | 80 | 45 | 2 | 1512.5 | 58972 | 65535 | 314 | 93.11 | 102.38 | 1796.7 | 17.55 | 1.27 | 0.364 | 0.404 | 0.208 | 0.21 | FAIL | saturated: 314 pixels at 65535 in 400-1800 cm-1 | bulk-lta/45sec-80p-2acc-diffrentspot/SP_20.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S21.csv | LTA_Bulk | LTA spot | 2026-10-28 03:13:20 | 80 | 45 | 2 | 1512.5 | 9810 | 16754 | 0 | 35.78 | 35.23 | 801.7 | 22.76 | 1.53 | 0.903 | 0.767 | 0.141 | 0.384 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_21.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S22.csv | LTA_Bulk | LTA spot | 2026-10-28 03:43:19 | 80 | 45 | 2 | 1510 | 27085 | 49937.5 | 0 | 47.86 | 63.14 | 1926.6 | 30.51 | 0.54 | 0.942 | 0.904 | 0.184 | 0.5 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_22.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S23.csv | LTA_Bulk | LTA spot | 2026-10-28 03:47:28 | 80 | 45 | 2 | 1510 | 18240 | 35447.5 | 0 | 45.09 | 48.01 | 1444.3 | 30.08 | 0.52 | 0.952 | 0.927 | 0.176 | 0.519 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_23.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S24.csv | LTA_Bulk | LTA spot | 2026-10-28 03:49:32 | 80 | 45 | 2 | 1510 | 14982 | 29599 | 0 | 37.76 | 45.35 | 1100.7 | 24.27 | 0.55 | 0.952 | 0.933 | 0.177 | 0.531 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_24.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S25.csv | LTA_Bulk | LTA spot | 2026-10-28 03:51:28 | 80 | 45 | 2 | 1510 | 8857 | 14969 | 0 | 39.15 | 33.14 | 892.7 | 26.93 | 4.89 | 0.483 | 0.264 | 0.016 | 0.124 | EXCLUDE | atypical shape (r to the core mean 0.48 < 0.80; r to April LTA 0.26): off-powder or glass-influenced spot | bulk-lta/45sec-80p-2acc-diffrentspot/SP_25.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S26.csv | LTA_Bulk | LTA spot | 2026-10-28 03:54:11 | 80 | 45 | 2 | 1510 | 19719 | 37484.5 | 0 | 40.48 | 49.86 | 1710.9 | 34.31 | 0.53 | 0.968 | 0.951 | 0.179 | 0.508 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_26.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S27.csv | LTA_Bulk | LTA spot | 2026-10-28 03:56:25 | 80 | 45 | 2 | 1512 | 43223 | 65535 | 105 | 68.92 | 88.34 | 2735 | 30.96 | 0.56 | 0.857 | 0.88 | 0.201 | 0.318 | FAIL | saturated: 105 pixels at 65535 in 400-1800 cm-1 | bulk-lta/45sec-80p-2acc-diffrentspot/SP_27.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S28.csv | LTA_Bulk | LTA spot | 2026-10-28 04:00:30 | 80 | 45 | 2 | 1512 | 24888 | 48114 | 0 | 54.46 | 61.64 | 1733.8 | 28.13 | 0.54 | 0.936 | 0.899 | 0.161 | 0.499 | REVIEW | ripple 54.5 counts (limit 50): marginal; brightest spots, ripple tracks noise | bulk-lta/45sec-80p-2acc-diffrentspot/SP_28.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S29.csv | LTA_Bulk | LTA spot | 2026-10-28 04:09:43 | 80 | 45 | 2 | 1512 | 21724 | 41935 | 0 | 46.09 | 56.4 | 1628.8 | 28.88 | 0.55 | 0.953 | 0.932 | 0.173 | 0.494 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_29.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S30.csv | LTA_Bulk | LTA spot | 2026-10-28 04:12:53 | 80 | 45 | 2 | 1512 | 15850 | 30741 | 0 | 39.94 | 44.27 | 1299.4 | 29.35 | 0.45 | 0.96 | 0.958 | 0.175 | 0.508 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_30.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S31.csv | LTA_Bulk | LTA spot | 2026-10-28 04:15:55 | 80 | 45 | 2 | 1512 | 23534 | 46086.5 | 0 | 50.94 | 58.55 | 1509.7 | 25.78 | 0.57 | 0.929 | 0.88 | 0.173 | 0.529 | REVIEW | ripple 50.9 counts (limit 50): marginal; brightest spots, ripple tracks noise | bulk-lta/45sec-80p-2acc-diffrentspot/SP_31.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S32.csv | LTA_Bulk | LTA spot | 2026-10-28 04:19:01 | 80 | 45 | 2 | 1512 | 12519 | 21378 | 0 | 38.87 | 40.12 | 1069.4 | 26.65 | 1.46 | 0.914 | 0.782 | 0.118 | 0.391 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_32.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S33.csv | LTA_Bulk | LTA spot | 2026-10-28 04:24:00 | 80 | 45 | 2 | 1512 | 35874 | 65085 | 0 | 58.38 | 73.14 | 2455 | 33.57 | 0.53 | 0.842 | 0.81 | 0.198 | 0.292 | REVIEW | ripple 58.4 counts (limit 50): marginal; brightest spots, ripple tracks noise | bulk-lta/45sec-80p-2acc-diffrentspot/SP_33.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S34.csv | LTA_Bulk | LTA spot | 2026-10-28 04:26:28 | 80 | 45 | 2 | 1512 | 21157 | 40171.5 | 0 | 43.63 | 50.58 | 2021.4 | 39.96 | 0.48 | 0.969 | 0.968 | 0.183 | 0.5 | PASS |  | bulk-lta/45sec-80p-2acc-diffrentspot/SP_34.csv | TRUE
LTA_Bulk_20261028_P080pct_T45s_A02_S35.csv | LTA_Bulk | LTA spot | 2026-10-28 04:28:41 | 80 | 45 | 2 | 1512 | 26964 | 52858.5 | 0 | 54.39 | 57.83 | 1849.2 | 31.98 | 0.56 | 0.94 | 0.905 | 0.169 | 0.494 | REVIEW | ripple 54.4 counts (limit 50): marginal; brightest spots, ripple tracks noise | bulk-lta/45sec-80p-2acc-diffrentspot/SP_35.csv | FALSE
LTA_Bulk_20261028_P080pct_T45s_A02_S36.csv | LTA_Bulk | LTA spot | 2026-10-28 04:30:39 | 80 | 45 | 2 | 1512 | 29134 | 57587.5 | 0 | 56.98 | 65.98 | 1800.6 | 27.29 | 0.54 | 0.914 | 0.858 | 0.163 | 0.489 | REVIEW | ripple 57.0 counts (limit 50): marginal; brightest spots, ripple tracks noise | bulk-lta/45sec-80p-2acc-diffrentspot/SP_36.csv | FALSE

## 45_LTA_Ref_Bands
class | band | detected | pos_mean | pos_sd | height_mean | height_sd | height_cv | snr_median | rel1092_cv | glass_share_pct | glass_zone
C2 - supporting peak | 419 | 10/10 | 418.7 | 1.792 | 603.697 | 212.547 | 35.207 | 13.128 | 17.938 | 0.718 | FALSE
C4 - uncertain / weak | 475 | 1/10 | 473 | 0 | 147.007 | 43.787 | 29.786 | 2.802 | 23.514 | 3.571 | FALSE
C1 - primary marker | 539 | 10/10 | 539.2 | 1.249 | 773.311 | 293.421 | 37.943 | 16.002 | 17.305 | 0.075 | FALSE
C4 - uncertain / weak | 601 | 0/10 | 601 | 0 | 344.907 | 176.767 | 51.251 | 7.577 | 35.391 | 0.315 | FALSE
C2 - supporting peak | 629 | 8/10 | 629.25 | 0.968 | 585.581 | 244.541 | 41.76 | 12.722 | 20.303 | 0.192 | FALSE
C2 - supporting peak | 723 | 10/10 | 721.7 | 1.9 | 463.012 | 156.945 | 33.897 | 9.947 | 13.63 | 0.54 | FALSE
C2 - supporting peak | 798 | 9/10 | 799.556 | 2.061 | 1060.293 | 289.321 | 27.287 | 21.905 | 4.459 | 0.351 | FALSE
C1 - primary marker | 861 | 10/10 | 861.6 | 2.107 | 747.43 | 235.008 | 31.442 | 15.578 | 7.55 | 0.159 | FALSE
C4 - uncertain / weak | 913 | 5/10 | 913.6 | 0.8 | 444.772 | 120.905 | 27.184 | 9.39 | 8.939 | 0.433 | FALSE
C4 - uncertain / weak | 961 | 5/10 | 960.4 | 0.8 | 100.041 | 42.645 | 42.628 | 2.136 | 34.175 | 2.25 | FALSE
C4 - uncertain / weak | 1009 | 7/10 | 1009.714 | 1.666 | 425.589 | 122.854 | 28.867 | 8.771 | 13.249 | 0.488 | FALSE
C4 - uncertain / weak | 1041 | 0/10 | 1041 | 0 | 87.932 | 50.637 | 57.586 | 1.86 | 76.972 | 3.53 | FALSE
C1 - primary marker | 1092 | 10/10 | 1092 | 0.447 | 1412.077 | 402.703 | 28.518 | 29.117 | 0 | 0.471 | FALSE
C4 - uncertain / weak | 1174 | 8/10 | 1172.25 | 2.537 | 145.816 | 75.159 | 51.544 | 2.947 | 34.414 | 0.237 | FALSE
C4 - uncertain / weak | 1246 | 7/10 | 1246.429 | 0.495 | 412.159 | 92.788 | 22.513 | 9.188 | 22.367 | 1.181 | FALSE
C4 - uncertain / weak | 1261 | 1/10 | 1262 | 0 | 388.159 | 103.91 | 26.77 | 7.736 | 44.527 | 1.866 | FALSE
C3 - glass zone | 1328 | 10/10 | 1327.5 | 1.91 | 712.07 | 272.933 | 38.33 | 11.666 | 65.238 | 6.664 | TRUE
C2 - supporting peak | 1459 | 9/10 | 1460.333 | 2.055 | 730.362 | 260.782 | 35.706 | 15.219 | 11.713 | 0.565 | FALSE
C2 - supporting peak | 1583 | 8/10 | 1583.25 | 1.09 | 351.883 | 130.901 | 37.2 | 7.316 | 15.5 | 0.718 | FALSE
C4 - uncertain / weak | 1658 | 2/10 | 1656 | 2 | 243.438 | 39.635 | 16.281 | 4.885 | 24.016 | 3.212 | FALSE
C2 - supporting peak | 1751 | 9/10 | 1750.667 | 0.667 | 247.026 | 87.068 | 35.246 | 5.26 | 16.897 | 0.588 | FALSE

## 46_LTA_Ref_Members
library_file | original_file | acq | spot | power_pct | int_s | acc | dark_med | raw_max | ripple | noise | h1092 | snr1092 | glass_ratio | r_core | rLTA_April | rNAM | rLPS
LTA_Bulk_20261028_P080pct_T45s_A02_S11.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_11.csv | 2026-10-28 02:36:44 | SP_11 | 80 | 45 | 2 | 1509.5 | 22959.5 | 38.29 | 40.16 | 1117.8 | 27.8 | 0.85 | 0.992 | 0.929 | 0.153 | 0.454
LTA_Bulk_20261028_P080pct_T45s_A02_S21.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_21.csv | 2026-10-28 03:13:20 | SP_21 | 80 | 45 | 2 | 1512.5 | 16754 | 35.78 | 35.23 | 801.7 | 22.8 | 1.53 | 0.903 | 0.767 | 0.141 | 0.384
LTA_Bulk_20261028_P080pct_T45s_A02_S22.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_22.csv | 2026-10-28 03:43:19 | SP_22 | 80 | 45 | 2 | 1510 | 49937.5 | 47.86 | 63.14 | 1926.6 | 30.5 | 0.54 | 0.942 | 0.904 | 0.184 | 0.5
LTA_Bulk_20261028_P080pct_T45s_A02_S23.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_23.csv | 2026-10-28 03:47:28 | SP_23 | 80 | 45 | 2 | 1510 | 35447.5 | 45.09 | 48.01 | 1444.3 | 30.1 | 0.52 | 0.952 | 0.927 | 0.176 | 0.519
LTA_Bulk_20261028_P080pct_T45s_A02_S24.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_24.csv | 2026-10-28 03:49:32 | SP_24 | 80 | 45 | 2 | 1510 | 29599 | 37.76 | 45.35 | 1100.7 | 24.3 | 0.55 | 0.952 | 0.933 | 0.177 | 0.531
LTA_Bulk_20261028_P080pct_T45s_A02_S26.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_26.csv | 2026-10-28 03:54:11 | SP_26 | 80 | 45 | 2 | 1510 | 37484.5 | 40.48 | 49.86 | 1710.9 | 34.3 | 0.53 | 0.968 | 0.951 | 0.179 | 0.508
LTA_Bulk_20261028_P080pct_T45s_A02_S29.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_29.csv | 2026-10-28 04:09:43 | SP_29 | 80 | 45 | 2 | 1512 | 41935 | 46.09 | 56.4 | 1628.8 | 28.9 | 0.55 | 0.953 | 0.932 | 0.173 | 0.494
LTA_Bulk_20261028_P080pct_T45s_A02_S30.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_30.csv | 2026-10-28 04:12:53 | SP_30 | 80 | 45 | 2 | 1512 | 30741 | 39.94 | 44.27 | 1299.4 | 29.4 | 0.45 | 0.96 | 0.958 | 0.175 | 0.508
LTA_Bulk_20261028_P080pct_T45s_A02_S32.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_32.csv | 2026-10-28 04:19:01 | SP_32 | 80 | 45 | 2 | 1512 | 21378 | 38.87 | 40.12 | 1069.4 | 26.7 | 1.46 | 0.914 | 0.782 | 0.118 | 0.391
LTA_Bulk_20261028_P080pct_T45s_A02_S34.csv | bulk-lta/45sec-80p-2acc-diffrentspot/SP_34.csv | 2026-10-28 04:26:28 | SP_34 | 80 | 45 | 2 | 1512 | 40171.5 | 43.63 | 50.58 | 2021.4 | 40 | 0.48 | 0.969 | 0.968 | 0.183 | 0.5

## 47_LTA_Final_Status
Item | Value
Status | ESTABLISHED, pending operator confirmation (28 Oct 2026)
Sample | LTA from Staphylococcus aureus, Sigma-Aldrich L2515-10MG, lot 0000419622, source number 0000404178 (label: bacterial cell wall polymer; no purity figure on the supplier page found, lot certificate not obtained; also the vial used on 3 Apr 2026)
Condition | 785 nm, 80 % / 45 s / 2 acc, bulk powder on glass, shutter-closed dark
Reference set | n = 10 independent spots (SP_11, SP_21, SP_22, SP_23, SP_24, SP_26, SP_29, SP_30, SP_32, SP_34)
Extended set | n = 18 (adds marginal-ripple spots); r to strict mean 0.986
Comparison | r to April LTA 0.95; to LPS 0.50; to NAM 0.17 (glass zone excluded)
Primary markers | 539, 861, 1092
Supporting | 419, 629, 723, 798, 1459, 1583, 1751